## Sheet
instruments: | Meter: HP | S: FL5700 | X: FL5730
Start Row | 346.0 | Stop Row | 961.0
| X Range | X Settings (V) | S Range | S Settings (V) | DVM Range | Nominal reading | #Readings | Delay (S) | DVM pause | DVM status | S status | X status | Mean | STD | start time | end time | readings...
| 2 | 0.05 | 0.2 | 0.0 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9813984E+00 | DC220MV |  | 0.050005414433499998 | 9.3526787299618658e-08 | 1517190251.164 | 1517190266.588 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9655049E+00 | DC220MV |  | 5.2869658231500008e-06 | 1.0475166962934471e-07 | 1517190281.423 | 1517190296.918 | 5.1209e-06 | 5.1562e-06 | 5.1562e-06 | 5.2106e-06 | 5.2797e-06 | 5.34e-06 | 5.3826e-06 | 5.49e-06 | 5.3562e-06 | 5.3606e-06 | 5.337e-06 | 5.3047e-06 | 5.4018e-06 | 5.2591e-06 | 5.3473e-06 | 5.3341e-06 | 5.3576e-06 | 5.2826e-06 | 5.2091e-06 | 5.0532e-06
| 2 | 0.05 | 0.2 | 0.05 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0107146E+00 | DC220MV |  | 5.0548333984999998e-06 | 6.1984025783800885e-08 | 1517190311.754 | 1517190327.238 | 5.0258e-06 | 5.1112e-06 | 5.12e-06 | 5.0582e-06 | 4.9493e-06 | 5.1406e-06 | 5.1127e-06 | 5.0303e-06 | 5.0347e-06 | 5.07e-06 | 5.1038e-06 | 5.0715e-06 | 5.0008e-06 | 4.8684e-06 | 5.0935e-06 | 5.0464e-06 | 5.0214e-06 | 5.0612e-06 | 5.0979e-06 | 5.0788e-06
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9988066E+00 | DC220MV |  | 5.1990499183499988e-06 | 5.2480321669004228e-08 | 1517190342.072 | 1517190357.547 | 5.1488e-06 | 5.1738e-06 | 5.1444e-06 | 5.1429e-06 | 5.1341e-06 | 5.24e-06 | 5.2533e-06 | 5.1782e-06 | 5.1414e-06 | 5.1768e-06 | 5.2135e-06 | 5.2121e-06 | 5.1826e-06 | 5.1576e-06 | 5.1561e-06 | 5.2842e-06 | 5.2003e-06 | 5.2518e-06 | 5.2842e-06 | 5.3048e-06
| 2 | 0.05 | 0.2 | 0.05 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9391214E+00 | DC220MV |  | 5.2991170321000005e-06 | 4.5582376322810435e-08 | 1517190372.368 | 1517190387.846 | 5.2544e-06 | 5.2456e-06 | 5.2868e-06 | 5.2603e-06 | 5.2912e-06 | 5.3413e-06 | 5.3222e-06 | 5.3325e-06 | 5.225e-06 | 5.3236e-06 | 5.2427e-06 | 5.3266e-06 | 5.2677e-06 | 5.3236e-06 | 5.3163e-06 | 5.2162e-06 | 5.3751e-06 | 5.3575e-06 | 5.3177e-06 | 5.356e-06
| 2 | 0.05 | 0.2 | 0.0 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9865744E+00 | DC220MV |  | 0.050005516006000006 | 1.2141039969463297e-07 | 1517190402.708 | 1517190418.175 | 0.050006 | 0.050005 | 0.050006 | 0.050005 | 0.050006 | 0.050005 | 0.050006 | 0.050005 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050005 | 0.050006 | 0.050006 | 0.050006 | 0.050005 | 0.050006 | 0.050006 | 0.050006
| 2 | 0.1 | 0.2 | 0.05 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0010378E+00 | DC220MV |  | 0.050005500380500015 | 6.0763065794940484e-08 | 1517190433.012 | 1517190448.46 | 0.050006 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050006 | 0.050005 | 0.050005 | 0.050005 | 0.050006 | 0.050006 | 0.050006 | 0.050005 | 0.050005 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050005 | 0.050006
| 2 | 0.1 | 0.2 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9837364E+00 | DC220MV |  | 0.10000529059499999 | 5.1769606573296943e-07 | 1517190463.295 | 1517190478.723 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001
| 2 | 0.15 | 0.2 | 0.05 | 1 | 0.09999999999999999 | 20.0 | 10.0 | 0.0 | 49.9951703E+00 | DC220MV |  | 0.10000570778499999 | 4.609030565220838e-08 | 1517190493.563 | 1517190509 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001
| 2 | 0.15 | 0.2 | 0.0 | 1 | 0.15 | 20.0 | 10.0 | 0.0 | 50.0190497E+00 | DC220MV |  | 0.15000557722999999 | 4.5553498778627674e-07 | 1517190523.843 | 1517190539.274 | 0.15001 | 0.15001 | 0.15001 | 0.15 | 0.15 | 0.15001 | 0.15 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001
| 2 | 0.2 | 0.2 | 0.05 | 1 | 0.15 | 20.0 | 10.0 | 0.0 | 49.9818651E+00 | DC220MV |  | 0.15000590408999998 | 4.6462015667352243e-08 | 1517190554.095 | 1517190569.526 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001
| 2 | 0.2 | 0.2 | 0.0 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0181009E+00 | DC220MV |  | 0.20000596213500002 | 6.8663624102467797e-08 | 1517190584.347 | 1517190599.785 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001
| 2 | 0.25 | 0.2 | 0.05 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9517918E+00 | DC220MV |  | 0.20000594656499998 | 6.4905864719188817e-08 | 1517190614.635 | 1517190630.058 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001
| 2 | 0.25 | 0.2 | 0.0 | 1 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0242146E+00 | DC220MV |  | 0.25000602625500001 | 7.9236509737293095e-08 | 1517190644.909 | 1517190660.343 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001
| 2 | 0.3 | 0.2 | 0.05 | 1 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0346230E+00 | DC220MV |  | 0.25000598145999997 | 6.9233260788424089e-08 | 1517190675.186 | 1517190690.626 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001
| 2 | 0.3 | 0.2 | 0.0 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9880547E+00 | DC220MV |  | 0.30000611621500001 | 2.9728021713348698e-08 | 1517190705.476 | 1517190720.876 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001
| 2 | 0.35 | 0.2 | 0.05 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9162255E+00 | DC220MV |  | 0.30000583723500002 | 4.8763671546771506e-07 | 1517190735.721 | 1517190751.155 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.3 | 0.3 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001
| 2 | 0.35 | 0.2 | 0.0 | 1 | 0.35 | 20.0 | 10.0 | 0.0 | 50.0142588E+00 | DC220MV |  | 0.35000618963000002 | 7.4626081903526073e-08 | 1517190765.992 | 1517190781.458 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001
| 2 | 0.4 | 0.2 | 0.05 | 1 | 0.35 | 20.0 | 10.0 | 0.0 | 50.0030474E+00 | DC220MV |  | 0.350006037145 | 4.8664488849773537e-08 | 1517190796.313 | 1517190811.743 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001
| 2 | 0.4 | 0.2 | 0.0 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9809601E+00 | DC220MV |  | 0.40000626821499996 | 8.4781503146031731e-08 | 1517190826.579 | 1517190841.999 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001
| 2 | 0.45 | 0.2 | 0.05 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0343443E+00 | DC220MV |  | 0.40000623631000004 | 4.5112114778677855e-08 | 1517190856.859 | 1517190872.295 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001
| 2 | 0.45 | 0.2 | 0.0 | 1 | 0.45 | 20.0 | 10.0 | 0.0 | 50.0493080E+00 | DC220MV |  | 0.45000649213499999 | 5.1447296088732246e-08 | 1517190887.132 | 1517190902.548 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001
| 2 | 0.5 | 0.2 | 0.05 | 1 | 0.45 | 20.0 | 10.0 | 0.0 | 50.0008774E+00 | DC220MV |  | 0.45000636849499998 | 5.9728606838730163e-08 | 1517190917.388 | 1517190932.789 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001
| 2 | 0.5 | 0.2 | 0.0 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 49.9885356E+00 | DC220MV |  | 0.50000649724000001 | 6.7472975321213078e-08 | 1517190947.65 | 1517190963.087 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001
| 2 | 0.5 | 0.2 | 0.05 | 1 | 0.45 | 20.0 | 10.0 | 0.0 | 50.0319166E+00 | DC220MV |  | 0.45000635155500002 | 4.8178579009426212e-08 | 1517190977.948 | 1517190993.38 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001
| 2 | 0.45 | 0.2 | 0.0 | 1 | 0.45 | 20.0 | 10.0 | 0.0 | 49.9891769E+00 | DC220MV |  | 0.45000643186000006 | 4.7469594478399957e-08 | 1517191008.241 | 1517191023.658 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001 | 0.45001
| 2 | 0.45 | 0.2 | 0.05 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9658394E+00 | DC220MV |  | 0.40000637407 | 5.2243718287846257e-08 | 1517191038.489 | 1517191053.912 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001
| 2 | 0.4 | 0.2 | 0.0 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0246962E+00 | DC220MV |  | 0.40000651997499997 | 4.2216879025588208e-08 | 1517191068.774 | 1517191084.216 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001
| 2 | 0.4 | 0.2 | 0.05 | 1 | 0.35 | 20.0 | 10.0 | 0.0 | 49.9924351E+00 | DC220MV |  | 0.35000640440500008 | 8.6221386423230202e-08 | 1517191099.081 | 1517191114.505 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001
| 2 | 0.35 | 0.2 | 0.0 | 1 | 0.35 | 20.0 | 10.0 | 0.0 | 50.0338862E+00 | DC220MV |  | 0.35000638055999994 | 1.8075978645897105e-07 | 1517191129.337 | 1517191144.779 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001 | 0.35001
| 2 | 0.35 | 0.2 | 0.05 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0572116E+00 | DC220MV |  | 0.30000601382499997 | 2.8069916436558445e-07 | 1517191159.811 | 1517191175.216 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001
| 2 | 0.3 | 0.2 | 0.0 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9435967E+00 | DC220MV |  | 0.30000636336499997 | 5.1388454681279527e-08 | 1517191190.069 | 1517191205.472 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001
| 2 | 0.3 | 0.2 | 0.05 | 1 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0075719E+00 | DC220MV |  | 0.25000611645500004 | 2.2701158445732028e-07 | 1517191220.328 | 1517191235.785 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001
| 2 | 0.25 | 0.2 | 0.0 | 1 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0004198E+00 | DC220MV |  | 0.25000632564999997 | 1.6758944925196845e-07 | 1517191250.633 | 1517191266.067 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001 | 0.25001
| 2 | 0.25 | 0.2 | 0.05 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0118943E+00 | DC220MV |  | 0.20000626090500001 | 5.4028904069590046e-08 | 1517191280.917 | 1517191296.336 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001
| 2 | 0.2 | 0.2 | 0.0 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9973916E+00 | DC220MV |  | 0.20000631058500001 | 4.6118719358967975e-08 | 1517191311.18 | 1517191326.612 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001
| 2 | 0.2 | 0.2 | 0.05 | 1 | 0.15 | 20.0 | 10.0 | 0.0 | 50.0100067E+00 | DC220MV |  | 0.15000617074500003 | 8.6682648060446603e-08 | 1517191341.44 | 1517191356.877 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001
| 2 | 0.15 | 0.2 | 0.0 | 1 | 0.15 | 20.0 | 10.0 | 0.0 | 49.9823836E+00 | DC220MV |  | 0.15000625577000001 | 5.4129456860820004e-08 | 1517191371.71 | 1517191387.143 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001 | 0.15001
| 2 | 0.15 | 0.2 | 0.05 | 1 | 0.09999999999999999 | 20.0 | 10.0 | 0.0 | 50.0019388E+00 | DC220MV |  | 0.10000619388500001 | 7.4893779949692653e-08 | 1517191401.993 | 1517191417.442 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001
| 2 | 0.1 | 0.2 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0290074E+00 | DC220MV |  | 0.10000639718 | 6.5001527675462791e-08 | 1517191432.295 | 1517191447.724 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001
| 2 | 0.1 | 0.2 | 0.05 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0122908E+00 | DC220MV |  | 0.05000579172099999 | 5.1228535956946215e-07 | 1517191462.568 | 1517191477.997 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050005 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006
| 2 | 0.05 | 0.2 | 0.0 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9433473E+00 | DC220MV |  | 0.0500062099245 | 4.8753033391009732e-08 | 1517191492.853 | 1517191508.282 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006
| 2 | 0.05 | 0.2 | 0.05 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9872815E+00 | DC220MV |  | 6.166916528449999e-06 | 5.0729810319534488e-08 | 1517191523.125 | 1517191538.632 | 6.1829e-06 | 6.127e-06 | 6.102e-06 | 6.1402e-06 | 6.1535e-06 | 6.1887e-06 | 6.2064e-06 | 6.1858e-06 | 6.2814e-06 | 6.1373e-06 | 6.2284e-06 | 6.0756e-06 | 6.1108e-06 | 6.152e-06 | 6.1505e-06 | 6.1314e-06 | 6.224e-06 | 6.1505e-06 | 6.1564e-06 | 6.2534e-06
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0100775E+00 | DC220MV |  | 6.2305827655000003e-06 | 4.2440450866403604e-08 | 1517191553.467 | 1517191568.952 | 6.1884e-06 | 6.187e-06 | 6.2311e-06 | 6.2443e-06 | 6.2708e-06 | 6.1796e-06 | 6.1811e-06 | 6.2311e-06 | 6.2605e-06 | 6.2473e-06 | 6.2782e-06 | 6.1575e-06 | 6.234e-06 | 6.2782e-06 | 6.2826e-06 | 6.1722e-06 | 6.2576e-06 | 6.1752e-06 | 6.2855e-06 | 6.2694e-06
| 2 | 0.05 | 0.2 | 0.05 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0000943E+00 | DC220MV |  | 6.1673826632999995e-06 | 7.1229930272252583e-08 | 1517191583.785 | 1517191599.259 | 6.2009e-06 | 6.1214e-06 | 6.1788e-06 | 6.2126e-06 | 6.248e-06 | 6.1685e-06 | 6.1803e-06 | 6.2141e-06 | 6.0964e-06 | 6.2391e-06 | 6.2583e-06 | 6.0669e-06 | 6.1332e-06 | 6.1965e-06 | 6.3407e-06 | 6.1008e-06 | 6.0846e-06 | 6.139e-06 | 6.1008e-06 | 6.0669e-06
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0205651E+00 | DC220MV |  | 6.1051248741499998e-06 | 5.4446629733478839e-08 | 1517191614.104 | 1517191629.583 | 6.0506e-06 | 6.1021e-06 | 6.0992e-06 | 6.1242e-06 | 6.1374e-06 | 6.0933e-06 | 6.0565e-06 | 6.0521e-06 | 6.1683e-06 | 6.0992e-06 | 6.2007e-06 | 6.03e-06 | 6.2507e-06 | 6.133e-06 | 6.0918e-06 | 6.0359e-06 | 6.058e-06 | 6.1139e-06 | 6.0756e-06 | 6.1301e-06
| 2 | 0.05 | 0.2 | 0.0 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9738203E+00 | DC220MV |  | 0.050006153549499999 | 2.1670347889440122e-07 | 1517191644.405 | 1517191659.829 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006 | 0.050006
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | 0.0 | 1 | -0.005 | 20.0 | 10.0 | 0.0 | 49.9809412E+00 | DC220MV |  | -0.049993802563499994 | 6.0875233902410485e-08 | 1517191674.659 | 1517191690.126 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0295883E+00 | DC220MV |  | 6.2277764628999999e-06 | 4.9674835989057519e-08 | 1517191704.948 | 1517191720.418 | 6.2094e-06 | 6.1696e-06 | 6.1564e-06 | 6.1932e-06 | 6.2521e-06 | 6.1696e-06 | 6.2418e-06 | 6.227e-06 | 6.2653e-06 | 6.1299e-06 | 6.2197e-06 | 6.2064e-06 | 6.2241e-06 | 6.227e-06 | 6.2697e-06 | 6.2079e-06 | 6.2638e-06 | 6.3065e-06 | 6.2845e-06 | 6.3316e-06
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | -0.005 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9753141E+00 | DC220MV |  | -0.044993783963999993 | 7.8177315532875353e-08 | 1517191735.272 | 1517191750.715 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0049109E+00 | DC220MV |  | 6.0903568527999988e-06 | 4.7168656980711985e-08 | 1517191765.539 | 1517191781.044 | 6.0539e-06 | 6.0333e-06 | 6.0848e-06 | 6.1377e-06 | 6.0759e-06 | 6.1112e-06 | 6.1715e-06 | 6.0759e-06 | 6.0407e-06 | 6.0304e-06 | 6.0539e-06 | 6.2068e-06 | 6.026e-06 | 6.1406e-06 | 6.0833e-06 | 6.0862e-06 | 6.0671e-06 | 6.0848e-06 | 6.1215e-06 | 6.1215e-06
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | -0.005 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0290404E+00 | DC220MV |  | -0.044993961300499999 | 2.7443021008128448e-07 | 1517191795.887 | 1517191811.322 | -0.044994 | -0.044995 | -0.044995 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994
| 2 | -0.05 | 0.2 | 0.0 | 1 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9874088E+00 | DC220MV |  | -0.049994341217999998 | 2.7886331970328089e-07 | 1517191826.162 | 1517191841.594 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049995 | -0.049995 | -0.049995 | -0.049994 | -0.049994 | -0.049995 | -0.049995 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049995 | -0.049994
| 2 | -0.1 | 0.2 | -0.005 | 1 | -0.095 | 20.0 | 10.0 | 0.0 | 49.9803379E+00 | DC220MV |  | -0.094993877350000006 | 6.6726868277107527e-08 | 1517191856.433 | 1517191871.89 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994
| 2 | -0.1 | 0.2 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9999623E+00 | DC220MV |  | -0.099993800081999998 | 5.3133046552571424e-08 | 1517191886.728 | 1517191902.189 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994
| 2 | -0.15 | 0.2 | -0.005 | 1 | -0.145 | 20.0 | 10.0 | 0.0 | 49.9844624E+00 | DC220MV |  | -0.14499413046000001 | 4.1077176141918525e-08 | 1517191917.041 | 1517191932.476 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499
| 2 | -0.15 | 0.2 | 0.0 | 1 | -0.15 | 20.0 | 10.0 | 0.0 | 50.0060337E+00 | DC220MV |  | -0.14999403895999999 | 5.4877330473575414e-08 | 1517191947.307 | 1517191962.753 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999
| 2 | -0.2 | 0.2 | -0.005 | 1 | -0.195 | 20.0 | 10.0 | 0.0 | 49.9798006E+00 | DC220MV |  | -0.19499438000499997 | 2.6731456465301684e-07 | 1517191977.592 | 1517191993.025 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499
| 2 | -0.2 | 0.2 | 0.0 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 50.0312648E+00 | DC220MV |  | -0.19999430018 | 7.065048903251115e-08 | 1517192007.853 | 1517192023.303 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999
| 2 | -0.25 | 0.2 | -0.005 | 1 | -0.245 | 20.0 | 10.0 | 0.0 | 49.9661597E+00 | DC220MV |  | -0.24499456519000001 | 6.4986790196272573e-08 | 1517192038.127 | 1517192053.558 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499
| 2 | -0.25 | 0.2 | 0.0 | 1 | -0.25 | 20.0 | 10.0 | 0.0 | 50.0045240E+00 | DC220MV |  | -0.24999455495 | 1.3897484844056438e-07 | 1517192068.42 | 1517192083.872 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999
| 2 | -0.3 | 0.2 | -0.005 | 1 | -0.295 | 20.0 | 10.0 | 0.0 | 50.0046325E+00 | DC220MV |  | -0.294994605475 | 1.1538752478189031e-07 | 1517192098.737 | 1517192114.182 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499
| 2 | -0.3 | 0.2 | 0.0 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9767278E+00 | DC220MV |  | -0.29999476541999998 | 5.6217013426749897e-08 | 1517192129.017 | 1517192144.444 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999
| 2 | -0.35 | 0.2 | -0.005 | 1 | -0.345 | 20.0 | 10.0 | 0.0 | 49.9714550E+00 | DC220MV |  | -0.34499464357499998 | 6.147696214915326e-08 | 1517192159.279 | 1517192174.743 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499
| 2 | -0.35 | 0.2 | 0.0 | 1 | -0.35 | 20.0 | 10.0 | 0.0 | 50.0204423E+00 | DC220MV |  | -0.34999469005500006 | 5.1576297604742108e-08 | 1517192189.584 | 1517192205.038 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999
| 2 | -0.4 | 0.2 | -0.005 | 1 | -0.395 | 20.0 | 10.0 | 0.0 | 49.9743245E+00 | DC220MV |  | -0.39499479245000002 | 3.9716149106381959e-08 | 1517192219.876 | 1517192235.302 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499
| 2 | -0.4 | 0.2 | 0.0 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 49.9703525E+00 | DC220MV |  | -0.39999474083500003 | 7.0605426669842733e-08 | 1517192250.137 | 1517192265.587 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999
| 2 | -0.45 | 0.2 | -0.005 | 1 | -0.445 | 20.0 | 10.0 | 0.0 | 49.9950948E+00 | DC220MV |  | -0.44499472016999986 | 4.4037337565560766e-08 | 1517192280.434 | 1517192295.885 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499
| 2 | -0.45 | 0.2 | 0.0 | 1 | -0.45 | 20.0 | 10.0 | 0.0 | 50.0123757E+00 | DC220MV |  | -0.44999471117500001 | 5.5377999918460315e-08 | 1517192310.736 | 1517192326.175 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999
| 2 | -0.5 | 0.2 | -0.005 | 1 | -0.495 | 20.0 | 10.0 | 0.0 | 50.0082231E+00 | DC220MV |  | -0.49499478907500005 | 5.6448834133017615e-08 | 1517192341.023 | 1517192356.452 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499
| 2 | -0.5 | 0.2 | 0.0 | 1 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9815964E+00 | DC220MV |  | -0.49999477580500001 | 7.6314936117501984e-08 | 1517192371.309 | 1517192386.745 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999 | -0.49999
| 2 | -0.5 | 0.2 | -0.005 | 1 | -0.495 | 20.0 | 10.0 | 0.0 | 50.0389454E+00 | DC220MV |  | -0.49499479178499994 | 3.3653681452074619e-08 | 1517192401.619 | 1517192417.061 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499 | -0.49499
| 2 | -0.45 | 0.2 | 0.0 | 1 | -0.45 | 20.0 | 10.0 | 0.0 | 49.9738345E+00 | DC220MV |  | -0.449994775545 | 6.5871332723506196e-08 | 1517192431.879 | 1517192447.306 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999 | -0.44999
| 2 | -0.45 | 0.2 | -0.005 | 1 | -0.445 | 20.0 | 10.0 | 0.0 | 50.1052210E+00 | DC220MV |  | -0.44499480254999996 | 6.2113947704901006e-08 | 1517192462.135 | 1517192477.587 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499 | -0.44499
| 2 | -0.4 | 0.2 | 0.0 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 49.9885262E+00 | DC220MV |  | -0.39999466128 | 4.8819633341915439e-08 | 1517192492.416 | 1517192507.798 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999 | -0.39999
| 2 | -0.4 | 0.2 | -0.005 | 1 | -0.395 | 20.0 | 10.0 | 0.0 | 50.0510804E+00 | DC220MV |  | -0.39499465459499994 | 7.0722072049701074e-08 | 1517192522.647 | 1517192538.077 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499 | -0.39499
| 2 | -0.35 | 0.2 | 0.0 | 1 | -0.35 | 20.0 | 10.0 | 0.0 | 50.0477862E+00 | DC220MV |  | -0.34999463465500003 | 5.4422894770638513e-08 | 1517192552.898 | 1517192568.321 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999 | -0.34999
| 2 | -0.35 | 0.2 | -0.005 | 1 | -0.345 | 20.0 | 10.0 | 0.0 | 49.9986085E+00 | DC220MV |  | -0.34499453608500003 | 5.3794993027379533e-08 | 1517192583.154 | 1517192598.578 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499 | -0.34499
| 2 | -0.3 | 0.2 | 0.0 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9857871E+00 | DC220MV |  | -0.29999444241499995 | 4.5282637680374731e-08 | 1517192613.439 | 1517192628.881 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999 | -0.29999
| 2 | -0.3 | 0.2 | -0.005 | 1 | -0.295 | 20.0 | 10.0 | 0.0 | 50.0387517E+00 | DC220MV |  | -0.29499434264000002 | 6.8599674921897465e-08 | 1517192643.705 | 1517192659.155 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499 | -0.29499
| 2 | -0.25 | 0.2 | 0.0 | 1 | -0.25 | 20.0 | 10.0 | 0.0 | 50.0241579E+00 | DC220MV |  | -0.24999447074 | 9.3952442223422511e-08 | 1517192674.014 | 1517192689.435 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999 | -0.24999
| 2 | -0.25 | 0.2 | -0.005 | 1 | -0.245 | 20.0 | 10.0 | 0.0 | 49.9379874E+00 | DC220MV |  | -0.24499419521000002 | 6.6517636005837715e-08 | 1517192704.284 | 1517192719.707 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499 | -0.24499
| 2 | -0.2 | 0.2 | 0.0 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 50.0142918E+00 | DC220MV |  | -0.19999411334999997 | 7.7333611711777391e-08 | 1517192734.571 | 1517192750.044 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999 | -0.19999
| 2 | -0.2 | 0.2 | -0.005 | 1 | -0.195 | 20.0 | 10.0 | 0.0 | 50.0271987E+00 | DC220MV |  | -0.19499414597500001 | 6.4134053941366587e-08 | 1517192764.885 | 1517192780.314 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499 | -0.19499
| 2 | -0.15 | 0.2 | 0.0 | 1 | -0.15 | 20.0 | 10.0 | 0.0 | 49.9835762E+00 | DC220MV |  | -0.14999407028 | 6.5661225999883249e-08 | 1517192795.141 | 1517192810.562 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999 | -0.14999
| 2 | -0.15 | 0.2 | -0.005 | 1 | -0.145 | 20.0 | 10.0 | 0.0 | 49.9978774E+00 | DC220MV |  | -0.14499406993500003 | 7.702382277593069e-08 | 1517192825.392 | 1517192840.84 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499 | -0.14499
| 2 | -0.1 | 0.2 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9912327E+00 | DC220MV |  | -0.099993909155499994 | 5.5408056586820634e-08 | 1517192855.682 | 1517192871.121 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994 | -0.099994
| 2 | -0.1 | 0.2 | -0.005 | 1 | -0.095 | 20.0 | 10.0 | 0.0 | 49.9671725E+00 | DC220MV |  | -0.094993833390499993 | 5.7503943644521797e-08 | 1517192885.954 | 1517192901.408 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994 | -0.094994
| 2 | -0.05 | 0.2 | 0.0 | 1 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9822280E+00 | DC220MV |  | -0.049993738377000004 | 5.9098348039059696e-08 | 1517192916.237 | 1517192931.68 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994
| 2 | -0.05 | 0.2 | -0.005 | 1 | -0.04500000000000001 | 20.0 | 10.0 | 0.0 | 50.0318788E+00 | DC220MV |  | -0.044993705325499997 | 8.51504742488874e-08 | 1517192946.524 | 1517192961.975 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9967644E+00 | DC220MV |  | 6.21315493265e-06 | 4.5634434893559405e-08 | 1517192976.794 | 1517192992.261 | 6.209e-06 | 6.1942e-06 | 6.1751e-06 | 6.2016e-06 | 6.2458e-06 | 6.1854e-06 | 6.1677e-06 | 6.2988e-06 | 6.2575e-06 | 6.1133e-06 | 6.2281e-06 | 6.234e-06 | 6.1928e-06 | 6.2031e-06 | 6.2001e-06 | 6.2472e-06 | 6.1692e-06 | 6.2384e-06 | 6.1854e-06 | 6.3164e-06
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | -0.005 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9889034E+00 | DC220MV |  | -0.044993790970999994 | 5.8764602006692297e-08 | 1517193007.117 | 1517193022.555 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994 | -0.044994
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9664423E+00 | DC220MV |  | 6.2680719859999991e-06 | 4.8321132378350092e-08 | 1517193037.41 | 1517193052.903 | 6.3674e-06 | 6.188e-06 | 6.2924e-06 | 6.2909e-06 | 6.2851e-06 | 6.2262e-06 | 6.2321e-06 | 6.2777e-06 | 6.2351e-06 | 6.2556e-06 | 6.2086e-06 | 6.263e-06 | 6.2865e-06 | 6.2924e-06 | 6.3468e-06 | 6.2895e-06 | 6.1792e-06 | 6.3189e-06 | 6.2248e-06 | 6.3012e-06
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | 0.0 | 1 | -0.005 | 20.0 | 10.0 | 0.0 | 49.9836704E+00 | DC220MV |  | -0.049993783317499998 | 6.2853730587814132e-08 | 1517193067.736 | 1517193083.204 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994 | -0.049994
| 2 | 0.1 | 2.0 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0085109E+00 | DC2_2V |  | 0.10000618191999999 | 5.8647451778118781e-08 | 1517193098.347 | 1517193113.78 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0626870E+00 | DC2_2V |  | 6.1337276263500019e-06 | 5.76296716569828e-08 | 1517193128.611 | 1517193144.067 | 6.2161e-06 | 6.1425e-06 | 6.169e-06 | 6.1543e-06 | 6.0792e-06 | 6.0703e-06 | 6.1057e-06 | 6.0865e-06 | 6.1027e-06 | 6.007e-06 | 6.2206e-06 | 6.0968e-06 | 6.2029e-06 | 6.2147e-06 | 6.0792e-06 | 6.1381e-06 | 6.1734e-06 | 6.1911e-06 | 6.1425e-06 | 6.0821e-06
| 2 | 0.1 | 2.0 | 0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0257916E+00 | DC2_2V |  | 5.6964616189000012e-06 | 6.2214699012116156e-08 | 1517193158.9 | 1517193174.351 | 5.6602e-06 | 5.7884e-06 | 5.6867e-06 | 5.6617e-06 | 5.644e-06 | 5.7486e-06 | 5.7221e-06 | 5.6219e-06 | 5.6985e-06 | 5.6971e-06 | 5.7221e-06 | 5.7884e-06 | 5.6425e-06 | 5.6013e-06 | 5.588e-06 | 5.7177e-06 | 5.809e-06 | 5.7589e-06 | 5.6381e-06 | 5.7339e-06
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9503042E+00 | DC2_2V |  | 6.1250143284000004e-06 | 5.0980923048492871e-08 | 1517193189.181 | 1517193204.644 | 6.1128e-06 | 6.1864e-06 | 6.1202e-06 | 6.0671e-06 | 6.1349e-06 | 6.0539e-06 | 6.1275e-06 | 6.182e-06 | 6.0318e-06 | 6.1511e-06 | 6.0524e-06 | 6.2115e-06 | 6.1231e-06 | 6.1584e-06 | 6.1967e-06 | 6.1452e-06 | 6.1246e-06 | 6.1599e-06 | 6.1143e-06 | 6.0465e-06
| 2 | 0.1 | 2.0 | 0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9943921E+00 | DC2_2V |  | 5.7014249555000003e-06 | 5.1379761440889908e-08 | 1517193219.478 | 1517193234.977 | 5.6602e-06 | 5.6823e-06 | 5.7573e-06 | 5.7117e-06 | 5.7705e-06 | 5.7088e-06 | 5.6779e-06 | 5.7014e-06 | 5.6602e-06 | 5.7455e-06 | 5.6558e-06 | 5.7882e-06 | 5.7352e-06 | 5.7044e-06 | 5.775e-06 | 5.5867e-06 | 5.6147e-06 | 5.7058e-06 | 5.6661e-06 | 5.7205e-06
| 2 | 0.1 | 2.0 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0125503E+00 | DC2_2V |  | 0.10000626964000001 | 6.5049192154355619e-08 | 1517193249.82 | 1517193265.238 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001
| 2 | 0.2 | 2.0 | 0.1 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0430701E+00 | DC2_2V |  | 0.10000601687499999 | 5.8157337241805593e-08 | 1517193280.088 | 1517193295.497 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001 | 0.10001
| 2 | 0.2 | 2.0 | 0.0 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0012925E+00 | DC2_2V |  | 0.20000649248500002 | 7.2725925742494488e-08 | 1517193310.338 | 1517193325.748 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001
| 2 | 0.3 | 2.0 | 0.1 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0043211E+00 | DC2_2V |  | 0.20000608103 | 7.4733922016883465e-08 | 1517193340.589 | 1517193356.018 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001
| 2 | 0.3 | 2.0 | 0.0 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9643226E+00 | DC2_2V |  | 0.300006646155 | 4.963417647888035e-08 | 1517193370.872 | 1517193386.287 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001
| 2 | 0.4 | 2.0 | 0.1 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9901953E+00 | DC2_2V |  | 0.30000625324500002 | 6.064340421632015e-08 | 1517193401.113 | 1517193416.558 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001
| 2 | 0.4 | 2.0 | 0.0 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9916193E+00 | DC2_2V |  | 0.40000673571500006 | 7.0175838257975796e-08 | 1517193431.384 | 1517193446.815 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001
| 2 | 0.5 | 2.0 | 0.1 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0465291E+00 | DC2_2V |  | 0.40000618057000004 | 6.156806072050186e-08 | 1517193461.661 | 1517193477.108 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001
| 2 | 0.5 | 2.0 | 0.0 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 49.9726518E+00 | DC2_2V |  | 0.50000681091999988 | 7.9581226432972394e-08 | 1517193491.947 | 1517193507.349 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001
| 2 | 0.6 | 2.0 | 0.1 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0591454E+00 | DC2_2V |  | 0.50000638788499996 | 6.2771237657903552e-08 | 1517193522.204 | 1517193537.645 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001
| 2 | 0.6 | 2.0 | 0.0 | 1 | 0.6 | 20.0 | 10.0 | 0.0 | 49.9304883E+00 | DC2_2V |  | 0.60000687648000006 | 4.2213914753199533e-08 | 1517193552.509 | 1517193567.912 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001
| 2 | 0.7 | 2.0 | 0.1 | 1 | 0.6 | 20.0 | 10.0 | 0.0 | 50.0458864E+00 | DC2_2V |  | 0.60000642935999993 | 8.4094603873812492e-08 | 1517193582.747 | 1517193598.205 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001
| 2 | 0.7 | 2.0 | 0.0 | 1 | 0.7 | 20.0 | 10.0 | 0.0 | 49.9816860E+00 | DC2_2V |  | 0.70000705999000001 | 4.8871913192493972e-08 | 1517193613.058 | 1517193628.456 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001
| 2 | 0.8 | 2.0 | 0.1 | 1 | 0.7000000000000001 | 20.0 | 10.0 | 0.0 | 49.9995330E+00 | DC2_2V |  | 0.70000662087499999 | 8.5327960688895968e-08 | 1517193643.31 | 1517193658.754 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001
| 2 | 0.8 | 2.0 | 0.0 | 1 | 0.8 | 20.0 | 10.0 | 0.0 | 49.9780238E+00 | DC2_2V |  | 0.80000731885499987 | 7.4015508336085674e-08 | 1517193673.584 | 1517193689.002 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001
| 2 | 0.9 | 2.0 | 0.1 | 1 | 0.8 | 20.0 | 10.0 | 0.0 | 49.9840711E+00 | DC2_2V |  | 0.80000689773500011 | 7.9475545136955374e-08 | 1517193703.865 | 1517193719.297 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001
| 2 | 0.9 | 2.0 | 0.0 | 1 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0234827E+00 | DC2_2V |  | 0.90000749094999988 | 7.7979327386033302e-08 | 1517193734.138 | 1517193749.577 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001
| 2 | 1.0 | 2.0 | 0.1 | 1 | 0.9 | 20.0 | 10.0 | 0.0 | 49.9729109E+00 | DC2_2V |  | 0.90000709997499995 | 8.2610113637535107e-08 | 1517193764.403 | 1517193779.826 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001
| 2 | 1.0 | 2.0 | 0.0 | 1 | 1.0 | 20.0 | 10.0 | 0.0 | 50.0320630E+00 | DC2_2V |  | 1.0000076757999998 | 6.7134640829421785e-08 | 1517193794.674 | 1517193810.095 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 2 | 1.0 | 2.0 | 0.1 | 1 | 0.9 | 20.0 | 10.0 | 0.0 | 49.9429425E+00 | DC2_2V |  | 0.90000711099000008 | 7.9532785068457113e-08 | 1517193824.931 | 1517193840.332 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001
| 2 | 0.9 | 2.0 | 0.0 | 1 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0520541E+00 | DC2_2V |  | 0.90000756413999983 | 7.4514001362506257e-08 | 1517193855.173 | 1517193870.636 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001 | 0.90001
| 2 | 0.9 | 2.0 | 0.1 | 1 | 0.8 | 20.0 | 10.0 | 0.0 | 50.0088837E+00 | DC2_2V |  | 0.80000690922499995 | 7.1503292756136441e-08 | 1517193885.491 | 1517193900.9 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001
| 2 | 0.8 | 2.0 | 0.0 | 1 | 0.8 | 20.0 | 10.0 | 0.0 | 50.0087043E+00 | DC2_2V |  | 0.80000687685000005 | 7.5193886047722566e-08 | 1517193915.746 | 1517193931.162 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001 | 0.80001
| 2 | 0.8 | 2.0 | 0.1 | 1 | 0.7000000000000001 | 20.0 | 10.0 | 0.0 | 49.9525827E+00 | DC2_2V |  | 0.70000569485499997 | 1.9105636726384761e-07 | 1517193946.013 | 1517193961.436 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001
| 2 | 0.7 | 2.0 | 0.0 | 1 | 0.7 | 20.0 | 10.0 | 0.0 | 50.0034625E+00 | DC2_2V |  | 0.70000583943499994 | 7.5144010227459737e-08 | 1517193976.288 | 1517193991.733 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001 | 0.70001
| 2 | 0.7 | 2.0 | 0.1 | 1 | 0.6 | 20.0 | 10.0 | 0.0 | 50.0273499E+00 | DC2_2V |  | 0.60000511863500017 | 7.1776495287555671e-08 | 1517194006.556 | 1517194021.992 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.6 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.6
| 2 | 0.6 | 2.0 | 0.0 | 1 | 0.6 | 20.0 | 10.0 | 0.0 | 49.9947506E+00 | DC2_2V |  | 0.600005496365 | 7.9728973877969257e-08 | 1517194036.835 | 1517194052.267 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001 | 0.60001
| 2 | 0.6 | 2.0 | 0.1 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 49.9879416E+00 | DC2_2V |  | 0.50000482199999996 | 6.9082233608157977e-08 | 1517194067.113 | 1517194082.554 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 2 | 0.5 | 2.0 | 0.0 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0114083E+00 | DC2_2V |  | 0.5000052600649999 | 7.1296123850391156e-08 | 1517194097.398 | 1517194112.833 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001 | 0.50001
| 2 | 0.5 | 2.0 | 0.1 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9483223E+00 | DC2_2V |  | 0.40000498214500002 | 1.1856459199644856e-07 | 1517194127.67 | 1517194143.098 | 0.4 | 0.4 | 0.40001 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.40001 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.40001 | 0.40001 | 0.40001 | 0.4
| 2 | 0.4 | 2.0 | 0.0 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9963588E+00 | DC2_2V |  | 0.40000539338499996 | 7.4072959134415157e-08 | 1517194157.957 | 1517194173.411 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001 | 0.40001
| 2 | 0.4 | 2.0 | 0.1 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0023020E+00 | DC2_2V |  | 0.30000488213499998 | 6.3060068786389986e-08 | 1517194188.252 | 1517194203.681 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.30001 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.3 | 2.0 | 0.0 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0022076E+00 | DC2_2V |  | 0.30000531460500002 | 8.1987410463264616e-08 | 1517194218.548 | 1517194233.967 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001 | 0.30001
| 2 | 0.3 | 2.0 | 0.1 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9701593E+00 | DC2_2V |  | 0.20000472935000002 | 5.3739431521871885e-08 | 1517194248.82 | 1517194264.252 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.2 | 2.0 | 0.0 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0561715E+00 | DC2_2V |  | 0.20000516369500004 | 8.1538429435806438e-08 | 1517194279.098 | 1517194294.537 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.20001 | 0.2 | 0.20001
| 2 | 0.2 | 2.0 | 0.1 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9908885E+00 | DC2_2V |  | 0.10000452136 | 5.3635588930876675e-08 | 1517194309.382 | 1517194324.786 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.1 | 2.0 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0190875E+00 | DC2_2V |  | 0.10000496146 | 9.1958036081773415e-08 | 1517194339.621 | 1517194355.054 | 0.1 | 0.10001 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10001 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10001 | 0.1 | 0.10001 | 0.10001 | 0.10001
| 2 | 0.1 | 2.0 | 0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0109835E+00 | DC2_2V |  | 4.5590561719499996e-06 | 5.9942336550090813e-08 | 1517194369.892 | 1517194385.368 | 4.414e-06 | 4.5244e-06 | 4.5862e-06 | 4.5892e-06 | 4.5465e-06 | 4.6098e-06 | 4.545e-06 | 4.5332e-06 | 4.5597e-06 | 4.5097e-06 | 4.6863e-06 | 4.4655e-06 | 4.5347e-06 | 4.5362e-06 | 4.5362e-06 | 4.6657e-06 | 4.5553e-06 | 4.5833e-06 | 4.6024e-06 | 4.598e-06
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0002972E+00 | DC2_2V |  | 4.9759527954999996e-06 | 6.3457428468268071e-08 | 1517194400.218 | 1517194415.701 | 4.9364e-06 | 4.982e-06 | 5.0188e-06 | 4.9187e-06 | 5.0261e-06 | 4.9305e-06 | 5.0614e-06 | 5.0482e-06 | 5.0423e-06 | 4.9187e-06 | 5.0909e-06 | 4.8716e-06 | 4.9349e-06 | 4.876e-06 | 4.9011e-06 | 5.0261e-06 | 5.0144e-06 | 5.0144e-06 | 4.929e-06 | 4.9776e-06
| 2 | 0.1 | 2.0 | 0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0171757E+00 | DC2_2V |  | 4.3751512309999999e-06 | 7.5323213424860368e-08 | 1517194430.531 | 1517194445.993 | 4.4709e-06 | 4.4017e-06 | 4.4783e-06 | 4.2707e-06 | 4.4474e-06 | 4.4621e-06 | 4.3723e-06 | 4.3973e-06 | 4.3973e-06 | 4.3561e-06 | 4.468e-06 | 4.2471e-06 | 4.3414e-06 | 4.3664e-06 | 4.3546e-06 | 4.2677e-06 | 4.2501e-06 | 4.2913e-06 | 4.412e-06 | 4.4503e-06
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9954485E+00 | DC2_2V |  | 4.8039370510000005e-06 | 8.2371552378997104e-08 | 1517194460.849 | 1517194476.311 | 4.9863e-06 | 4.7507e-06 | 4.8214e-06 | 4.8877e-06 | 4.8479e-06 | 4.8803e-06 | 4.733e-06 | 4.8877e-06 | 4.7846e-06 | 4.8832e-06 | 4.7507e-06 | 4.7433e-06 | 4.7168e-06 | 4.8803e-06 | 4.7625e-06 | 4.6403e-06 | 4.7919e-06 | 4.7389e-06 | 4.7154e-06 | 4.8759e-06
| 2 | 0.1 | 2.0 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0158871E+00 | DC2_2V |  | 0.10000488734000001 | 5.9023973095946374e-08 | 1517194491.161 | 1517194506.601 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | -0.1 | 2.0 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 50.0053119E+00 | DC2_2V |  | -0.099995281979999989 | 6.6598230155776134e-08 | 1517194521.449 | 1517194536.893 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9785894E+00 | DC2_2V |  | 4.7152948178999993e-06 | 6.7624615421587493e-08 | 1517194551.715 | 1517194567.186 | 4.6226e-06 | 4.7713e-06 | 4.6064e-06 | 4.6903e-06 | 4.5976e-06 | 4.7286e-06 | 4.8184e-06 | 4.7581e-06 | 4.6432e-06 | 4.764e-06 | 4.7551e-06 | 4.7168e-06 | 4.7168e-06 | 4.7743e-06 | 4.7743e-06 | 4.6815e-06 | 4.7801e-06 | 4.5902e-06 | 4.761e-06 | 4.7551e-06
| 2 | -0.1 | 2.0 | -0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0341979E+00 | DC2_2V |  | 5.2797885426999999e-06 | 7.6145054641721963e-08 | 1517194582.012 | 1517194597.501 | 5.3344e-06 | 5.3124e-06 | 5.4757e-06 | 5.2579e-06 | 5.3197e-06 | 5.2182e-06 | 5.2388e-06 | 5.2521e-06 | 5.3138e-06 | 5.38e-06 | 5.2373e-06 | 5.3389e-06 | 5.3124e-06 | 5.1741e-06 | 5.2006e-06 | 5.2256e-06 | 5.1976e-06 | 5.3845e-06 | 5.2167e-06 | 5.205e-06
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9806254E+00 | DC2_2V |  | 4.7071551711499999e-06 | 5.9646936581040287e-08 | 1517194612.335 | 1517194627.791 | 4.7152e-06 | 4.6872e-06 | 4.6209e-06 | 4.7343e-06 | 4.6563e-06 | 4.7137e-06 | 4.7181e-06 | 4.668e-06 | 4.6548e-06 | 4.755e-06 | 4.6739e-06 | 4.8478e-06 | 4.7476e-06 | 4.696e-06 | 4.7019e-06 | 4.7918e-06 | 4.6622e-06 | 4.5885e-06 | 4.7918e-06 | 4.7181e-06
| 2 | -0.1 | 2.0 | -0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9898983E+00 | DC2_2V |  | 5.1989188054499996e-06 | 5.8176734471302751e-08 | 1517194642.626 | 1517194658.109 | 5.2754e-06 | 5.1812e-06 | 5.1385e-06 | 5.2989e-06 | 5.2695e-06 | 5.2533e-06 | 5.1827e-06 | 5.112e-06 | 5.2224e-06 | 5.1606e-06 | 5.2974e-06 | 5.1606e-06 | 5.1797e-06 | 5.1253e-06 | 5.1341e-06 | 5.1856e-06 | 5.1532e-06 | 5.1562e-06 | 5.2312e-06 | 5.2606e-06
| 2 | -0.1 | 2.0 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 50.0064348E+00 | DC2_2V |  | -0.099995319090000007 | 8.8411310532624663e-08 | 1517194672.952 | 1517194688.397 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099996 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995
| 2 | -0.2 | 2.0 | -0.1 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9689014E+00 | DC2_2V |  | -0.099994872566999987 | 5.3493694497591124e-08 | 1517194703.233 | 1517194718.686 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995 | -0.099995
| 2 | -0.2 | 2.0 | 0.0 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 50.0208389E+00 | DC2_2V |  | -0.19999552871499998 | 6.6642451000153211e-08 | 1517194733.528 | 1517194748.987 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 2 | -0.3 | 2.0 | -0.1 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9611999E+00 | DC2_2V |  | -0.19999505812500001 | 9.5800223774893413e-08 | 1517194763.84 | 1517194779.261 | -0.2 | -0.2 | -0.19999 | -0.2 | -0.19999 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.19999 | -0.2 | -0.2 | -0.2 | -0.19999 | -0.2 | -0.2 | -0.19999 | -0.2
| 2 | -0.3 | 2.0 | 0.0 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9625093E+00 | DC2_2V |  | -0.29999571644500006 | 5.406398500428807e-08 | 1517194794.107 | 1517194809.572 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 2 | -0.4 | 2.0 | -0.1 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0320205E+00 | DC2_2V |  | -0.29999554469999995 | 6.594452213894555e-08 | 1517194824.41 | 1517194839.86 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 2 | -0.4 | 2.0 | 0.0 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 49.9304836E+00 | DC2_2V |  | -0.39999626104499997 | 7.2059499540919813e-08 | 1517194854.698 | 1517194870.131 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 2 | -0.5 | 2.0 | -0.1 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 49.9726659E+00 | DC2_2V |  | -0.39999581654500005 | 5.7234600333346674e-08 | 1517194884.997 | 1517194900.47 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 2 | -0.5 | 2.0 | 0.0 | 1 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9914920E+00 | DC2_2V |  | -0.49999642525999999 | 6.9231866942213561e-08 | 1517194915.295 | 1517194930.742 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 2 | -0.6 | 2.0 | -0.1 | 1 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9980048E+00 | DC2_2V |  | -0.49999593036000001 | 6.306550086838673e-08 | 1517194945.592 | 1517194961.029 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 2 | -0.6 | 2.0 | 0.0 | 1 | -0.6 | 20.0 | 10.0 | 0.0 | 50.0020944E+00 | DC2_2V |  | -0.59999647815500001 | 8.1454972074246608e-08 | 1517194975.883 | 1517194991.325 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 2 | -0.7 | 2.0 | -0.1 | 1 | -0.6 | 20.0 | 10.0 | 0.0 | 49.9806443E+00 | DC2_2V |  | -0.59999603410999991 | 8.1109709034275247e-08 | 1517195006.181 | 1517195021.621 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 2 | -0.7 | 2.0 | 0.0 | 1 | -0.7 | 20.0 | 10.0 | 0.0 | 50.0585874E+00 | DC2_2V |  | -0.69999659948500004 | 7.2521102269470275e-08 | 1517195036.49 | 1517195051.939 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 2 | -0.8 | 2.0 | -0.1 | 1 | -0.7000000000000001 | 20.0 | 10.0 | 0.0 | 49.9589677E+00 | DC2_2V |  | -0.69999637899 | 6.7463078060021276e-08 | 1517195066.766 | 1517195082.179 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 2 | -0.8 | 2.0 | 0.0 | 1 | -0.8 | 20.0 | 10.0 | 0.0 | 49.9962408E+00 | DC2_2V |  | -0.79999706482499999 | 5.9624289310950938e-08 | 1517195097.028 | 1517195112.488 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 2 | -0.9 | 2.0 | -0.1 | 1 | -0.8 | 20.0 | 10.0 | 0.0 | 50.0133385E+00 | DC2_2V |  | -0.79999649655000005 | 6.7627838207219301e-08 | 1517195127.339 | 1517195142.783 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 2 | -0.9 | 2.0 | 0.0 | 1 | -0.9 | 20.0 | 10.0 | 0.0 | 50.0238179E+00 | DC2_2V |  | -0.89999707284499997 | 4.8213664825304702e-08 | 1517195157.633 | 1517195173.074 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 2 | -1.0 | 2.0 | -0.1 | 1 | -0.9 | 20.0 | 10.0 | 0.0 | 50.0164441E+00 | DC2_2V |  | -0.89999675761500009 | 8.4291484000924727e-08 | 1517195187.925 | 1517195203.366 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 2 | -1.0 | 2.0 | 0.0 | 1 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9722089E+00 | DC2_2V |  | -0.99999739764999995 | 5.5165075000376293e-08 | 1517195218.227 | 1517195233.658 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 2 | -1.0 | 2.0 | -0.1 | 1 | -0.9 | 20.0 | 10.0 | 0.0 | 49.9862161E+00 | DC2_2V |  | -0.89999678534999994 | 5.974670284048037e-08 | 1517195248.507 | 1517195263.964 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 2 | -0.9 | 2.0 | 0.0 | 1 | -0.9 | 20.0 | 10.0 | 0.0 | 50.0197720E+00 | DC2_2V |  | -0.89999717381000011 | 6.6459844261444095e-08 | 1517195278.838 | 1517195294.298 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1517195297.004 | 1517195297.004