## Dict
label | Keysight | label | FL5520 | label | FL5700
NoError | 0\r\n | NoError | 0\n | NoError | 0\n
reset | RESET | reset | *RST | reset | *RST
init | END 2; DCV AUTO; AZERO ON;DELAY 0;NPLC 20 | init | OUT 0V, 0HZ | init | OUT 0V, 0HZ
Make_Safe | DCV AUTO | Make_Safe | OUT 0V,0HZ;STBY | Make_Safe | OUT 0V,0HZ;STBY
error | ERR? | error | FAULT? | error | FAULT?
SettleTime | 0.0 | SettleTime | 0.3 | SettleTime | 0.3
measure_seperation | 0 | operate | OPER | operate | OPER
DCVRange | DCV $ | standby | STBY | standby | STBY
MeasureSetup | LFREQ LINE | DCVRange | RANGELCK OFF; OUT $V; RANGELCK ON | DCVRange | RANGELCK OFF; OUT $V; RANGELCK ON
status | LINE? | status | RANGE? | SetVoltage | OUT $V
SingleMsmntSetup | SKIP | SetVoltage | OUT $V | status | RANGE?
Ranges | (0,0.1,'0.1'),(0,1,'1'),(0,10,'10'),(0,100,'100'),(0,1000,'1000') | Ranges | (0,0.22,'0.2'),(0,2.2,'2'),(0,10,'10'),(0,22,'20'),(0,200,'220'),(0,200,'220'),(100,1100,'1000') | Ranges | (0,0.22,'0.2'),(0,2.2,'2'),(0,10,'10'),(0,22,'20'),(0,200,'220'),(0,200,'220'),(100,1100,'1000')

## Control
instruments: | Meter: HP | S: FL5700 | X: FL5730
Start Row | 8 | Stop Row | 1008
| X Range | X Settings (V) | S Range | S Settings (V) | DVM Range | Nominal reading | #Readings | Delay (S) | DVM pause | DVM status | S status | X status | Mean | STD | start time | end time | readings...
| 0.2 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 49.9921616E+00 | DC220MV | DC220MV | 0.010005 | 5.2663e-08
| 0.1 | 0.0001 | 0.2 | 0 | 0.1 | 0.0001 | 20 | 10 | 0 | 50.0131308E+00 | DC220MV | DC220MV | 0.010005 | 5.2663e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9719827E+00 | DC220MV | DC220MV | 5.6183e-06 | 3.6452e-08
| 0.1 | 0.0001 | 0.2 | 0.0001 | 0.1 | 0 | 20 | 10 | 0 | 50.0056421E+00 | DC220MV | DC220MV | 5.1692e-06 | 4.3858e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9625988E+00 | DC220MV | DC220MV | 5.5943e-06 | 3.7356e-08
| 0.1 | 0.0001 | 0.2 | 0.0001 | 0.1 | 0 | 20 | 10 | 0 | 50.0264432E+00 | DC220MV | DC220MV | 5.2605e-06 | 3.5703e-08
| 0.1 | 0.0001 | 0.2 | 0 | 0.1 | 0.0001 | 20 | 10 | 0 | 50.0096339E+00 | DC220MV | DC220MV | 0.010005 | 3.5711e-08
| 0.1 | 0.0002 | 0.2 | 0.0001 | 0.1 | 0.0001 | 20 | 10 | 0 | 49.9681288E+00 | DC220MV | DC220MV | 0.010005 | 3.7356e-08
| 0.1 | 0.0002 | 0.2 | 0 | 0.1 | 0.0002 | 20 | 10 | 0 | 49.9975426E+00 | DC220MV | DC220MV | 0.020005 | 4.5103e-08
| 0.1 | 0.0003 | 0.2 | 0.0001 | 0.1 | 0.0002 | 20 | 10 | 0 | 50.0329037E+00 | DC220MV | DC220MV | 0.020005 | 5.1348e-08
| 0.1 | 0.0003 | 0.2 | 0 | 0.1 | 0.0003 | 20 | 10 | 0 | 50.0029530E+00 | DC220MV | DC220MV | 0.030005 | 4.6575e-08
| 0.1 | 0.0004 | 0.2 | 0.0001 | 0.1 | 0.0003 | 20 | 10 | 0 | 49.9937083E+00 | DC220MV | DC220MV | 0.030004 | 3.8575e-08
| 0.1 | 0.0004 | 0.2 | 0 | 0.1 | 0.0004 | 20 | 10 | 0 | 50.0149054E+00 | DC220MV | DC220MV | 0.040004 | 4.146e-08
| 0.1 | 0.0005 | 0.2 | 0.0001 | 0.1 | 0.0004 | 20 | 10 | 0 | 49.9931330E+00 | DC220MV | DC220MV | 0.040004 | 5.3093e-08
| 0.1 | 0.0005 | 0.2 | 0 | 0.1 | 0.0005 | 20 | 10 | 0 | 49.9917466E+00 | DC220MV | DC220MV | 0.050004 | 4.4105e-08
| 0.1 | 0.0006 | 0.2 | 0.0001 | 0.1 | 0.0005 | 20 | 10 | 0 | 50.0594385E+00 | DC220MV | DC220MV | 0.050003 | 4.989e-08
| 0.1 | 0.0006 | 0.2 | 0 | 0.1 | 0.0006 | 20 | 10 | 0 | 49.9829162E+00 | DC220MV | DC220MV | 0.060003 | 4.2866e-08
| 0.1 | 0.0007 | 0.2 | 0.0001 | 0.1 | 0.0006 | 20 | 10 | 0 | 49.9492356E+00 | DC220MV | DC220MV | 0.060003 | 4.4228e-08
| 0.1 | 0.0007 | 0.2 | 0 | 0.1 | 0.0007 | 20 | 10 | 0 | 50.0236433E+00 | DC220MV | DC220MV | 0.070003 | 7.9555e-08
| 0.1 | 0.0008 | 0.2 | 0.0001 | 0.1 | 0.0007 | 20 | 10 | 0 | 49.9746355E+00 | DC220MV | DC220MV | 0.070002 | 4.1837e-08
| 0.1 | 0.0008 | 0.2 | 0 | 0.1 | 0.0008 | 20 | 10 | 0 | 49.9932698E+00 | DC220MV | DC220MV | 0.080002 | 7.1368e-08
| 0.1 | 0.0009 | 0.2 | 0.0001 | 0.1 | 0.0008 | 20 | 10 | 0 | 49.9617510E+00 | DC220MV | DC220MV | 0.080002 | 5.1504e-08
| 0.1 | 0.0009 | 0.2 | 0 | 0.1 | 0.0009 | 20 | 10 | 0 | 49.9962880E+00 | DC220MV | DC220MV | 0.090002 | 4.66e-08
| 0.1 | 0.001 | 0.2 | 0.0001 | 0.1 | 0.0009 | 20 | 10 | 0 | 49.9896012E+00 | DC220MV | DC220MV | 0.090002 | 3.2379e-08
| 0.1 | 0.001 | 0.2 | 0 | 0.1 | 0.001 | 20 | 10 | 0 | 50.0031417E+00 | DC220MV | DC220MV | 0.1 | 5.4103e-08
| 0.1 | 0.001 | 0.2 | 0.0001 | 0.1 | 0.0009 | 20 | 10 | 0 | 49.9789805E+00 | DC220MV | DC220MV | 0.090002 | 5.163e-08
| 0.1 | 0.0009 | 0.2 | 0 | 0.1 | 0.0009 | 20 | 10 | 0 | 50.0171238E+00 | DC220MV | DC220MV | 0.090002 | 5.37e-07
| 0.1 | 0.0009 | 0.2 | 0.0001 | 0.1 | 0.0008 | 20 | 10 | 0 | 49.9998915E+00 | DC220MV | DC220MV | 0.080002 | 4.9294e-08
| 0.1 | 0.0008 | 0.2 | 0 | 0.1 | 0.0008 | 20 | 10 | 0 | 49.9831943E+00 | DC220MV | DC220MV | 0.080003 | 4.2433e-08
| 0.1 | 0.0008 | 0.2 | 0.0001 | 0.1 | 0.0007 | 20 | 10 | 0 | 49.9900492E+00 | DC220MV | DC220MV | 0.070003 | 5.5564e-08
| 0.1 | 0.0007 | 0.2 | 0 | 0.1 | 0.0007 | 20 | 10 | 0 | 49.9985614E+00 | DC220MV | DC220MV | 0.070003 | 4.4479e-08
| 0.1 | 0.0007 | 0.2 | 0.0001 | 0.1 | 0.0006 | 20 | 10 | 0 | 50.0033021E+00 | DC220MV | DC220MV | 0.060003 | 3.8506e-08
| 0.1 | 0.0006 | 0.2 | 0 | 0.1 | 0.0006 | 20 | 10 | 0 | 49.9917985E+00 | DC220MV | DC220MV | 0.060004 | 5.6544e-08
| 0.1 | 0.0006 | 0.2 | 0.0001 | 0.1 | 0.0005 | 20 | 10 | 0 | 49.9843775E+00 | DC220MV | DC220MV | 0.050004 | 4.5149e-08
| 0.1 | 0.0005 | 0.2 | 0 | 0.1 | 0.0005 | 20 | 10 | 0 | 50.0249511E+00 | DC220MV | DC220MV | 0.050004 | 5.6065e-08
| 0.1 | 0.0005 | 0.2 | 0.0001 | 0.1 | 0.0004 | 20 | 10 | 0 | 49.9716387E+00 | DC220MV | DC220MV | 0.040004 | 6.4814e-08
| 0.1 | 0.0004 | 0.2 | 0 | 0.1 | 0.0004 | 20 | 10 | 0 | 49.9675352E+00 | DC220MV | DC220MV | 0.040005 | 4.3282e-08
| 0.1 | 0.0004 | 0.2 | 0.0001 | 0.1 | 0.0003 | 20 | 10 | 0 | 50.0280393E+00 | DC220MV | DC220MV | 0.030005 | 3.3781e-08
| 0.1 | 0.0003 | 0.2 | 0 | 0.1 | 0.0003 | 20 | 10 | 0 | 50.0554009E+00 | DC220MV | DC220MV | 0.030005 | 4.149e-08
| 0.1 | 0.0003 | 0.2 | 0.0001 | 0.1 | 0.0002 | 20 | 10 | 0 | 49.9716906E+00 | DC220MV | DC220MV | 0.020005 | 5.0905e-08
| 0.1 | 0.0002 | 0.2 | 0 | 0.1 | 0.0002 | 20 | 10 | 0 | 49.9874135E+00 | DC220MV | DC220MV | 0.020005 | 3.9469e-08
| 0.1 | 0.0002 | 0.2 | 0.0001 | 0.1 | 0.0001 | 20 | 10 | 0 | 49.9689156E+00 | DC220MV | DC220MV | 0.010005 | 3.5281e-08
| 0.1 | 0.0001 | 0.2 | 0 | 0.1 | 0.0001 | 20 | 10 | 0 | 49.9946185E+00 | DC220MV | DC220MV | 0.010006 | 3.0983e-08
| 0.1 | 0.0001 | 0.2 | 0.0001 | 0.1 | 0 | 20 | 10 | 0 | 49.9880877E+00 | DC220MV | DC220MV | 5.4446e-06 | 5.5812e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0087987E+00 | DC220MV | DC220MV | 5.4484e-06 | 8.7916e-08
| 0.1 | 0.0001 | 0.2 | 0.0001 | 0.1 | 0 | 20 | 10 | 0 | 49.9794990E+00 | DC220MV | DC220MV | 4.7549e-06 | 6.0539e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9480446E+00 | DC220MV | DC220MV | 4.8223e-06 | 5.8733e-08
| 0.1 | 0.0001 | 0.2 | 0 | 0.1 | 0.0001 | 20 | 10 | 0 | 50.0089497E+00 | DC220MV | DC220MV | 0.010004 | 4.5554e-08
| 0.1 | -0.0001 | 0.2 | 0 | 0.1 | -0.0001 | 20 | 10 | 0 | 49.9811958E+00 | DC220MV | DC220MV | -0.0099949 | 2.9958e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9946138E+00 | DC220MV | DC220MV | 4.6438e-06 | 6.5219e-08
| 0.1 | -0.0001 | 0.2 | -0.0001 | 0.1 | 0 | 20 | 10 | 0 | 50.0182567E+00 | DC220MV | DC220MV | 5.0585e-06 | 2.973e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0069538E+00 | DC220MV | DC220MV | 4.5116e-06 | 4.2838e-08
| 0.1 | -0.0001 | 0.2 | -0.0001 | 0.1 | 0 | 20 | 10 | 0 | 49.9680582E+00 | DC220MV | DC220MV | 4.9348e-06 | 3.9575e-08
| 0.1 | -0.0001 | 0.2 | 0 | 0.1 | -0.0001 | 20 | 10 | 0 | 50.0065008E+00 | DC220MV | DC220MV | -0.0099951 | 3.8644e-08
| 0.1 | -0.0002 | 0.2 | -0.0001 | 0.1 | -0.0001 | 20 | 10 | 0 | 50.0455414E+00 | DC220MV | DC220MV | -0.0099946 | 6.0552e-08
| 0.1 | -0.0002 | 0.2 | 0 | 0.1 | -0.0002 | 20 | 10 | 0 | 50.0172890E+00 | DC220MV | DC220MV | -0.019995 | 7.5968e-08
| 0.1 | -0.0003 | 0.2 | -0.0001 | 0.1 | -0.0002 | 20 | 10 | 0 | 50.0313923E+00 | DC220MV | DC220MV | -0.019994 | 3.544e-08
| 0.1 | -0.0003 | 0.2 | 0 | 0.1 | -0.0003 | 20 | 10 | 0 | 50.0098180E+00 | DC220MV | DC220MV | -0.029994 | 4.0054e-08
| 0.1 | -0.0004 | 0.2 | -0.0001 | 0.1 | -0.0003 | 20 | 10 | 0 | 49.9871212E+00 | DC220MV | DC220MV | -0.029994 | 3.9743e-08
| 0.1 | -0.0004 | 0.2 | 0 | 0.1 | -0.0004 | 20 | 10 | 0 | 49.9888091E+00 | DC220MV | DC220MV | -0.039994 | 5.294e-08
| 0.1 | -0.0005 | 0.2 | -0.0001 | 0.1 | -0.0004 | 20 | 10 | 0 | 50.0417802E+00 | DC220MV | DC220MV | -0.039993 | 3.5819e-08
| 0.1 | -0.0005 | 0.2 | 0 | 0.1 | -0.0005 | 20 | 10 | 0 | 49.9707247E+00 | DC220MV | DC220MV | -0.049993 | 4.8468e-08
| 0.1 | -0.0006 | 0.2 | -0.0001 | 0.1 | -0.0005 | 20 | 10 | 0 | 49.9903839E+00 | DC220MV | DC220MV | -0.049993 | 3.4955e-08
| 0.1 | -0.0006 | 0.2 | 0 | 0.1 | -0.0006 | 20 | 10 | 0 | 50.0350529E+00 | DC220MV | DC220MV | -0.059993 | 4.5685e-08
| 0.1 | -0.0007 | 0.2 | -0.0001 | 0.1 | -0.0006 | 20 | 10 | 0 | 49.9981086E+00 | DC220MV | DC220MV | -0.059993 | 3.6657e-08
| 0.1 | -0.0007 | 0.2 | 0 | 0.1 | -0.0007 | 20 | 10 | 0 | 49.9734151E+00 | DC220MV | DC220MV | -0.069993 | 5.076e-08
| 0.1 | -0.0008 | 0.2 | -0.0001 | 0.1 | -0.0007 | 20 | 10 | 0 | 50.0092753E+00 | DC220MV | DC220MV | -0.069992 | 4.5716e-08
| 0.1 | -0.0008 | 0.2 | 0 | 0.1 | -0.0008 | 20 | 10 | 0 | 49.9634418E+00 | DC220MV | DC220MV | -0.079993 | 5.4041e-08
| 0.1 | -0.0009 | 0.2 | -0.0001 | 0.1 | -0.0008 | 20 | 10 | 0 | 50.0123144E+00 | DC220MV | DC220MV | -0.079992 | 4.7165e-08
| 0.1 | -0.0009 | 0.2 | 0 | 0.1 | -0.0009 | 20 | 10 | 0 | 49.9921569E+00 | DC220MV | DC220MV | -0.089992 | 4.5394e-08
| 0.1 | -0.001 | 0.2 | -0.0001 | 0.1 | -0.0009 | 20 | 10 | 0 | 49.9707859E+00 | DC220MV | DC220MV | -0.089992 | 2.9758e-08
| 0.1 | -0.001 | 0.2 | 0 | 0.1 | -0.001 | 20 | 10 | 0 | 50.0001604E+00 | DC220MV | DC220MV | -0.099992 | 4.8841e-08
| 0.1 | -0.001 | 0.2 | -0.0001 | 0.1 | -0.0009 | 20 | 10 | 0 | 49.9567874E+00 | DC220MV | DC220MV | -0.089992 | 3.9385e-08
| 0.1 | -0.0009 | 0.2 | 0 | 0.1 | -0.0009 | 20 | 10 | 0 | 50.0067839E+00 | DC220MV | DC220MV | -0.089992 | 7.4541e-08
| 0.1 | -0.0009 | 0.2 | -0.0001 | 0.1 | -0.0008 | 20 | 10 | 0 | 49.9751774E+00 | DC220MV | DC220MV | -0.079992 | 3.2109e-08
| 0.1 | -0.0008 | 0.2 | 0 | 0.1 | -0.0008 | 20 | 10 | 0 | 49.9859945E+00 | DC220MV | DC220MV | -0.079993 | 5.3007e-08
| 0.1 | -0.0008 | 0.2 | -0.0001 | 0.1 | -0.0007 | 20 | 10 | 0 | 50.0301172E+00 | DC220MV | DC220MV | -0.069993 | 4.2062e-08
| 0.1 | -0.0007 | 0.2 | 0 | 0.1 | -0.0007 | 20 | 10 | 0 | 50.0593676E+00 | DC220MV | DC220MV | -0.069993 | 3.84e-08
| 0.1 | -0.0007 | 0.2 | -0.0001 | 0.1 | -0.0006 | 20 | 10 | 0 | 50.0273924E+00 | DC220MV | DC220MV | -0.059993 | 4.2363e-08
| 0.1 | -0.0006 | 0.2 | 0 | 0.1 | -0.0006 | 20 | 10 | 0 | 50.0151697E+00 | DC220MV | DC220MV | -0.059993 | 4.3653e-08
| 0.1 | -0.0006 | 0.2 | -0.0001 | 0.1 | -0.0005 | 20 | 10 | 0 | 49.9966182E+00 | DC220MV | DC220MV | -0.049993 | 6.0246e-08
| 0.1 | -0.0005 | 0.2 | 0 | 0.1 | -0.0005 | 20 | 10 | 0 | 49.9998443E+00 | DC220MV | DC220MV | -0.049993 | 3.7076e-08
| 0.1 | -0.0005 | 0.2 | -0.0001 | 0.1 | -0.0004 | 20 | 10 | 0 | 50.0308634E+00 | DC220MV | DC220MV | -0.039993 | 4.2491e-08
| 0.1 | -0.0004 | 0.2 | 0 | 0.1 | -0.0004 | 20 | 10 | 0 | 49.9736742E+00 | DC220MV | DC220MV | -0.039994 | 4.7936e-08
| 0.1 | -0.0004 | 0.2 | -0.0001 | 0.1 | -0.0003 | 20 | 10 | 0 | 50.0024577E+00 | DC220MV | DC220MV | -0.029994 | 6.0326e-08
| 0.1 | -0.0003 | 0.2 | 0 | 0.1 | -0.0003 | 20 | 10 | 0 | 50.0028256E+00 | DC220MV | DC220MV | -0.029994 | 2.7836e-08
| 0.1 | -0.0003 | 0.2 | -0.0001 | 0.1 | -0.0002 | 20 | 10 | 0 | 50.0306981E+00 | DC220MV | DC220MV | -0.019994 | 4.6542e-08
| 0.1 | -0.0002 | 0.2 | 0 | 0.1 | -0.0002 | 20 | 10 | 0 | 49.9751397E+00 | DC220MV | DC220MV | -0.019995 | 4.1285e-08
| 0.1 | -0.0002 | 0.2 | -0.0001 | 0.1 | -0.0001 | 20 | 10 | 0 | 50.0084212E+00 | DC220MV | DC220MV | -0.0099947 | 4.8006e-08
| 0.1 | -0.0001 | 0.2 | 0 | 0.1 | -0.0001 | 20 | 10 | 0 | 49.9834489E+00 | DC220MV | DC220MV | -0.0099951 | 4.321e-08
| 0.1 | -0.0001 | 0.2 | -0.0001 | 0.1 | 0 | 20 | 10 | 0 | 49.9227751E+00 | DC220MV | DC220MV | 4.9641e-06 | 3.9357e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9635784E+00 | DC220MV | DC220MV | 4.5697e-06 | 7.0682e-08
| 0.1 | -0.0001 | 0.2 | -0.0001 | 0.1 | 0 | 20 | 10 | 0 | 49.9812806E+00 | DC220MV | DC220MV | 5.163e-06 | 4.253e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9246373E+00 | DC220MV | DC220MV | 4.6397e-06 | 3.1806e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0779657E+00 | DC220MV | DC220MV | -0.009995 | 4.4132e-08
| 0.1 | 0.001 | 0.2 | 0 | 0.1 | 0.001 | 20 | 10 | 0 | 49.9783961E+00 | DC220MV | DC220MV | 0.1 | 3.0148e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9276849E+00 | DC220MV | DC220MV | 4.6034e-06 | 3.6841e-08
| 0.1 | 0.001 | 0.2 | 0.001 | 0.1 | 0 | 20 | 10 | 0 | 49.9834819E+00 | DC220MV | DC220MV | 3.8322e-07 | 3.3318e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9816624E+00 | DC220MV | DC2_2V | 4.4662e-06 | 4.7154e-08
| 0.1 | 0.001 | 0.2 | 0.001 | 0.1 | 0 | 20 | 10 | 0 | 49.9695751E+00 | DC220MV | DC2_2V | 0.010005 | 5.2663e-08
| 0.1 | 0.001 | 0.2 | 0 | 0.1 | 0.001 | 20 | 10 | 0 | 50.0398430E+00 | DC220MV | DC2_2V | 5.6183e-06 | 3.6452e-08
| 0.1 | 0.002 | 0.2 | 0.001 | 0.1 | 0.001 | 20 | 10 | 0 | 50.0250739E+00 | DC220MV | DC2_2V | 5.1692e-06 | 4.3858e-08
| 0.1 | 0.002 | 0.2 | 0 | 0.1 | 0.002 | 20 | 10 | 0 | 50.0171096E+00 | DC220MV | DC2_2V | 5.5943e-06 | 3.7356e-08
| 0.1 | 0.003 | 0.2 | 0.001 | 0.1 | 0.002 | 20 | 10 | 0 | 49.9844105E+00 | DC220MV | DC2_2V | 5.2605e-06 | 3.5703e-08
| 0.1 | 0.003 | 0.2 | 0 | 0.1 | 0.003 | 20 | 10 | 0 | 50.0200316E+00 | DC220MV | DC2_2V | 0.010005 | 3.5711e-08
| 0.1 | 0.004 | 0.2 | 0.001 | 0.1 | 0.003 | 20 | 10 | 0 | 50.0360873E+00 | DC220MV | DC2_2V | 0.010005 | 3.7356e-08
| 0.1 | 0.004 | 0.2 | 0 | 0.1 | 0.004 | 20 | 10 | 0 | 49.9919258E+00 | DC220MV | DC2_2V | 0.020005 | 4.5103e-08
| 0.1 | 0.005 | 0.2 | 0.001 | 0.1 | 0.004 | 20 | 10 | 0 | 50.0050854E+00 | DC220MV | DC2_2V | 0.020005 | 5.1348e-08
| 0.1 | 0.005 | 0.2 | 0 | 0.1 | 0.005 | 20 | 10 | 0 | 50.0059394E+00 | DC220MV | DC2_2V | 0.030005 | 4.6575e-08
| 0.1 | 0.006 | 0.2 | 0.001 | 0.1 | 0.005 | 20 | 10 | 0 | 50.0113469E+00 | DC220MV | DC2_2V | 0.030004 | 3.8575e-08
| 0.1 | 0.006 | 0.2 | 0 | 0.1 | 0.006 | 20 | 10 | 0 | 49.9915014E+00 | DC220MV | DC2_2V | 0.040004 | 4.146e-08
| 0.1 | 0.007 | 0.2 | 0.001 | 0.1 | 0.006 | 20 | 10 | 0 | 50.0096103E+00 | DC220MV | DC2_2V | 0.040004 | 5.3093e-08
| 0.1 | 0.007 | 0.2 | 0 | 0.1 | 0.007 | 20 | 10 | 0 | 50.0684095E+00 | DC220MV | DC2_2V | 0.050004 | 4.4105e-08
| 0.1 | 0.008 | 0.2 | 0.001 | 0.1 | 0.007 | 20 | 10 | 0 | 50.0477106E+00 | DC220MV | DC2_2V | 0.050003 | 4.989e-08
| 0.1 | 0.008 | 0.2 | 0 | 0.1 | 0.008 | 20 | 10 | 0 | 49.9717565E+00 | DC220MV | DC2_2V | 0.060003 | 4.2866e-08
| 0.1 | 0.009 | 0.2 | 0.001 | 0.1 | 0.008 | 20 | 10 | 0 | 50.0544838E+00 | DC220MV | DC2_2V | 0.060003 | 4.4228e-08
| 0.1 | 0.009 | 0.2 | 0 | 0.1 | 0.009 | 20 | 10 | 0 | 50.0214290E+00 | DC220MV | DC2_2V | 0.070003 | 7.9555e-08
| 0.1 | 0.01 | 0.2 | 0.001 | 0.1 | 0.009 | 20 | 10 | 0 | 49.9219569E+00 | DC220MV | DC2_2V | 0.070002 | 4.1837e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0338814E+00 | DC220MV | DC2_2V | 0.080002 | 7.1368e-08
| 0.1 | 0.01 | 0.2 | 0.001 | 0.1 | 0.009 | 20 | 10 | 0 | 50.1087360E+00 | DC220MV | DC2_2V | 0.080002 | 5.1504e-08
| 0.1 | 0.009 | 0.2 | 0 | 0.1 | 0.009 | 20 | 10 | 0 | 50.0456029E+00 | DC220MV | DC2_2V | 0.090002 | 4.66e-08
| 0.1 | 0.009 | 0.2 | 0.001 | 0.1 | 0.008 | 20 | 10 | 0 | 49.9975662E+00 | DC220MV | DC2_2V | 0.090002 | 3.2379e-08
| 0.1 | 0.008 | 0.2 | 0 | 0.1 | 0.008 | 20 | 10 | 0 | 50.0492891E+00 | DC220MV | DC2_2V | 0.1 | 5.4103e-08
| 0.1 | 0.008 | 0.2 | 0.001 | 0.1 | 0.007 | 20 | 10 | 0 | 49.9829162E+00 | DC220MV | DC2_2V | 0.090002 | 5.163e-08
| 0.1 | 0.007 | 0.2 | 0 | 0.1 | 0.007 | 20 | 10 | 0 | 49.9883046E+00 | DC220MV | DC2_2V | 0.090002 | 5.37e-07
| 0.1 | 0.007 | 0.2 | 0.001 | 0.1 | 0.006 | 20 | 10 | 0 | 49.9992547E+00 | DC220MV | DC2_2V | 0.080002 | 4.9294e-08
| 0.1 | 0.006 | 0.2 | 0 | 0.1 | 0.006 | 20 | 10 | 0 | 49.9908602E+00 | DC220MV | DC2_2V | 0.080003 | 4.2433e-08
| 0.1 | 0.006 | 0.2 | 0.001 | 0.1 | 0.005 | 20 | 10 | 0 | 49.9872721E+00 | DC220MV | DC2_2V | 0.070003 | 5.5564e-08
| 0.1 | 0.005 | 0.2 | 0 | 0.1 | 0.005 | 20 | 10 | 0 | 50.0022878E+00 | DC220MV | DC2_2V | 0.070003 | 4.4479e-08
| 0.1 | 0.005 | 0.2 | 0.001 | 0.1 | 0.004 | 20 | 10 | 0 | 49.9996226E+00 | DC220MV | DC2_2V | 0.060003 | 3.8506e-08
| 0.1 | 0.004 | 0.2 | 0 | 0.1 | 0.004 | 20 | 10 | 0 | 50.0800194E+00 | DC220MV | DC2_2V | 0.060004 | 5.6544e-08
| 0.1 | 0.004 | 0.2 | 0.001 | 0.1 | 0.003 | 20 | 10 | 0 | 49.9637951E+00 | DC220MV | DC2_2V | 0.050004 | 4.5149e-08
| 0.1 | 0.003 | 0.2 | 0 | 0.1 | 0.003 | 20 | 10 | 0 | 49.9802531E+00 | DC220MV | DC2_2V | 0.050004 | 5.6065e-08
| 0.1 | 0.003 | 0.2 | 0.001 | 0.1 | 0.002 | 20 | 10 | 0 | 49.9646712E+00 | DC220MV | DC2_2V | 0.040004 | 6.4814e-08
| 0.1 | 0.002 | 0.2 | 0 | 0.1 | 0.002 | 20 | 10 | 0 | 49.9537362E+00 | DC220MV | DC2_2V | 0.040005 | 4.3282e-08
| 0.1 | 0.002 | 0.2 | 0.001 | 0.1 | 0.001 | 20 | 10 | 0 | 49.9582519E+00 | DC220MV | DC2_2V | 0.030005 | 3.3781e-08
| 0.1 | 0.001 | 0.2 | 0 | 0.1 | 0.001 | 20 | 10 | 0 | 50.0248756E+00 | DC220MV | DC2_2V | 0.030005 | 4.149e-08
| 0.1 | 0.001 | 0.2 | 0.001 | 0.1 | 0 | 20 | 10 | 0 | 50.0470206E+00 | DC220MV | DC2_2V | 0.020005 | 5.0905e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0172276E+00 | DC220MV | DC2_2V | 0.020005 | 3.9469e-08
| 0.1 | 0.001 | 0.2 | 0.001 | 0.1 | 0 | 20 | 10 | 0 | 49.9859143E+00 | DC220MV | DC220MV | 0.010005 | 3.5281e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0120690E+00 | DC220MV | DC220MV | 0.010006 | 3.0983e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0356528E+00 | DC220MV | DC220MV | 5.4446e-06 | 5.5812e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9932603E+00 | DC220MV | DC220MV | 5.4484e-06 | 8.7916e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0500075E+00 | DC220MV | DC220MV | 4.7549e-06 | 6.0539e-08
| 0.1 | -0.001 | 0.2 | -0.001 | 0.1 | 0 | 20 | 10 | 0 | 50.0102144E+00 | DC220MV | DC220MV | 4.8223e-06 | 5.8733e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9887808E+00 | DC220MV | DC2_2V | 0.010004 | 4.5554e-08
| 0.1 | -0.001 | 0.2 | -0.001 | 0.1 | 0 | 20 | 10 | 0 | 50.0200505E+00 | DC220MV | DC2_2V | -0.0099949 | 2.9958e-08
| 0.1 | -0.001 | 0.2 | 0 | 0.1 | -0.001 | 20 | 10 | 0 | 49.9935952E+00 | DC220MV | DC2_2V | 4.6438e-06 | 6.5219e-08
| 0.1 | -0.002 | 0.2 | -0.001 | 0.1 | -0.001 | 20 | 10 | 0 | 50.0334138E+00 | DC220MV | DC2_2V | 5.0585e-06 | 2.973e-08
| 0.1 | -0.002 | 0.2 | 0 | 0.1 | -0.002 | 20 | 10 | 0 | 50.0612684E+00 | DC220MV | DC2_2V | 4.5116e-06 | 4.2838e-08
| 0.1 | -0.003 | 0.2 | -0.001 | 0.1 | -0.002 | 20 | 10 | 0 | 50.0152783E+00 | DC220MV | DC2_2V | 4.9348e-06 | 3.9575e-08
| 0.1 | -0.003 | 0.2 | 0 | 0.1 | -0.003 | 20 | 10 | 0 | 49.9900256E+00 | DC220MV | DC2_2V | -0.0099951 | 3.8644e-08
| 0.1 | -0.004 | 0.2 | -0.001 | 0.1 | -0.003 | 20 | 10 | 0 | 49.9713042E+00 | DC220MV | DC2_2V | -0.0099946 | 6.0552e-08
| 0.1 | -0.004 | 0.2 | 0 | 0.1 | -0.004 | 20 | 10 | 0 | 49.9393284E+00 | DC220MV | DC2_2V | -0.019995 | 7.5968e-08
| 0.1 | -0.005 | 0.2 | -0.001 | 0.1 | -0.004 | 20 | 10 | 0 | 49.9646194E+00 | DC220MV | DC2_2V | -0.019994 | 3.544e-08
| 0.1 | -0.005 | 0.2 | 0 | 0.1 | -0.005 | 20 | 10 | 0 | 49.9501535E+00 | DC220MV | DC2_2V | -0.029994 | 4.0054e-08
| 0.1 | -0.006 | 0.2 | -0.001 | 0.1 | -0.005 | 20 | 10 | 0 | 50.0500501E+00 | DC220MV | DC2_2V | -0.029994 | 3.9743e-08
| 0.1 | -0.006 | 0.2 | 0 | 0.1 | -0.006 | 20 | 10 | 0 | 49.9714738E+00 | DC220MV | DC2_2V | -0.039994 | 5.294e-08
| 0.1 | -0.007 | 0.2 | -0.001 | 0.1 | -0.006 | 20 | 10 | 0 | 50.0005236E+00 | DC220MV | DC2_2V | -0.039993 | 3.5819e-08
| 0.1 | -0.007 | 0.2 | 0 | 0.1 | -0.007 | 20 | 10 | 0 | 50.0048354E+00 | DC220MV | DC2_2V | -0.049993 | 4.8468e-08
| 0.1 | -0.008 | 0.2 | -0.001 | 0.1 | -0.007 | 20 | 10 | 0 | 50.0362243E+00 | DC220MV | DC2_2V | -0.049993 | 3.4955e-08
| 0.1 | -0.008 | 0.2 | 0 | 0.1 | -0.008 | 20 | 10 | 0 | 49.9836091E+00 | DC220MV | DC2_2V | -0.059993 | 4.5685e-08
| 0.1 | -0.009 | 0.2 | -0.001 | 0.1 | -0.008 | 20 | 10 | 0 | 50.0367298E+00 | DC220MV | DC2_2V | -0.059993 | 3.6657e-08
| 0.1 | -0.009 | 0.2 | 0 | 0.1 | -0.009 | 20 | 10 | 0 | 49.9842738E+00 | DC220MV | DC2_2V | -0.069993 | 5.076e-08
| 0.1 | -0.01 | 0.2 | -0.001 | 0.1 | -0.009 | 20 | 10 | 0 | 49.9949250E+00 | DC220MV | DC2_2V | -0.069992 | 4.5716e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9986793E+00 | DC220MV | DC2_2V | -0.079993 | 5.4041e-08
| 0.1 | -0.01 | 0.2 | -0.001 | 0.1 | -0.009 | 20 | 10 | 0 | 49.9566037E+00 | DC220MV | DC2_2V | -0.079992 | 4.7165e-08
| 0.1 | -0.009 | 0.2 | 0 | 0.1 | -0.009 | 20 | 10 | 0 | 50.1014460E+00 | DC220MV | DC2_2V | -0.089992 | 4.5394e-08
| 0.1 | -0.009 | 0.2 | -0.001 | 0.1 | -0.008 | 20 | 10 | 0 | 49.9959437E+00 | DC220MV | DC2_2V | -0.089992 | 2.9758e-08
| 0.1 | -0.008 | 0.2 | 0 | 0.1 | -0.008 | 20 | 10 | 0 | 50.0459715E+00 | DC220MV | DC2_2V | -0.099992 | 4.8841e-08
| 0.1 | -0.008 | 0.2 | -0.001 | 0.1 | -0.007 | 20 | 10 | 0 | 49.9684774E+00 | DC220MV | DC2_2V | -0.089992 | 3.9385e-08
| 0.1 | -0.007 | 0.2 | 0 | 0.1 | -0.007 | 20 | 10 | 0 | 49.9719874E+00 | DC220MV | DC2_2V | -0.089992 | 7.4541e-08
| 0.1 | -0.007 | 0.2 | -0.001 | 0.1 | -0.006 | 20 | 10 | 0 | 49.9837600E+00 | DC220MV | DC2_2V | -0.079992 | 3.2109e-08
| 0.1 | -0.006 | 0.2 | 0 | 0.1 | -0.006 | 20 | 10 | 0 | 50.0258577E+00 | DC220MV | DC2_2V | -0.079993 | 5.3007e-08
| 0.1 | -0.006 | 0.2 | -0.001 | 0.1 | -0.005 | 20 | 10 | 0 | 49.9915910E+00 | DC220MV | DC2_2V | -0.069993 | 4.2062e-08
| 0.1 | -0.005 | 0.2 | 0 | 0.1 | -0.005 | 20 | 10 | 0 | 50.0515247E+00 | DC220MV | DC2_2V | -0.069993 | 3.84e-08
| 0.1 | -0.005 | 0.2 | -0.001 | 0.1 | -0.004 | 20 | 10 | 0 | 50.0354119E+00 | DC220MV | DC2_2V | -0.059993 | 4.2363e-08
| 0.1 | -0.004 | 0.2 | 0 | 0.1 | -0.004 | 20 | 10 | 0 | 50.0273499E+00 | DC220MV | DC2_2V | -0.059993 | 4.3653e-08
| 0.1 | -0.004 | 0.2 | -0.001 | 0.1 | -0.003 | 20 | 10 | 0 | 49.9984812E+00 | DC220MV | DC2_2V | -0.049993 | 6.0246e-08
| 0.1 | -0.003 | 0.2 | 0 | 0.1 | -0.003 | 20 | 10 | 0 | 49.9898417E+00 | DC220MV | DC2_2V | -0.049993 | 3.7076e-08
| 0.1 | -0.003 | 0.2 | -0.001 | 0.1 | -0.002 | 20 | 10 | 0 | 49.9354655E+00 | DC220MV | DC2_2V | -0.039993 | 4.2491e-08
| 0.1 | -0.002 | 0.2 | 0 | 0.1 | -0.002 | 20 | 10 | 0 | 49.9777363E+00 | DC220MV | DC2_2V | -0.039994 | 4.7936e-08
| 0.1 | -0.002 | 0.2 | -0.001 | 0.1 | -0.001 | 20 | 10 | 0 | 50.0357094E+00 | DC220MV | DC2_2V | -0.029994 | 6.0326e-08
| 0.1 | -0.001 | 0.2 | 0 | 0.1 | -0.001 | 20 | 10 | 0 | 49.9519283E+00 | DC220MV | DC2_2V | -0.029994 | 2.7836e-08
| 0.1 | -0.001 | 0.2 | -0.001 | 0.1 | 0 | 20 | 10 | 0 | 49.9450228E+00 | DC220MV | DC2_2V | -0.019994 | 4.6542e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0212165E+00 | DC220MV | DC2_2V | -0.019995 | 4.1285e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9804463E+00 | DC220MV | DC220MV | -0.0099947 | 4.8006e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9891109E+00 | DC220MV | DC220MV | -0.0099951 | 4.321e-08
| 0.1 | -0.001 | 0.2 | 0 | 0.1 | -0.001 | 20 | 10 | 0 | 49.9998066E+00 | DC220MV | DC220MV | 4.9641e-06 | 3.9357e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0422385E+00 | DC2_2V | DC2_2V | 4.5697e-06 | 7.0682e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9700039E+00 | DC2_2V | DC2_2V | 5.163e-06 | 4.253e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0 | 20 | 10 | 0 | 50.0165526E+00 | DC2_2V | DC2_2V | 4.6397e-06 | 3.1806e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9795602E+00 | DC2_2V | DC11V | -0.009995 | 4.4132e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.8893445E+00 | DC2_2V | DC11V | 0.1 | 3.0148e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 49.9553088E+00 | DC2_2V | DC11V | 4.6034e-06 | 3.6841e-08
| 0.1 | 0.02 | 0.2 | 0.01 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0452768E+00 | DC2_2V | DC11V | 3.8322e-07 | 3.3318e-08
| 0.1 | 0.02 | 0.2 | 0 | 0.1 | 0.02 | 20 | 10 | 0 | 49.9977501E+00 | DC2_2V | DC11V | 4.4662e-06 | 4.7154e-08
| 0.1 | 0.03 | 0.2 | 0.01 | 0.1 | 0.02 | 20 | 10 | 0 | 50.0547486E+00 | DC2_2V | DC11V | 0.010005 | 5.2663e-08
| 0.1 | 0.03 | 0.2 | 0 | 0.1 | 0.03 | 20 | 10 | 0 | 50.0525741E+00 | DC2_2V | DC11V | 5.6183e-06 | 3.6452e-08
| 0.1 | 0.04 | 0.2 | 0.01 | 0.1 | 0.03 | 20 | 10 | 0 | 49.9871212E+00 | DC2_2V | DC11V | 5.1692e-06 | 4.3858e-08
| 0.1 | 0.04 | 0.2 | 0 | 0.1 | 0.04 | 20 | 10 | 0 | 49.9714785E+00 | DC2_2V | DC11V | 5.5943e-06 | 3.7356e-08
| 0.1 | 0.05 | 0.2 | 0.01 | 0.1 | 0.04 | 20 | 10 | 0 | 50.0065527E+00 | DC2_2V | DC11V | 5.2605e-06 | 3.5703e-08
| 0.1 | 0.05 | 0.2 | 0 | 0.1 | 0.05 | 20 | 10 | 0 | 49.9839957E+00 | DC2_2V | DC11V | 0.010005 | 3.5711e-08
| 0.1 | 0.05 | 0.2 | 0.01 | 0.1 | 0.04 | 20 | 10 | 0 | 49.9931142E+00 | DC2_2V | DC11V | 0.010005 | 3.7356e-08
| 0.1 | 0.04 | 0.2 | 0 | 0.1 | 0.04 | 20 | 10 | 0 | 49.9825250E+00 | DC2_2V | DC11V | 0.020005 | 4.5103e-08
| 0.1 | 0.04 | 0.2 | 0.01 | 0.1 | 0.03 | 20 | 10 | 0 | 50.0452579E+00 | DC2_2V | DC11V | 0.020005 | 5.1348e-08
| 0.1 | 0.03 | 0.2 | 0 | 0.1 | 0.03 | 20 | 10 | 0 | 49.8609679E+00 | DC2_2V | DC11V | 0.030005 | 4.6575e-08
| 0.1 | 0.03 | 0.2 | 0.01 | 0.1 | 0.02 | 20 | 10 | 0 | 49.9730900E+00 | DC2_2V | DC11V | 0.030004 | 3.8575e-08
| 0.1 | 0.02 | 0.2 | 0 | 0.1 | 0.02 | 20 | 10 | 0 | 50.0250172E+00 | DC2_2V | DC11V | 0.040004 | 4.146e-08
| 0.1 | 0.02 | 0.2 | 0.01 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0696488E+00 | DC2_2V | DC11V | 0.040004 | 5.3093e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0135508E+00 | DC2_2V | DC11V | 0.050004 | 4.4105e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9570887E+00 | DC2_2V | DC11V | 0.050003 | 4.989e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9908696E+00 | DC2_2V | DC11V | 0.060003 | 4.2866e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9654861E+00 | DC2_2V | DC11V | 0.060003 | 4.4228e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0630369E+00 | DC2_2V | DC11V | 0.070003 | 7.9555e-08
| 0.1 | 0.1 | 0.2 | 0 | 0.1 | 0.1 | 20 | 10 | 0 | 50.0308728E+00 | DC2_2V | DC11V | 0.070002 | 4.1837e-08
| 0.1 | -0.1 | 0.2 | 0 | 0.1 | -0.1 | 20 | 10 | 0 | 50.0105777E+00 | DC2_2V | DC11V | 0.080002 | 7.1368e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0179546E+00 | DC2_2V | DC11V | 0.080002 | 5.1504e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9165076E+00 | DC2_2V | DC11V | 0.090002 | 4.66e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9685198E+00 | DC2_2V | DC11V | 0.090002 | 3.2379e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9781322E+00 | DC2_2V | DC11V | 0.1 | 5.4103e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9635549E+00 | DC2_2V | DC11V | 0.090002 | 5.163e-08
| 0.1 | -0.02 | 0.2 | -0.01 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9867064E+00 | DC2_2V | DC11V | 0.090002 | 5.37e-07
| 0.1 | -0.02 | 0.2 | 0 | 0.1 | -0.02 | 20 | 10 | 0 | 49.9604323E+00 | DC2_2V | DC11V | 0.080002 | 4.9294e-08
| 0.1 | -0.03 | 0.2 | -0.01 | 0.1 | -0.02 | 20 | 10 | 0 | 49.9561470E+00 | DC2_2V | DC11V | 0.080003 | 4.2433e-08
| 0.1 | -0.03 | 0.2 | 0 | 0.1 | -0.03 | 20 | 10 | 0 | 50.0151980E+00 | DC2_2V | DC11V | 0.070003 | 5.5564e-08
| 0.1 | -0.04 | 0.2 | -0.01 | 0.1 | -0.03 | 20 | 10 | 0 | 50.0420353E+00 | DC2_2V | DC11V | 0.070003 | 4.4479e-08
| 0.1 | -0.04 | 0.2 | 0 | 0.1 | -0.04 | 20 | 10 | 0 | 49.9588546E+00 | DC2_2V | DC11V | 0.060003 | 3.8506e-08
| 0.1 | -0.05 | 0.2 | -0.01 | 0.1 | -0.04 | 20 | 10 | 0 | 49.9365852E+00 | DC2_2V | DC11V | 0.060004 | 5.6544e-08
| 0.1 | -0.05 | 0.2 | 0 | 0.1 | -0.05 | 20 | 10 | 0 | 49.9921333E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 0.1 | -0.05 | 0.2 | -0.01 | 0.1 | -0.04 | 20 | 10 | 0 | 50.0283699E+00 | DC2_2V | DC11V | 0.050004 | 5.6065e-08
| 0.1 | -0.04 | 0.2 | 0 | 0.1 | -0.04 | 20 | 10 | 0 | 50.0458392E+00 | DC2_2V | DC11V | 0.040004 | 6.4814e-08
| 0.1 | -0.04 | 0.2 | -0.01 | 0.1 | -0.03 | 20 | 10 | 0 | 50.0737454E+00 | DC2_2V | DC11V | 0.040005 | 4.3282e-08
| 0.1 | -0.03 | 0.2 | 0 | 0.1 | -0.03 | 20 | 10 | 0 | 50.0049297E+00 | DC2_2V | DC2_2V | 0.030005 | 3.3781e-08
| 0.1 | -0.03 | 0.2 | -0.01 | 0.1 | -0.02 | 20 | 10 | 0 | 49.9964012E+00 | DC2_2V | DC2_2V | 0.030005 | 4.149e-08
| 0.1 | -0.02 | 0.2 | 0 | 0.1 | -0.02 | 20 | 10 | 0 | 49.9888138E+00 | DC2_2V | DC2_2V | 0.020005 | 5.0905e-08
| 0.1 | -0.02 | 0.2 | -0.01 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9775714E+00 | DC2_2V | DC2_2V | 0.020005 | 3.9469e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0598120E+00 | DC2_2V | DC2_2V | 0.010005 | 3.5281e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9908837E+00 | DC2_2V | DC2_2V | 0.010006 | 3.0983e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0086666E+00 | DC2_2V | DC11V | 5.4446e-06 | 5.5812e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9911761E+00 | DC2_2V | DC11V | 5.4484e-06 | 8.7916e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 4.7549e-06 | 6.0539e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0652170E+00 | DC2_2V | DC11V | 4.8223e-06 | 5.8733e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0098227E+00 | DC2_2V | DC11V | 4.9348e-06 | 3.9575e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0548715E+00 | DC2_2V | DC11V | -0.0099951 | 3.8644e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0 | 20 | 10 | 0 | 50.0107523E+00 | DC2_2V | DC11V | -0.0099946 | 6.0552e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0186626E+00 | DC2_2V | DC11V | -0.019995 | 7.5968e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0 | 20 | 10 | 0 | 50.0054346E+00 | DC2_2V | DC11V | -0.019994 | 3.544e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 49.9804275E+00 | DC2_2V | DC11V | -0.029994 | 4.0054e-08
| 0.1 | 0.02 | 0.2 | 0.01 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0244601E+00 | DC2_2V | DC11V | -0.029994 | 3.9743e-08
| 0.1 | 0.02 | 0.2 | 0 | 0.1 | 0.02 | 20 | 10 | 0 | 49.9897568E+00 | DC2_2V | DC11V | -0.039994 | 5.294e-08
| 0.1 | 0.03 | 0.2 | 0.01 | 0.1 | 0.02 | 20 | 10 | 0 | 50.0130742E+00 | DC2_2V | DC11V | -0.039993 | 3.5819e-08
| 0.1 | 0.03 | 0.2 | 0 | 0.1 | 0.03 | 20 | 10 | 0 | 49.9879746E+00 | DC2_2V | DC11V | -0.049993 | 4.8468e-08
| 0.1 | 0.04 | 0.2 | 0.01 | 0.1 | 0.03 | 20 | 10 | 0 | 49.9643792E+00 | DC2_2V | DC11V | -0.049993 | 3.4955e-08
| 0.1 | 0.04 | 0.2 | 0 | 0.1 | 0.04 | 20 | 10 | 0 | 49.9486471E+00 | DC2_2V | DC11V | -0.059993 | 4.5685e-08
| 0.1 | 0.05 | 0.2 | 0.01 | 0.1 | 0.04 | 20 | 10 | 0 | 49.9932603E+00 | DC2_2V | DC11V | -0.059993 | 3.6657e-08
| 0.1 | 0.05 | 0.2 | 0 | 0.1 | 0.05 | 20 | 10 | 0 | 50.0016038E+00 | DC2_2V | DC11V | -0.069993 | 5.076e-08
| 0.1 | 0.06 | 0.2 | 0.01 | 0.1 | 0.05 | 20 | 10 | 0 | 50.0230484E+00 | DC2_2V | DC11V | -0.069992 | 4.5716e-08
| 0.1 | 0.06 | 0.2 | 0 | 0.1 | 0.06 | 20 | 10 | 0 | 49.9626506E+00 | DC2_2V | DC11V | -0.079993 | 5.4041e-08
| 0.1 | 0.07 | 0.2 | 0.01 | 0.1 | 0.06 | 20 | 10 | 0 | 49.9980991E+00 | DC2_2V | DC11V | -0.079992 | 4.7165e-08
| 0.1 | 0.07 | 0.2 | 0 | 0.1 | 0.07 | 20 | 10 | 0 | 49.9965851E+00 | DC2_2V | DC11V | -0.089992 | 4.5394e-08
| 0.1 | 0.08 | 0.2 | 0.01 | 0.1 | 0.07 | 20 | 10 | 0 | 49.9741407E+00 | DC2_2V | DC11V | -0.089992 | 2.9758e-08
| 0.1 | 0.08 | 0.2 | 0 | 0.1 | 0.08 | 20 | 10 | 0 | 50.0042975E+00 | DC2_2V | DC11V | -0.099992 | 4.8841e-08
| 0.1 | 0.09 | 0.2 | 0.01 | 0.1 | 0.08 | 20 | 10 | 0 | 49.9474892E+00 | DC2_2V | DC11V | -0.089992 | 3.9385e-08
| 0.1 | 0.09 | 0.2 | 0 | 0.1 | 0.09 | 20 | 10 | 0 | 50.0383738E+00 | DC2_2V | DC11V | -0.089992 | 7.4541e-08
| 0.1 | 0.1 | 0.2 | 0.01 | 0.1 | 0.09 | 20 | 10 | 0 | 49.9884791E+00 | DC2_2V | DC11V | -0.079992 | 3.2109e-08
| 0.1 | 0.1 | 0.2 | 0 | 0.1 | 0.1 | 20 | 10 | 0 | 50.0185021E+00 | DC2_2V | DC11V | -0.079993 | 5.3007e-08
| 0.1 | 0.1 | 0.2 | 0.01 | 0.1 | 0.09 | 20 | 10 | 0 | 50.0553773E+00 | DC2_2V | DC11V | -0.069993 | 4.2062e-08
| 0.1 | 0.09 | 0.2 | 0 | 0.1 | 0.09 | 20 | 10 | 0 | 49.9801353E+00 | DC2_2V | DC11V | -0.069993 | 3.84e-08
| 0.1 | 0.09 | 0.2 | 0.01 | 0.1 | 0.08 | 20 | 10 | 0 | 49.9973822E+00 | DC2_2V | DC11V | -0.059993 | 4.2363e-08
| 0.1 | 0.08 | 0.2 | 0 | 0.1 | 0.08 | 20 | 10 | 0 | 50.0116820E+00 | DC2_2V | DC11V | -0.059993 | 4.3653e-08
| 0.1 | 0.08 | 0.2 | 0.01 | 0.1 | 0.07 | 20 | 10 | 0 | 50.0381470E+00 | DC2_2V | DC11V | -0.049993 | 6.0246e-08
| 0.1 | 0.07 | 0.2 | 0 | 0.1 | 0.07 | 20 | 10 | 0 | 49.9937932E+00 | DC2_2V | DC11V | -0.049993 | 3.7076e-08
| 0.1 | 0.07 | 0.2 | 0.01 | 0.1 | 0.06 | 20 | 10 | 0 | 50.0414069E+00 | DC2_2V | DC11V | -0.039993 | 4.2491e-08
| 0.1 | 0.06 | 0.2 | 0 | 0.1 | 0.06 | 20 | 10 | 0 | 50.0424417E+00 | DC2_2V | DC11V | -0.039994 | 4.7936e-08
| 0.1 | 0.06 | 0.2 | 0.01 | 0.1 | 0.05 | 20 | 10 | 0 | 49.9540752E+00 | DC2_2V | DC2_2V | -0.029994 | 6.0326e-08
| 0.1 | 0.05 | 0.2 | 0 | 0.1 | 0.05 | 20 | 10 | 0 | 49.9572441E+00 | DC2_2V | DC2_2V | -0.029994 | 2.7836e-08
| 0.1 | 0.05 | 0.2 | 0.01 | 0.1 | 0.04 | 20 | 10 | 0 | 50.0156511E+00 | DC2_2V | DC2_2V | -0.019994 | 4.6542e-08
| 0.1 | 0.04 | 0.2 | 0 | 0.1 | 0.04 | 20 | 10 | 0 | 50.0131308E+00 | DC220MV | DC220MV | -0.019995 | 4.1285e-08
| 0.1 | 0.04 | 0.2 | 0.01 | 0.1 | 0.03 | 20 | 10 | 0 | 49.9719827E+00 | DC220MV | DC220MV | -0.0099947 | 4.8006e-08
| 0.1 | 0.03 | 0.2 | 0 | 0.1 | 0.03 | 20 | 10 | 0 | 50.0056421E+00 | DC220MV | DC220MV | -0.0099951 | 4.321e-08
| 0.1 | 0.03 | 0.2 | 0.01 | 0.1 | 0.02 | 20 | 10 | 0 | 49.9625988E+00 | DC220MV | DC220MV | 4.9641e-06 | 3.9357e-08
| 0.1 | 0.02 | 0.2 | 0 | 0.1 | 0.02 | 20 | 10 | 0 | 50.0264432E+00 | DC220MV | DC220MV | 4.5697e-06 | 7.0682e-08
| 0.1 | 0.02 | 0.2 | 0.01 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0096339E+00 | DC220MV | DC220MV | 5.163e-06 | 4.253e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 49.9681288E+00 | DC220MV | DC220MV | 4.6397e-06 | 3.1806e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9975426E+00 | DC220MV | DC220MV | -0.009995 | 4.4132e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0329037E+00 | DC220MV | DC220MV | 0.1 | 3.0148e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0 | 20 | 10 | 0 | 50.0029530E+00 | DC220MV | DC220MV | 4.6034e-06 | 3.6841e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9937083E+00 | DC220MV | DC220MV | 3.8322e-07 | 3.3318e-08
| 0.1 | 0.01 | 0.2 | 0 | 0.1 | 0.01 | 20 | 10 | 0 | 50.0149054E+00 | DC220MV | DC220MV | 4.4662e-06 | 4.7154e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9931330E+00 | DC220MV | DC220MV | 0.010005 | 5.2663e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9917466E+00 | DC220MV | DC220MV | 5.6183e-06 | 3.6452e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 50.0594385E+00 | DC220MV | DC220MV | 5.1692e-06 | 4.3858e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9829162E+00 | DC220MV | DC220MV | 5.5943e-06 | 3.7356e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 49.9492356E+00 | DC220MV | DC220MV | 5.2605e-06 | 3.5703e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0236433E+00 | DC220MV | DC220MV | 0.010005 | 3.5711e-08
| 0.1 | -0.02 | 0.2 | -0.01 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9746355E+00 | DC220MV | DC220MV | 0.010005 | 3.7356e-08
| 0.1 | -0.02 | 0.2 | 0 | 0.1 | -0.02 | 20 | 10 | 0 | 49.9932698E+00 | DC220MV | DC220MV | 0.020005 | 4.5103e-08
| 0.1 | -0.03 | 0.2 | -0.01 | 0.1 | -0.02 | 20 | 10 | 0 | 49.9617510E+00 | DC220MV | DC220MV | 0.020005 | 5.1348e-08
| 0.1 | -0.03 | 0.2 | 0 | 0.1 | -0.03 | 20 | 10 | 0 | 49.9962880E+00 | DC220MV | DC220MV | 0.030005 | 4.6575e-08
| 0.1 | -0.04 | 0.2 | -0.01 | 0.1 | -0.03 | 20 | 10 | 0 | 49.9896012E+00 | DC220MV | DC220MV | 0.030004 | 3.8575e-08
| 0.1 | -0.04 | 0.2 | 0 | 0.1 | -0.04 | 20 | 10 | 0 | 50.0031417E+00 | DC220MV | DC220MV | 0.040004 | 4.146e-08
| 0.1 | -0.05 | 0.2 | -0.01 | 0.1 | -0.04 | 20 | 10 | 0 | 49.9789805E+00 | DC220MV | DC220MV | 0.040004 | 5.3093e-08
| 0.1 | -0.05 | 0.2 | 0 | 0.1 | -0.05 | 20 | 10 | 0 | 50.0171238E+00 | DC220MV | DC220MV | 0.050004 | 4.4105e-08
| 0.1 | -0.06 | 0.2 | -0.01 | 0.1 | -0.05 | 20 | 10 | 0 | 49.9998915E+00 | DC220MV | DC220MV | 0.050003 | 4.989e-08
| 0.1 | -0.06 | 0.2 | 0 | 0.1 | -0.06 | 20 | 10 | 0 | 49.9831943E+00 | DC220MV | DC220MV | 0.060003 | 4.2866e-08
| 0.1 | -0.07 | 0.2 | -0.01 | 0.1 | -0.06 | 20 | 10 | 0 | 49.9900492E+00 | DC220MV | DC220MV | 0.060003 | 4.4228e-08
| 0.1 | -0.07 | 0.2 | 0 | 0.1 | -0.07 | 20 | 10 | 0 | 49.9985614E+00 | DC220MV | DC220MV | 0.070003 | 7.9555e-08
| 0.1 | -0.08 | 0.2 | -0.01 | 0.1 | -0.07 | 20 | 10 | 0 | 50.0033021E+00 | DC220MV | DC220MV | 0.070002 | 4.1837e-08
| 0.1 | -0.08 | 0.2 | 0 | 0.1 | -0.08 | 20 | 10 | 0 | 49.9917985E+00 | DC220MV | DC220MV | 0.080002 | 7.1368e-08
| 0.1 | -0.09 | 0.2 | -0.01 | 0.1 | -0.08 | 20 | 10 | 0 | 49.9843775E+00 | DC220MV | DC220MV | 0.080002 | 5.1504e-08
| 0.1 | -0.09 | 0.2 | 0 | 0.1 | -0.09 | 20 | 10 | 0 | 50.0249511E+00 | DC220MV | DC220MV | 0.090002 | 4.66e-08
| 0.1 | -0.1 | 0.2 | -0.01 | 0.1 | -0.09 | 20 | 10 | 0 | 49.9716387E+00 | DC220MV | DC220MV | 0.090002 | 3.2379e-08
| 0.1 | -0.1 | 0.2 | 0 | 0.1 | -0.1 | 20 | 10 | 0 | 49.9675352E+00 | DC220MV | DC220MV | 0.1 | 5.4103e-08
| 0.1 | -0.1 | 0.2 | -0.01 | 0.1 | -0.09 | 20 | 10 | 0 | 50.0280393E+00 | DC220MV | DC220MV | 0.090002 | 5.163e-08
| 0.1 | -0.09 | 0.2 | 0 | 0.1 | -0.09 | 20 | 10 | 0 | 50.0554009E+00 | DC220MV | DC220MV | 0.090002 | 5.37e-07
| 0.1 | -0.09 | 0.2 | -0.01 | 0.1 | -0.08 | 20 | 10 | 0 | 49.9716906E+00 | DC220MV | DC220MV | 0.080002 | 4.9294e-08
| 0.1 | -0.08 | 0.2 | 0 | 0.1 | -0.08 | 20 | 10 | 0 | 49.9874135E+00 | DC220MV | DC220MV | 0.080003 | 4.2433e-08
| 0.1 | -0.08 | 0.2 | -0.01 | 0.1 | -0.07 | 20 | 10 | 0 | 49.9689156E+00 | DC220MV | DC220MV | 0.070003 | 5.5564e-08
| 0.1 | -0.07 | 0.2 | 0 | 0.1 | -0.07 | 20 | 10 | 0 | 49.9946185E+00 | DC220MV | DC220MV | 0.070003 | 4.4479e-08
| 0.1 | -0.07 | 0.2 | -0.01 | 0.1 | -0.06 | 20 | 10 | 0 | 49.9880877E+00 | DC220MV | DC220MV | 0.060003 | 3.8506e-08
| 0.1 | -0.06 | 0.2 | 0 | 0.1 | -0.06 | 20 | 10 | 0 | 50.0087987E+00 | DC220MV | DC220MV | 0.060004 | 5.6544e-08
| 0.1 | -0.06 | 0.2 | -0.01 | 0.1 | -0.05 | 20 | 10 | 0 | 49.9794990E+00 | DC220MV | DC220MV | 0.050004 | 4.5149e-08
| 0.1 | -0.05 | 0.2 | 0 | 0.1 | -0.05 | 20 | 10 | 0 | 49.9480446E+00 | DC220MV | DC220MV | 0.050004 | 5.6065e-08
| 0.1 | -0.05 | 0.2 | -0.01 | 0.1 | -0.04 | 20 | 10 | 0 | 50.0089497E+00 | DC220MV | DC220MV | 0.040004 | 6.4814e-08
| 0.1 | -0.04 | 0.2 | 0 | 0.1 | -0.04 | 20 | 10 | 0 | 49.9811958E+00 | DC220MV | DC220MV | 0.040005 | 4.3282e-08
| 0.1 | -0.04 | 0.2 | -0.01 | 0.1 | -0.03 | 20 | 10 | 0 | 49.9946138E+00 | DC220MV | DC220MV | 0.030005 | 3.3781e-08
| 0.1 | -0.03 | 0.2 | 0 | 0.1 | -0.03 | 20 | 10 | 0 | 50.0182567E+00 | DC220MV | DC220MV | 0.030005 | 4.149e-08
| 0.1 | -0.03 | 0.2 | -0.01 | 0.1 | -0.02 | 20 | 10 | 0 | 50.0069538E+00 | DC220MV | DC220MV | 0.020005 | 5.0905e-08
| 0.1 | -0.02 | 0.2 | 0 | 0.1 | -0.02 | 20 | 10 | 0 | 49.9680582E+00 | DC220MV | DC220MV | 0.020005 | 3.9469e-08
| 0.1 | -0.02 | 0.2 | -0.01 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0065008E+00 | DC220MV | DC220MV | 0.010005 | 3.5281e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0455414E+00 | DC220MV | DC220MV | 0.010006 | 3.0983e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 50.0172890E+00 | DC220MV | DC220MV | 5.4446e-06 | 5.5812e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0313923E+00 | DC220MV | DC220MV | 5.4484e-06 | 8.7916e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0 | 20 | 10 | 0 | 50.0098180E+00 | DC220MV | DC220MV | 4.7549e-06 | 6.0539e-08
| 0.1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9871212E+00 | DC220MV | DC220MV | 4.8223e-06 | 5.8733e-08
| 0.1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9888091E+00 | DC220MV | DC220MV | 0.010004 | 4.5554e-08
| 1 | 0.1 | 0.2 | 0 | 0.1 | 0.1 | 20 | 10 | 0 | 50.0417802E+00 | DC220MV | DC220MV | -0.0099949 | 2.9958e-08
| 1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9707247E+00 | DC220MV | DC220MV | 4.6438e-06 | 6.5219e-08
| 1 | 0.1 | 0.2 | 0.05 | 0.1 | 0.05 | 20 | 10 | 0 | 49.9903839E+00 | DC220MV | DC220MV | 5.0585e-06 | 2.973e-08
| 1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0350529E+00 | DC220MV | DC220MV | 4.5116e-06 | 4.2838e-08
| 1 | 0.05 | 0.2 | 0.05 | 0.1 | 0 | 20 | 10 | 0 | 49.9981086E+00 | DC220MV | DC220MV | 4.9348e-06 | 3.9575e-08
| 1 | 0.05 | 0.2 | 0 | 0.1 | 0.05 | 20 | 10 | 0 | 49.9734151E+00 | DC220MV | DC220MV | -0.0099951 | 3.8644e-08
| 1 | 0.1 | 0.2 | 0.05 | 0.1 | 0.05 | 20 | 10 | 0 | 50.0092753E+00 | DC220MV | DC220MV | -0.0099946 | 6.0552e-08
| 1 | 0.1 | 0.2 | 0 | 0.1 | 0.1 | 20 | 10 | 0 | 49.9634418E+00 | DC220MV | DC220MV | -0.019995 | 7.5968e-08
| 1 | 0.15 | 0.2 | 0.05 | 0.1 | 0.1 | 20 | 10 | 0 | 50.0123144E+00 | DC220MV | DC220MV | -0.019994 | 3.544e-08
| 1 | 0.15 | 0.2 | 0 | 0.1 | 0.15 | 20 | 10 | 0 | 49.9921569E+00 | DC220MV | DC220MV | -0.029994 | 4.0054e-08
| 1 | 0.2 | 0.2 | 0.05 | 0.1 | 0.15 | 20 | 10 | 0 | 49.9707859E+00 | DC220MV | DC220MV | -0.029994 | 3.9743e-08
| 1 | 0.2 | 0.2 | 0 | 0.1 | 0.2 | 20 | 10 | 0 | 50.0001604E+00 | DC220MV | DC220MV | -0.039994 | 5.294e-08
| 1 | 0.25 | 0.2 | 0.05 | 0.1 | 0.2 | 20 | 10 | 0 | 49.9567874E+00 | DC220MV | DC220MV | -0.039993 | 3.5819e-08
| 1 | 0.25 | 0.2 | 0 | 0.1 | 0.25 | 20 | 10 | 0 | 50.0067839E+00 | DC220MV | DC220MV | -0.049993 | 4.8468e-08
| 1 | 0.3 | 0.2 | 0.05 | 0.1 | 0.25 | 20 | 10 | 0 | 49.9751774E+00 | DC220MV | DC220MV | -0.049993 | 3.4955e-08
| 1 | 0.3 | 0.2 | 0 | 0.1 | 0.3 | 20 | 10 | 0 | 49.9859945E+00 | DC220MV | DC220MV | -0.059993 | 4.5685e-08
| 1 | 0.35 | 0.2 | 0.05 | 0.1 | 0.3 | 20 | 10 | 0 | 50.0301172E+00 | DC220MV | DC220MV | -0.059993 | 3.6657e-08
| 1 | 0.35 | 0.2 | 0 | 0.1 | 0.35 | 20 | 10 | 0 | 50.0593676E+00 | DC220MV | DC220MV | -0.069993 | 5.076e-08
| 1 | 0.4 | 0.2 | 0.05 | 0.1 | 0.35 | 20 | 10 | 0 | 50.0273924E+00 | DC220MV | DC220MV | -0.069992 | 4.5716e-08
| 1 | 0.4 | 0.2 | 0 | 0.1 | 0.4 | 20 | 10 | 0 | 50.0151697E+00 | DC220MV | DC220MV | -0.079993 | 5.4041e-08
| 1 | 0.45 | 0.2 | 0.05 | 0.1 | 0.4 | 20 | 10 | 0 | 49.9966182E+00 | DC220MV | DC220MV | -0.079992 | 4.7165e-08
| 1 | 0.45 | 0.2 | 0 | 0.1 | 0.45 | 20 | 10 | 0 | 49.9998443E+00 | DC220MV | DC220MV | -0.089992 | 4.5394e-08
| 1 | 0.5 | 0.2 | 0.05 | 0.1 | 0.45 | 20 | 10 | 0 | 50.0308634E+00 | DC220MV | DC220MV | -0.089992 | 2.9758e-08
| 1 | 0.5 | 0.2 | 0 | 0.1 | 0.5 | 20 | 10 | 0 | 49.9736742E+00 | DC220MV | DC220MV | -0.099992 | 4.8841e-08
| 1 | 0.5 | 0.2 | 0.05 | 0.1 | 0.45 | 20 | 10 | 0 | 50.0024577E+00 | DC220MV | DC220MV | -0.089992 | 3.9385e-08
| 1 | 0.45 | 0.2 | 0 | 0.1 | 0.45 | 20 | 10 | 0 | 50.0028256E+00 | DC220MV | DC220MV | -0.089992 | 7.4541e-08
| 1 | 0.45 | 0.2 | 0.05 | 0.1 | 0.4 | 20 | 10 | 0 | 50.0306981E+00 | DC220MV | DC220MV | -0.079992 | 3.2109e-08
| 1 | 0.4 | 0.2 | 0 | 0.1 | 0.4 | 20 | 10 | 0 | 49.9751397E+00 | DC220MV | DC220MV | -0.079993 | 5.3007e-08
| 1 | 0.4 | 0.2 | 0.05 | 0.1 | 0.35 | 20 | 10 | 0 | 50.0084212E+00 | DC220MV | DC220MV | -0.069993 | 4.2062e-08
| 1 | 0.35 | 0.2 | 0 | 0.1 | 0.35 | 20 | 10 | 0 | 49.9834489E+00 | DC220MV | DC220MV | -0.069993 | 3.84e-08
| 1 | 0.35 | 0.2 | 0.05 | 0.1 | 0.3 | 20 | 10 | 0 | 49.9227751E+00 | DC220MV | DC220MV | -0.059993 | 4.2363e-08
| 1 | 0.3 | 0.2 | 0 | 0.1 | 0.3 | 20 | 10 | 0 | 49.9635784E+00 | DC220MV | DC220MV | -0.059993 | 4.3653e-08
| 1 | 0.3 | 0.2 | 0.05 | 0.1 | 0.25 | 20 | 10 | 0 | 49.9812806E+00 | DC220MV | DC220MV | -0.049993 | 6.0246e-08
| 1 | 0.25 | 0.2 | 0 | 0.1 | 0.25 | 20 | 10 | 0 | 49.9246373E+00 | DC220MV | DC220MV | -0.049993 | 3.7076e-08
| 1 | 0.25 | 0.2 | 0.05 | 0.1 | 0.2 | 20 | 10 | 0 | 50.0779657E+00 | DC220MV | DC220MV | -0.039993 | 4.2491e-08
| 1 | 0.2 | 0.2 | 0 | 0.1 | 0.2 | 20 | 10 | 0 | 49.9783961E+00 | DC220MV | DC220MV | -0.039994 | 4.7936e-08
| 1 | 0.2 | 0.2 | 0.05 | 0.1 | 0.15 | 20 | 10 | 0 | 49.9276849E+00 | DC220MV | DC220MV | -0.029994 | 6.0326e-08
| 1 | 0.15 | 0.2 | 0 | 0.1 | 0.15 | 20 | 10 | 0 | 49.9834819E+00 | DC220MV | DC220MV | -0.029994 | 2.7836e-08
| 1 | 0.15 | 0.2 | 0.05 | 0.1 | 0.1 | 20 | 10 | 0 | 49.9816624E+00 | DC220MV | DC2_2V | -0.019994 | 4.6542e-08
| 1 | 0.1 | 0.2 | 0 | 0.1 | 0.1 | 20 | 10 | 0 | 49.9695751E+00 | DC220MV | DC2_2V | -0.019995 | 4.1285e-08
| 1 | 0.1 | 0.2 | 0.05 | 0.1 | 0.05 | 20 | 10 | 0 | 50.0398430E+00 | DC220MV | DC2_2V | -0.0099947 | 4.8006e-08
| 1 | 0.05 | 0.2 | 0 | 0.1 | 0.05 | 20 | 10 | 0 | 50.0250739E+00 | DC220MV | DC2_2V | -0.0099951 | 4.321e-08
| 1 | 0.05 | 0.2 | 0.05 | 0.1 | 0 | 20 | 10 | 0 | 50.0171096E+00 | DC220MV | DC2_2V | 4.9641e-06 | 3.9357e-08
| 1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9844105E+00 | DC220MV | DC2_2V | 4.5697e-06 | 7.0682e-08
| 1 | 0.05 | 0.2 | 0.05 | 0.1 | 0 | 20 | 10 | 0 | 50.0360873E+00 | DC220MV | DC2_2V | 4.6397e-06 | 3.1806e-08
| 1 | 0 | 0.2 | 0 |  | 0 | 20 | 10 | 0 | 50.0360873E+00 | DC220MV | DC2_2V | 4.6397e-06 | 3.1806e-08
| 1 | 0.05 | 0.2 | 0 |  | 0.05 | 20 | 10 | 0 | 50.0360873E+00 | DC220MV | DC2_2V | 4.6397e-06 | 3.1806e-08
| 1 | -0.005 | 0.2 | 0 |  | -0.005 | 20 | 10 | 0 | 50.0360873E+00 | DC220MV | DC2_2V | 4.6397e-06 | 3.1806e-08
| 1 | 0 | 0.2 | 0 |  | 0 | 20 | 10 | 0 | 50.0360873E+00 | DC220MV | DC2_2V | 4.6397e-06 | 3.1806e-08
| 1 | -0.005 | 0.2 | -0.005 | 0.1 | 0 | 20 | 10 | 0 | 49.9919258E+00 | DC220MV | DC2_2V | -0.009995 | 4.4132e-08
| 1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0050854E+00 | DC220MV | DC2_2V | 0.1 | 3.0148e-08
| 1 | -0.005 | 0.2 | -0.005 | 0.1 | 0 | 20 | 10 | 0 | 50.0059394E+00 | DC220MV | DC2_2V | 4.6034e-06 | 3.6841e-08
| 1 | -0.005 | 0.2 | 0 | 0.1 | -0.005 | 20 | 10 | 0 | 50.0113469E+00 | DC220MV | DC2_2V | 3.8322e-07 | 3.3318e-08
| 1 | -0.01 | 0.2 | -0.005 | 0.1 | -0.005 | 20 | 10 | 0 | 49.9915014E+00 | DC220MV | DC2_2V | 4.4662e-06 | 4.7154e-08
| 1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0096103E+00 | DC220MV | DC2_2V | 0.010005 | 5.2663e-08
| 1 | -0.015 | 0.2 | -0.005 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0684095E+00 | DC220MV | DC2_2V | 5.6183e-06 | 3.6452e-08
| 1 | -0.015 | 0.2 | 0 | 0.1 | -0.015 | 20 | 10 | 0 | 50.0477106E+00 | DC220MV | DC2_2V | 5.1692e-06 | 4.3858e-08
| 1 | -0.02 | 0.2 | -0.005 | 0.1 | -0.015 | 20 | 10 | 0 | 49.9717565E+00 | DC220MV | DC2_2V | 5.5943e-06 | 3.7356e-08
| 1 | -0.02 | 0.2 | 0 | 0.1 | -0.02 | 20 | 10 | 0 | 50.0544838E+00 | DC220MV | DC2_2V | 5.2605e-06 | 3.5703e-08
| 1 | -0.025 | 0.2 | -0.005 | 0.1 | -0.02 | 20 | 10 | 0 | 50.0214290E+00 | DC220MV | DC2_2V | 0.010005 | 3.5711e-08
| 1 | -0.025 | 0.2 | 0 | 0.1 | -0.025 | 20 | 10 | 0 | 49.9219569E+00 | DC220MV | DC2_2V | 0.010005 | 3.7356e-08
| 1 | -0.03 | 0.2 | -0.005 | 0.1 | -0.025 | 20 | 10 | 0 | 50.0338814E+00 | DC220MV | DC2_2V | 0.020005 | 4.5103e-08
| 1 | -0.03 | 0.2 | 0 | 0.1 | -0.03 | 20 | 10 | 0 | 50.1087360E+00 | DC220MV | DC2_2V | 0.020005 | 5.1348e-08
| 1 | -0.035 | 0.2 | -0.005 | 0.1 | -0.03 | 20 | 10 | 0 | 50.0456029E+00 | DC220MV | DC2_2V | 0.030005 | 4.6575e-08
| 1 | -0.035 | 0.2 | 0 | 0.1 | -0.035 | 20 | 10 | 0 | 49.9975662E+00 | DC220MV | DC2_2V | 0.030004 | 3.8575e-08
| 1 | -0.04 | 0.2 | -0.005 | 0.1 | -0.035 | 20 | 10 | 0 | 50.0492891E+00 | DC220MV | DC2_2V | 0.040004 | 4.146e-08
| 1 | -0.04 | 0.2 | 0 | 0.1 | -0.04 | 20 | 10 | 0 | 49.9829162E+00 | DC220MV | DC2_2V | 0.040004 | 5.3093e-08
| 1 | -0.045 | 0.2 | -0.005 | 0.1 | -0.04 | 20 | 10 | 0 | 49.9883046E+00 | DC220MV | DC2_2V | 0.050004 | 4.4105e-08
| 1 | -0.045 | 0.2 | 0 | 0.1 | -0.045 | 20 | 10 | 0 | 49.9992547E+00 | DC220MV | DC2_2V | 0.050003 | 4.989e-08
| 1 | -0.05 | 0.2 | -0.005 | 0.1 | -0.045 | 20 | 10 | 0 | 49.9908602E+00 | DC220MV | DC2_2V | 0.060003 | 4.2866e-08
| 1 | -0.05 | 0.2 | 0 | 0.1 | -0.05 | 20 | 10 | 0 | 49.9872721E+00 | DC220MV | DC2_2V | 0.060003 | 4.4228e-08
| 1 | -0.05 | 0.2 | -0.005 | 0.1 | -0.045 | 20 | 10 | 0 | 50.0022878E+00 | DC220MV | DC2_2V | 0.070003 | 7.9555e-08
| 1 | -0.045 | 0.2 | 0 | 0.1 | -0.045 | 20 | 10 | 0 | 49.9996226E+00 | DC220MV | DC2_2V | 0.070002 | 4.1837e-08
| 1 | -0.045 | 0.2 | -0.005 | 0.1 | -0.04 | 20 | 10 | 0 | 50.0800194E+00 | DC220MV | DC2_2V | 0.080002 | 7.1368e-08
| 1 | -0.04 | 0.2 | 0 | 0.1 | -0.04 | 20 | 10 | 0 | 49.9637951E+00 | DC220MV | DC2_2V | 0.080002 | 5.1504e-08
| 1 | -0.04 | 0.2 | -0.005 | 0.1 | -0.035 | 20 | 10 | 0 | 49.9802531E+00 | DC220MV | DC2_2V | 0.090002 | 4.66e-08
| 1 | -0.035 | 0.2 | 0 | 0.1 | -0.035 | 20 | 10 | 0 | 49.9646712E+00 | DC220MV | DC2_2V | 0.090002 | 3.2379e-08
| 1 | -0.035 | 0.2 | -0.005 | 0.1 | -0.03 | 20 | 10 | 0 | 49.9537362E+00 | DC220MV | DC2_2V | 0.1 | 5.4103e-08
| 1 | -0.03 | 0.2 | 0 | 0.1 | -0.03 | 20 | 10 | 0 | 49.9582519E+00 | DC220MV | DC2_2V | 0.090002 | 5.163e-08
| 1 | -0.03 | 0.2 | -0.005 | 0.1 | -0.025 | 20 | 10 | 0 | 50.0248756E+00 | DC220MV | DC2_2V | 0.090002 | 5.37e-07
| 1 | -0.025 | 0.2 | 0 | 0.1 | -0.025 | 20 | 10 | 0 | 50.0470206E+00 | DC220MV | DC2_2V | 0.080002 | 4.9294e-08
| 1 | -0.025 | 0.2 | -0.005 | 0.1 | -0.02 | 20 | 10 | 0 | 50.0172276E+00 | DC220MV | DC2_2V | 0.080003 | 4.2433e-08
| 1 | -0.02 | 0.2 | 0 | 0.1 | -0.02 | 20 | 10 | 0 | 49.9859143E+00 | DC220MV | DC220MV | 0.070003 | 5.5564e-08
| 1 | -0.02 | 0.2 | -0.005 | 0.1 | -0.015 | 20 | 10 | 0 | 50.0120690E+00 | DC220MV | DC220MV | 0.070003 | 4.4479e-08
| 1 | -0.015 | 0.2 | 0 | 0.1 | -0.015 | 20 | 10 | 0 | 50.0356528E+00 | DC220MV | DC220MV | 0.060003 | 3.8506e-08
| 1 | -0.015 | 0.2 | -0.005 | 0.1 | -0.01 | 20 | 10 | 0 | 49.9932603E+00 | DC220MV | DC220MV | 0.060004 | 5.6544e-08
| 1 | -0.01 | 0.2 | 0 | 0.1 | -0.01 | 20 | 10 | 0 | 50.0500075E+00 | DC220MV | DC220MV | 0.050004 | 4.5149e-08
| 1 | -0.01 | 0.2 | -0.005 | 0.1 | -0.005 | 20 | 10 | 0 | 50.0102144E+00 | DC220MV | DC220MV | 0.050004 | 5.6065e-08
| 1 | -0.005 | 0.2 | 0 | 0.1 | -0.005 | 20 | 10 | 0 | 49.9887808E+00 | DC220MV | DC2_2V | 0.040004 | 6.4814e-08
| 1 | -0.005 | 0.2 | -0.005 | 0.1 | 0 | 20 | 10 | 0 | 50.0200505E+00 | DC220MV | DC2_2V | 0.040005 | 4.3282e-08
| 1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 49.9935952E+00 | DC220MV | DC2_2V | 0.030005 | 3.3781e-08
| 1 | -0.005 | 0.2 | -0.005 | 0.1 | 0 | 20 | 10 | 0 | 50.0334138E+00 | DC220MV | DC2_2V | 0.030005 | 4.149e-08
| 1 | 0 | 0.2 | 0 | 0.1 | 0 | 20 | 10 | 0 | 50.0612684E+00 | DC220MV | DC2_2V | 0.020005 | 5.0905e-08
| 1 | -0.005 | 0.2 | 0 | 0.1 | -0.005 | 20 | 10 | 0 | 50.0152783E+00 | DC220MV | DC2_2V | 0.020005 | 3.9469e-08
| 1 | 0.1 | 2 | 0 | 1 | 0.1 | 20 | 10 | 0 | 49.9900256E+00 | DC220MV | DC2_2V | 0.010005 | 3.5281e-08
| 1 | 0 | 2 | 0 | 1 | 0 | 20 | 10 | 0 | 49.9713042E+00 | DC220MV | DC2_2V | 0.010006 | 3.0983e-08
| 1 | 0.1 | 2 | 0.1 | 1 | 0 | 20 | 10 | 0 | 49.9393284E+00 | DC220MV | DC2_2V | 5.4446e-06 | 5.5812e-08
| 1 | 0 | 2 | 0 | 1 | 0 | 20 | 10 | 0 | 49.9646194E+00 | DC220MV | DC2_2V | 5.4484e-06 | 8.7916e-08
| 1 | 0.1 | 2 | 0.1 | 1 | 0 | 20 | 10 | 0 | 49.9501535E+00 | DC220MV | DC2_2V | 4.7549e-06 | 6.0539e-08
| 1 | 0.1 | 2 | 0 | 1 | 0.1 | 20 | 10 | 0 | 50.0500501E+00 | DC220MV | DC2_2V | 4.8223e-06 | 5.8733e-08
| 1 | 0.2 | 2 | 0.1 | 1 | 0.1 | 20 | 10 | 0 | 49.9714738E+00 | DC220MV | DC2_2V | 0.010004 | 4.5554e-08
| 1 | 0.2 | 2 | 0 | 1 | 0.2 | 20 | 10 | 0 | 50.0005236E+00 | DC220MV | DC2_2V | -0.0099949 | 2.9958e-08
| 1 | 0.3 | 2 | 0.1 | 1 | 0.2 | 20 | 10 | 0 | 50.0048354E+00 | DC220MV | DC2_2V | 4.6438e-06 | 6.5219e-08
| 1 | 0.3 | 2 | 0 | 1 | 0.3 | 20 | 10 | 0 | 50.0362243E+00 | DC220MV | DC2_2V | 5.0585e-06 | 2.973e-08
| 1 | 0.4 | 2 | 0.1 | 1 | 0.3 | 20 | 10 | 0 | 49.9836091E+00 | DC220MV | DC2_2V | 4.5116e-06 | 4.2838e-08
| 1 | 0.4 | 2 | 0 | 1 | 0.4 | 20 | 10 | 0 | 50.0367298E+00 | DC220MV | DC2_2V | 4.9348e-06 | 3.9575e-08
| 1 | 0.5 | 2 | 0.1 | 1 | 0.4 | 20 | 10 | 0 | 49.9842738E+00 | DC220MV | DC2_2V | -0.0099951 | 3.8644e-08
| 1 | 0.5 | 2 | 0 | 1 | 0.5 | 20 | 10 | 0 | 49.9949250E+00 | DC220MV | DC2_2V | -0.0099946 | 6.0552e-08
| 1 | 0.6 | 2 | 0.1 | 1 | 0.5 | 20 | 10 | 0 | 49.9986793E+00 | DC220MV | DC2_2V | -0.019995 | 7.5968e-08
| 1 | 0.6 | 2 | 0 | 1 | 0.6 | 20 | 10 | 0 | 49.9566037E+00 | DC220MV | DC2_2V | -0.019994 | 3.544e-08
| 1 | 0.7 | 2 | 0.1 | 1 | 0.6 | 20 | 10 | 0 | 50.1014460E+00 | DC220MV | DC2_2V | -0.029994 | 4.0054e-08
| 1 | 0.7 | 2 | 0 | 1 | 0.7 | 20 | 10 | 0 | 49.9959437E+00 | DC220MV | DC2_2V | -0.029994 | 3.9743e-08
| 1 | 0.8 | 2 | 0.1 | 1 | 0.7 | 20 | 10 | 0 | 50.0459715E+00 | DC220MV | DC2_2V | -0.039994 | 5.294e-08
| 1 | 0.8 | 2 | 0 | 1 | 0.8 | 20 | 10 | 0 | 49.9684774E+00 | DC220MV | DC2_2V | -0.039993 | 3.5819e-08
| 1 | 0.9 | 2 | 0.1 | 1 | 0.8 | 20 | 10 | 0 | 49.9719874E+00 | DC220MV | DC2_2V | -0.049993 | 4.8468e-08
| 1 | 0.9 | 2 | 0 | 1 | 0.9 | 20 | 10 | 0 | 49.9837600E+00 | DC220MV | DC2_2V | -0.049993 | 3.4955e-08
| 1 | 1 | 2 | 0.1 | 1 | 0.9 | 20 | 10 | 0 | 50.0258577E+00 | DC220MV | DC2_2V | -0.059993 | 4.5685e-08
| 1 | 1 | 2 | 0 | 1 | 1 | 20 | 10 | 0 | 49.9915910E+00 | DC220MV | DC2_2V | -0.059993 | 3.6657e-08
| 1 | 1 | 2 | 0.1 | 1 | 0.9 | 20 | 10 | 0 | 50.0515247E+00 | DC220MV | DC2_2V | -0.069993 | 5.076e-08
| 1 | 0.9 | 2 | 0 | 1 | 0.9 | 20 | 10 | 0 | 50.0354119E+00 | DC220MV | DC2_2V | -0.069992 | 4.5716e-08
| 1 | 0.9 | 2 | 0.1 | 1 | 0.8 | 20 | 10 | 0 | 50.0273499E+00 | DC220MV | DC2_2V | -0.079993 | 5.4041e-08
| 1 | 0.8 | 2 | 0 | 1 | 0.8 | 20 | 10 | 0 | 49.9984812E+00 | DC220MV | DC2_2V | -0.079992 | 4.7165e-08
| 1 | 0.8 | 2 | 0.1 | 1 | 0.7 | 20 | 10 | 0 | 49.9898417E+00 | DC220MV | DC2_2V | -0.089992 | 4.5394e-08
| 1 | 0.7 | 2 | 0 | 1 | 0.7 | 20 | 10 | 0 | 49.9354655E+00 | DC220MV | DC2_2V | -0.089992 | 2.9758e-08
| 1 | 0.7 | 2 | 0.1 | 1 | 0.6 | 20 | 10 | 0 | 49.9777363E+00 | DC220MV | DC2_2V | -0.099992 | 4.8841e-08
| 1 | 0.6 | 2 | 0 | 1 | 0.6 | 20 | 10 | 0 | 50.0357094E+00 | DC220MV | DC2_2V | -0.089992 | 3.9385e-08
| 1 | 0.6 | 2 | 0.1 | 1 | 0.5 | 20 | 10 | 0 | 49.9519283E+00 | DC220MV | DC2_2V | -0.089992 | 7.4541e-08
| 1 | 0.5 | 2 | 0 | 1 | 0.5 | 20 | 10 | 0 | 49.9450228E+00 | DC220MV | DC2_2V | -0.079992 | 3.2109e-08
| 1 | 0.5 | 2 | 0.1 | 1 | 0.4 | 20 | 10 | 0 | 50.0212165E+00 | DC220MV | DC2_2V | -0.079993 | 5.3007e-08
| 1 | 0.4 | 2 | 0 | 1 | 0.4 | 20 | 10 | 0 | 49.9804463E+00 | DC220MV | DC220MV | -0.069993 | 4.2062e-08
| 1 | 0.4 | 2 | 0.1 | 1 | 0.3 | 20 | 10 | 0 | 49.9891109E+00 | DC220MV | DC220MV | -0.069993 | 3.84e-08
| 1 | 0.3 | 2 | 0 | 1 | 0.3 | 20 | 10 | 0 | 49.9998066E+00 | DC220MV | DC220MV | -0.059993 | 4.2363e-08
| 1 | 0.3 | 2 | 0.1 | 1 | 0.2 | 20 | 10 | 0 | 50.0422385E+00 | DC2_2V | DC2_2V | -0.059993 | 4.3653e-08
| 1 | 0.2 | 2 | 0 | 1 | 0.2 | 20 | 10 | 0 | 49.9700039E+00 | DC2_2V | DC2_2V | -0.049993 | 6.0246e-08
| 1 | 0.2 | 2 | 0.1 | 1 | 0.1 | 20 | 10 | 0 | 50.0165526E+00 | DC2_2V | DC2_2V | -0.049993 | 3.7076e-08
| 1 | 0.1 | 2 | 0 | 1 | 0.1 | 20 | 10 | 0 | 49.9795602E+00 | DC2_2V | DC11V | -0.039993 | 4.2491e-08
| 1 | 0.1 | 2 | 0.1 | 1 | 0 | 20 | 10 | 0 | 49.8893445E+00 | DC2_2V | DC11V | -0.039994 | 4.7936e-08
| 1 | 0 | 2 | 0 | 1 | 0 | 20 | 10 | 0 | 49.9553088E+00 | DC2_2V | DC11V | -0.029994 | 6.0326e-08
| 1 | 0.1 | 2 | 0.1 | 1 | 0 | 20 | 10 | 0 | 50.0452768E+00 | DC2_2V | DC11V | -0.029994 | 2.7836e-08
| 1 | 0 | 2 | 0 | 1 | 0 | 20 | 10 | 0 | 49.9977501E+00 | DC2_2V | DC11V | -0.019994 | 4.6542e-08
| 1 | 0.1 | 2 | 0.1 | 1 | 0 | 20 | 10 | 0 | 50.0547486E+00 | DC2_2V | DC11V | -0.019995 | 4.1285e-08
| 1 | -0.1 | 2 | -0.1 | 1 | 0 | 20 | 10 | 0 | 50.0525741E+00 | DC2_2V | DC11V | -0.0099947 | 4.8006e-08
| 1 | 0 | 2 | 0 | 1 | 0 | 20 | 10 | 0 | 49.9871212E+00 | DC2_2V | DC11V | -0.0099951 | 4.321e-08
| 1 | -0.1 | 2 | -0.1 | 1 | 0 | 20 | 10 | 0 | 49.9714785E+00 | DC2_2V | DC11V | 4.9641e-06 | 3.9357e-08
| 1 | 0 | 2 | 0 | 1 | 0 | 20 | 10 | 0 | 50.0065527E+00 | DC2_2V | DC11V | 4.5697e-06 | 7.0682e-08
| 1 | -0.1 | 2 | -0.1 | 1 | 0 | 20 | 10 | 0 | 49.9839957E+00 | DC2_2V | DC11V | 5.163e-06 | 4.253e-08
| 1 | -0.1 | 2 | 0 | 1 | -0.1 | 20 | 10 | 0 | 49.9931142E+00 | DC2_2V | DC11V | 4.6397e-06 | 3.1806e-08
| 1 | -0.2 | 2 | -0.1 | 1 | -0.1 | 20 | 10 | 0 | 49.9825250E+00 | DC2_2V | DC11V | -0.009995 | 4.4132e-08
| 1 | -0.2 | 2 | 0 | 1 | -0.2 | 20 | 10 | 0 | 50.0452579E+00 | DC2_2V | DC11V | 0.1 | 3.0148e-08
| 1 | -0.3 | 2 | -0.1 | 1 | -0.2 | 20 | 10 | 0 | 49.8609679E+00 | DC2_2V | DC11V | 4.6034e-06 | 3.6841e-08
| 1 | -0.3 | 2 | 0 | 1 | -0.3 | 20 | 10 | 0 | 49.9730900E+00 | DC2_2V | DC11V | 3.8322e-07 | 3.3318e-08
| 1 | -0.4 | 2 | -0.1 | 1 | -0.3 | 20 | 10 | 0 | 50.0250172E+00 | DC2_2V | DC11V | 4.4662e-06 | 4.7154e-08
| 1 | -0.4 | 2 | 0 | 1 | -0.4 | 20 | 10 | 0 | 50.0696488E+00 | DC2_2V | DC11V | 0.010005 | 5.2663e-08
| 1 | -0.5 | 2 | -0.1 | 1 | -0.4 | 20 | 10 | 0 | 50.0135508E+00 | DC2_2V | DC11V | 5.6183e-06 | 3.6452e-08
| 1 | -0.5 | 2 | 0 | 1 | -0.5 | 20 | 10 | 0 | 49.9570887E+00 | DC2_2V | DC11V | 5.1692e-06 | 4.3858e-08
| 1 | -0.6 | 2 | -0.1 | 1 | -0.5 | 20 | 10 | 0 | 49.9908696E+00 | DC2_2V | DC11V | 5.5943e-06 | 3.7356e-08
| 1 | -0.6 | 2 | 0 | 1 | -0.6 | 20 | 10 | 0 | 49.9654861E+00 | DC2_2V | DC11V | 5.2605e-06 | 3.5703e-08
| 1 | -0.7 | 2 | -0.1 | 1 | -0.6 | 20 | 10 | 0 | 50.0630369E+00 | DC2_2V | DC11V | 0.010005 | 3.5711e-08
| 1 | -0.7 | 2 | 0 | 1 | -0.7 | 20 | 10 | 0 | 50.0308728E+00 | DC2_2V | DC11V | 0.010005 | 3.7356e-08
| 1 | -0.8 | 2 | -0.1 | 1 | -0.7 | 20 | 10 | 0 | 50.0105777E+00 | DC2_2V | DC11V | 0.020005 | 4.5103e-08
| 1 | -0.8 | 2 | 0 | 1 | -0.8 | 20 | 10 | 0 | 50.0179546E+00 | DC2_2V | DC11V | 0.020005 | 5.1348e-08
| 1 | -0.9 | 2 | -0.1 | 1 | -0.8 | 20 | 10 | 0 | 49.9165076E+00 | DC2_2V | DC11V | 0.030005 | 4.6575e-08
| 1 | -0.9 | 2 | 0 | 1 | -0.9 | 20 | 10 | 0 | 49.9685198E+00 | DC2_2V | DC11V | 0.030004 | 3.8575e-08
| 1 | -1 | 2 | -0.1 | 1 | -0.9 | 20 | 10 | 0 | 49.9781322E+00 | DC2_2V | DC11V | 0.040004 | 4.146e-08
| 1 | -1 | 2 | 0 | 1 | -1 | 20 | 10 | 0 | 49.9635549E+00 | DC2_2V | DC11V | 0.040004 | 5.3093e-08
| 1 | -1 | 2 | -0.1 | 1 | -0.9 | 20 | 10 | 0 | 49.9867064E+00 | DC2_2V | DC11V | 0.050004 | 4.4105e-08
| 1 | -0.9 | 2 | 0 | 1 | -0.9 | 20 | 10 | 0 | 49.9604323E+00 | DC2_2V | DC11V | 0.050003 | 4.989e-08
| 1 | -0.9 | 2 | -0.1 | 1 | -0.8 | 20 | 10 | 0 | 49.9561470E+00 | DC2_2V | DC11V | 0.060003 | 4.2866e-08
| 1 | -0.8 | 2 | 0 | 1 | -0.8 | 20 | 10 | 0 | 50.0151980E+00 | DC2_2V | DC11V | 0.060003 | 4.4228e-08
| 1 | -0.8 | 2 | -0.1 | 1 | -0.7 | 20 | 10 | 0 | 50.0420353E+00 | DC2_2V | DC11V | 0.070003 | 7.9555e-08
| 1 | -0.7 | 2 | 0 | 1 | -0.7 | 20 | 10 | 0 | 49.9588546E+00 | DC2_2V | DC11V | 0.070002 | 4.1837e-08
| 1 | -0.7 | 2 | -0.1 | 1 | -0.6 | 20 | 10 | 0 | 49.9365852E+00 | DC2_2V | DC11V | 0.080002 | 7.1368e-08
| 1 | -0.6 | 2 | 0 | 1 | -0.6 | 20 | 10 | 0 | 49.9921333E+00 | DC2_2V | DC11V | 0.080002 | 5.1504e-08
| 1 | -0.6 | 2 | -0.1 | 1 | -0.5 | 20 | 10 | 0 | 50.0283699E+00 | DC2_2V | DC11V | 0.090002 | 4.66e-08
| 1 | -0.5 | 2 | 0 | 1 | -0.5 | 20 | 10 | 0 | 50.0458392E+00 | DC2_2V | DC11V | 0.090002 | 3.2379e-08
| 1 | -0.5 | 2 | -0.1 | 1 | -0.4 | 20 | 10 | 0 | 50.0737454E+00 | DC2_2V | DC11V | 0.1 | 5.4103e-08
| 1 | -0.4 | 2 | 0 | 1 | -0.4 | 20 | 10 | 0 | 50.0049297E+00 | DC2_2V | DC2_2V | 0.090002 | 5.163e-08
| 1 | -0.4 | 2 | -0.1 | 1 | -0.3 | 20 | 10 | 0 | 49.9964012E+00 | DC2_2V | DC2_2V | 0.090002 | 5.37e-07
| 1 | -0.3 | 2 | 0 | 1 | -0.3 | 20 | 10 | 0 | 49.9888138E+00 | DC2_2V | DC2_2V | 0.080002 | 4.9294e-08
| 1 | -0.3 | 2 | -0.1 | 1 | -0.2 | 20 | 10 | 0 | 49.9775714E+00 | DC2_2V | DC2_2V | 0.080003 | 4.2433e-08
| 1 | -0.2 | 2 | 0 | 1 | -0.2 | 20 | 10 | 0 | 50.0598120E+00 | DC2_2V | DC2_2V | 0.070003 | 5.5564e-08
| 1 | -0.2 | 2 | -0.1 | 1 | -0.1 | 20 | 10 | 0 | 49.9908837E+00 | DC2_2V | DC2_2V | 0.070003 | 4.4479e-08
| 1 | -0.1 | 2 | 0 | 1 | -0.1 | 20 | 10 | 0 | 50.0086666E+00 | DC2_2V | DC11V | 0.060003 | 3.8506e-08
| 1 | -0.1 | 2 | -0.1 | 1 | 0 | 20 | 10 | 0 | 49.9911761E+00 | DC2_2V | DC11V | 0.060004 | 5.6544e-08
| 1 | 0 | 2 | 0 | 1 | 0 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 1 | -0.1 | 2 | -0.1 | 1 | 0 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 1 | 0 | 2 | 0 | 1 | 0 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 1 | -0.1 | 2 | 0 | 1 | -0.1 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 10 | 0.05 | 0.2 | 0 | 10 | 0.05 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 10 | 0.05 | 0.2 | 0.05 | 10 | 0 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9616662E+00 | DC2_2V | DC11V | 0.030005 | 3.3781e-08
| 10 | 0.05 | 0.2 | 0.05 | 10 | 0 | 20 | 10 | 0 | 49.9760822E+00 | DC2_2V | DC11V | 0.030005 | 4.149e-08
| 10 | 0.05 | 0.2 | 0 | 10 | 0.05 | 20 | 10 | 0 | 50.0089497E+00 | DC2_2V | DC11V | 0.020005 | 5.0905e-08
| 10 | 0.1 | 0.2 | 0.05 | 10 | 0.05 | 20 | 10 | 0 | 50.0098227E+00 | DC2_2V | DC11V | 0.020005 | 3.9469e-08
| 10 | 0.1 | 0.2 | 0 | 10 | 0.1 | 20 | 10 | 0 | 50.0548715E+00 | DC2_2V | DC11V | 0.010005 | 3.5281e-08
| 10 | 0.15 | 0.2 | 0.05 | 10 | 0.1 | 20 | 10 | 0 | 50.0107523E+00 | DC2_2V | DC11V | 0.010006 | 3.0983e-08
| 10 | 0.15 | 0.2 | 0 | 10 | 0.15 | 20 | 10 | 0 | 50.0186626E+00 | DC2_2V | DC11V | 5.4446e-06 | 5.5812e-08
| 10 | 0.2 | 0.2 | 0.05 | 10 | 0.15 | 20 | 10 | 0 | 50.0054346E+00 | DC2_2V | DC11V | 5.4484e-06 | 8.7916e-08
| 10 | 0.2 | 0.2 | 0 | 10 | 0.2 | 20 | 10 | 0 | 49.9804275E+00 | DC2_2V | DC11V | 4.7549e-06 | 6.0539e-08
| 10 | 0.25 | 0.2 | 0.05 | 10 | 0.2 | 20 | 10 | 0 | 50.0244601E+00 | DC2_2V | DC11V | 4.8223e-06 | 5.8733e-08
| 10 | 0.25 | 0.2 | 0 | 10 | 0.25 | 20 | 10 | 0 | 49.9897568E+00 | DC2_2V | DC11V | 0.010004 | 4.5554e-08
| 10 | 0.3 | 0.2 | 0.05 | 10 | 0.25 | 20 | 10 | 0 | 50.0130742E+00 | DC2_2V | DC11V | -0.0099949 | 2.9958e-08
| 10 | 0.3 | 0.2 | 0 | 10 | 0.3 | 20 | 10 | 0 | 49.9879746E+00 | DC2_2V | DC11V | 4.6438e-06 | 6.5219e-08
| 10 | 0.35 | 0.2 | 0.05 | 10 | 0.3 | 20 | 10 | 0 | 49.9643792E+00 | DC2_2V | DC11V | 5.0585e-06 | 2.973e-08
| 10 | 0.35 | 0.2 | 0 | 10 | 0.35 | 20 | 10 | 0 | 49.9486471E+00 | DC2_2V | DC11V | 4.5116e-06 | 4.2838e-08
| 10 | 0.4 | 0.2 | 0.05 | 10 | 0.35 | 20 | 10 | 0 | 49.9932603E+00 | DC2_2V | DC11V | 4.9348e-06 | 3.9575e-08
| 10 | 0.4 | 0.2 | 0 | 10 | 0.4 | 20 | 10 | 0 | 50.0016038E+00 | DC2_2V | DC11V | -0.0099951 | 3.8644e-08
| 10 | 0.45 | 0.2 | 0.05 | 10 | 0.4 | 20 | 10 | 0 | 50.0230484E+00 | DC2_2V | DC11V | -0.0099946 | 6.0552e-08
| 10 | 0.45 | 0.2 | 0 | 10 | 0.45 | 20 | 10 | 0 | 49.9626506E+00 | DC2_2V | DC11V | -0.019995 | 7.5968e-08
| 10 | 0.5 | 0.2 | 0.05 | 10 | 0.45 | 20 | 10 | 0 | 49.9980991E+00 | DC2_2V | DC11V | -0.019994 | 3.544e-08
| 10 | 0.5 | 0.2 | 0 | 10 | 0.5 | 20 | 10 | 0 | 49.9965851E+00 | DC2_2V | DC11V | -0.029994 | 4.0054e-08
| 10 | 0.5 | 0.2 | 0.05 | 10 | 0.45 | 20 | 10 | 0 | 49.9741407E+00 | DC2_2V | DC11V | -0.029994 | 3.9743e-08
| 10 | 0.45 | 0.2 | 0 | 10 | 0.45 | 20 | 10 | 0 | 50.0042975E+00 | DC2_2V | DC11V | -0.039994 | 5.294e-08
| 10 | 0.45 | 0.2 | 0.05 | 10 | 0.4 | 20 | 10 | 0 | 49.9474892E+00 | DC2_2V | DC11V | -0.039993 | 3.5819e-08
| 10 | 0.4 | 0.2 | 0 | 10 | 0.4 | 20 | 10 | 0 | 50.0383738E+00 | DC2_2V | DC11V | -0.049993 | 4.8468e-08
| 10 | 0.4 | 0.2 | 0.05 | 10 | 0.35 | 20 | 10 | 0 | 49.9884791E+00 | DC2_2V | DC11V | -0.049993 | 3.4955e-08
| 10 | 0.35 | 0.2 | 0 | 10 | 0.35 | 20 | 10 | 0 | 50.0185021E+00 | DC2_2V | DC11V | -0.059993 | 4.5685e-08
| 10 | 0.35 | 0.2 | 0.05 | 10 | 0.3 | 20 | 10 | 0 | 50.0553773E+00 | DC2_2V | DC11V | -0.059993 | 3.6657e-08
| 10 | 0.3 | 0.2 | 0 | 10 | 0.3 | 20 | 10 | 0 | 49.9801353E+00 | DC2_2V | DC11V | -0.069993 | 5.076e-08
| 10 | 0.3 | 0.2 | 0.05 | 10 | 0.25 | 20 | 10 | 0 | 49.9973822E+00 | DC2_2V | DC11V | -0.069992 | 4.5716e-08
| 10 | 0.25 | 0.2 | 0 | 10 | 0.25 | 20 | 10 | 0 | 50.0116820E+00 | DC2_2V | DC11V | -0.079993 | 5.4041e-08
| 10 | 0.25 | 0.2 | 0.05 | 10 | 0.2 | 20 | 10 | 0 | 50.0381470E+00 | DC2_2V | DC11V | -0.079992 | 4.7165e-08
| 10 | 0.2 | 0.2 | 0 | 10 | 0.2 | 20 | 10 | 0 | 49.9937932E+00 | DC2_2V | DC11V | -0.089992 | 4.5394e-08
| 10 | 0.2 | 0.2 | 0.05 | 10 | 0.15 | 20 | 10 | 0 | 50.0414069E+00 | DC2_2V | DC11V | -0.089992 | 2.9758e-08
| 10 | 0.15 | 0.2 | 0 | 10 | 0.15 | 20 | 10 | 0 | 50.0424417E+00 | DC2_2V | DC11V | -0.099992 | 4.8841e-08
| 10 | 0.15 | 0.2 | 0.05 | 10 | 0.1 | 20 | 10 | 0 | 49.9540752E+00 | DC2_2V | DC2_2V | -0.089992 | 3.9385e-08
| 10 | 0.1 | 0.2 | 0 | 10 | 0.1 | 20 | 10 | 0 | 49.9572441E+00 | DC2_2V | DC2_2V | -0.089992 | 7.4541e-08
| 10 | 0.1 | 0.2 | 0.05 | 10 | 0.05 | 20 | 10 | 0 | 50.0156511E+00 | DC2_2V | DC2_2V | -0.079992 | 3.2109e-08
| 10 | 0.05 | 0.2 | 0 | 10 | 0.05 | 20 | 10 | 0 | 50.0131308E+00 | DC220MV | DC220MV | -0.079993 | 5.3007e-08
| 10 | 0.05 | 0.2 | 0.05 | 10 | 0 | 20 | 10 | 0 | 49.9719827E+00 | DC220MV | DC220MV | -0.069993 | 4.2062e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0056421E+00 | DC220MV | DC220MV | -0.069993 | 3.84e-08
| 10 | 0.05 | 0.2 | 0.05 | 10 | 0 | 20 | 10 | 0 | 49.9625988E+00 | DC220MV | DC220MV | -0.059993 | 4.2363e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9625988E+00 | DC220MV | DC220MV | -0.059993 | 4.2363e-08
| 10 | 0.05 | 0.2 | 0 | 10 | 0.05 | 20 | 10 | 0 | 50.0264432E+00 | DC220MV | DC220MV | -0.059993 | 4.3653e-08
| 10 | -0.05 | 0.2 | 0 | 10 | -0.05 | 20 | 10 | 0 | 49.9625988E+00 | DC220MV | DC220MV | -0.059993 | 4.2363e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9625988E+00 | DC220MV | DC220MV | -0.059993 | 4.2363e-08
| 10 | -0.05 | 0.2 | -0.05 | 10 | 0 | 20 | 10 | 0 | 50.0096339E+00 | DC220MV | DC220MV | -0.049993 | 6.0246e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9681288E+00 | DC220MV | DC220MV | -0.049993 | 3.7076e-08
| 10 | -0.05 | 0.2 | -0.05 | 10 | 0 | 20 | 10 | 0 | 49.9975426E+00 | DC220MV | DC220MV | -0.039993 | 4.2491e-08
| 10 | -0.05 | 0.2 | 0 | 10 | -0.05 | 20 | 10 | 0 | 50.0329037E+00 | DC220MV | DC220MV | -0.039994 | 4.7936e-08
| 10 | -0.1 | 0.2 | -0.05 | 10 | -0.05 | 20 | 10 | 0 | 50.0029530E+00 | DC220MV | DC220MV | -0.029994 | 6.0326e-08
| 10 | -0.1 | 0.2 | 0 | 10 | -0.1 | 20 | 10 | 0 | 49.9937083E+00 | DC220MV | DC220MV | -0.029994 | 2.7836e-08
| 10 | -0.15 | 0.2 | -0.05 | 10 | -0.1 | 20 | 10 | 0 | 50.0149054E+00 | DC220MV | DC220MV | -0.019994 | 4.6542e-08
| 10 | -0.15 | 0.2 | 0 | 10 | -0.15 | 20 | 10 | 0 | 49.9931330E+00 | DC220MV | DC220MV | -0.019995 | 4.1285e-08
| 10 | -0.2 | 0.2 | -0.05 | 10 | -0.15 | 20 | 10 | 0 | 49.9917466E+00 | DC220MV | DC220MV | -0.0099947 | 4.8006e-08
| 10 | -0.2 | 0.2 | 0 | 10 | -0.2 | 20 | 10 | 0 | 50.0594385E+00 | DC220MV | DC220MV | -0.0099951 | 4.321e-08
| 10 | -0.25 | 0.2 | -0.05 | 10 | -0.2 | 20 | 10 | 0 | 49.9829162E+00 | DC220MV | DC220MV | 4.9641e-06 | 3.9357e-08
| 10 | -0.25 | 0.2 | 0 | 10 | -0.25 | 20 | 10 | 0 | 49.9492356E+00 | DC220MV | DC220MV | 4.5697e-06 | 7.0682e-08
| 10 | -0.3 | 0.2 | -0.05 | 10 | -0.25 | 20 | 10 | 0 | 50.0236433E+00 | DC220MV | DC220MV | 5.163e-06 | 4.253e-08
| 10 | -0.3 | 0.2 | 0 | 10 | -0.3 | 20 | 10 | 0 | 49.9746355E+00 | DC220MV | DC220MV | 4.6397e-06 | 3.1806e-08
| 10 | -0.35 | 0.2 | -0.05 | 10 | -0.3 | 20 | 10 | 0 | 49.9932698E+00 | DC220MV | DC220MV | -0.009995 | 4.4132e-08
| 10 | -0.35 | 0.2 | 0 | 10 | -0.35 | 20 | 10 | 0 | 49.9617510E+00 | DC220MV | DC220MV | 0.1 | 3.0148e-08
| 10 | -0.4 | 0.2 | -0.05 | 10 | -0.35 | 20 | 10 | 0 | 49.9962880E+00 | DC220MV | DC220MV | 4.6034e-06 | 3.6841e-08
| 10 | -0.4 | 0.2 | 0 | 10 | -0.4 | 20 | 10 | 0 | 49.9896012E+00 | DC220MV | DC220MV | 3.8322e-07 | 3.3318e-08
| 10 | -0.45 | 0.2 | -0.05 | 10 | -0.4 | 20 | 10 | 0 | 50.0031417E+00 | DC220MV | DC220MV | 4.4662e-06 | 4.7154e-08
| 10 | -0.45 | 0.2 | 0 | 10 | -0.45 | 20 | 10 | 0 | 49.9789805E+00 | DC220MV | DC220MV | 0.010005 | 5.2663e-08
| 10 | -0.5 | 0.2 | -0.05 | 10 | -0.45 | 20 | 10 | 0 | 50.0171238E+00 | DC220MV | DC220MV | 5.6183e-06 | 3.6452e-08
| 10 | -0.5 | 0.2 | 0 | 10 | -0.5 | 20 | 10 | 0 | 49.9998915E+00 | DC220MV | DC220MV | 5.1692e-06 | 4.3858e-08
| 10 | -0.5 | 0.2 | -0.05 | 10 | -0.45 | 20 | 10 | 0 | 49.9831943E+00 | DC220MV | DC220MV | 5.5943e-06 | 3.7356e-08
| 10 | -0.45 | 0.2 | 0 | 10 | -0.45 | 20 | 10 | 0 | 49.9900492E+00 | DC220MV | DC220MV | 5.2605e-06 | 3.5703e-08
| 10 | -0.45 | 0.2 | -0.05 | 10 | -0.4 | 20 | 10 | 0 | 49.9985614E+00 | DC220MV | DC220MV | 0.010005 | 3.5711e-08
| 10 | -0.4 | 0.2 | 0 | 10 | -0.4 | 20 | 10 | 0 | 50.0033021E+00 | DC220MV | DC220MV | 0.010005 | 3.7356e-08
| 10 | -0.4 | 0.2 | -0.05 | 10 | -0.35 | 20 | 10 | 0 | 49.9917985E+00 | DC220MV | DC220MV | 0.020005 | 4.5103e-08
| 10 | -0.35 | 0.2 | 0 | 10 | -0.35 | 20 | 10 | 0 | 49.9843775E+00 | DC220MV | DC220MV | 0.020005 | 5.1348e-08
| 10 | -0.35 | 0.2 | -0.05 | 10 | -0.3 | 20 | 10 | 0 | 50.0249511E+00 | DC220MV | DC220MV | 0.030005 | 4.6575e-08
| 10 | -0.3 | 0.2 | 0 | 10 | -0.3 | 20 | 10 | 0 | 49.9716387E+00 | DC220MV | DC220MV | 0.030004 | 3.8575e-08
| 10 | -0.3 | 0.2 | -0.05 | 10 | -0.25 | 20 | 10 | 0 | 49.9675352E+00 | DC220MV | DC220MV | 0.040004 | 4.146e-08
| 10 | -0.25 | 0.2 | 0 | 10 | -0.25 | 20 | 10 | 0 | 50.0280393E+00 | DC220MV | DC220MV | 0.040004 | 5.3093e-08
| 10 | -0.25 | 0.2 | -0.05 | 10 | -0.2 | 20 | 10 | 0 | 50.0554009E+00 | DC220MV | DC220MV | 0.050004 | 4.4105e-08
| 10 | -0.2 | 0.2 | 0 | 10 | -0.2 | 20 | 10 | 0 | 49.9716906E+00 | DC220MV | DC220MV | 0.050003 | 4.989e-08
| 10 | -0.2 | 0.2 | -0.05 | 10 | -0.15 | 20 | 10 | 0 | 49.9874135E+00 | DC220MV | DC220MV | 0.060003 | 4.2866e-08
| 10 | -0.15 | 0.2 | 0 | 10 | -0.15 | 20 | 10 | 0 | 49.9689156E+00 | DC220MV | DC220MV | 0.060003 | 4.4228e-08
| 10 | -0.15 | 0.2 | -0.05 | 10 | -0.1 | 20 | 10 | 0 | 49.9946185E+00 | DC220MV | DC220MV | 0.070003 | 7.9555e-08
| 10 | -0.1 | 0.2 | 0 | 10 | -0.1 | 20 | 10 | 0 | 49.9880877E+00 | DC220MV | DC220MV | 0.070002 | 4.1837e-08
| 10 | -0.1 | 0.2 | -0.05 | 10 | -0.05 | 20 | 10 | 0 | 50.0087987E+00 | DC220MV | DC220MV | 0.080002 | 7.1368e-08
| 10 | -0.05 | 0.2 | 0 | 10 | -0.05 | 20 | 10 | 0 | 49.9794990E+00 | DC220MV | DC220MV | 0.080002 | 5.1504e-08
| 10 | -0.05 | 0.2 | -0.05 | 10 | 0 | 20 | 10 | 0 | 49.9480446E+00 | DC220MV | DC220MV | 0.090002 | 4.66e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0089497E+00 | DC220MV | DC220MV | 0.090002 | 3.2379e-08
| 10 | -0.05 | 0.2 | -0.05 | 10 | 0 | 20 | 10 | 0 | 49.9811958E+00 | DC220MV | DC220MV | 0.1 | 5.4103e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9946138E+00 | DC220MV | DC220MV | 0.090002 | 5.163e-08
| 10 | -0.05 | 0.2 | 0 | 10 | -0.05 | 20 | 10 | 0 | 50.0182567E+00 | DC220MV | DC220MV | 0.090002 | 5.37e-07
| 10 | 0.1 | 0.2 | 0 | 10 | 0.1 | 20 | 10 | 0 | 50.0069538E+00 | DC220MV | DC220MV | 0.080002 | 4.9294e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9680582E+00 | DC220MV | DC220MV | 0.080003 | 4.2433e-08
| 10 | 0.1 | 0.2 | 0.1 | 10 | 0 | 20 | 10 | 0 | 50.0065008E+00 | DC220MV | DC220MV | 0.070003 | 5.5564e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0455414E+00 | DC220MV | DC220MV | 0.070003 | 4.4479e-08
| 10 | 0.1 | 0.2 | 0.1 | 10 | 0 | 20 | 10 | 0 | 50.0172890E+00 | DC220MV | DC220MV | 0.060003 | 3.8506e-08
| 10 | 0.1 | 0.2 | 0 | 10 | 0.1 | 20 | 10 | 0 | 50.0313923E+00 | DC220MV | DC220MV | 0.060004 | 5.6544e-08
| 10 | 0.2 | 0.2 | 0.1 | 10 | 0.1 | 20 | 10 | 0 | 50.0098180E+00 | DC220MV | DC220MV | 0.050004 | 4.5149e-08
| 10 | 0.2 | 0.2 | 0 | 10 | 0.2 | 20 | 10 | 0 | 49.9871212E+00 | DC220MV | DC220MV | 0.050004 | 5.6065e-08
| 10 | 0.3 | 0.2 | 0.1 | 10 | 0.2 | 20 | 10 | 0 | 49.9888091E+00 | DC220MV | DC220MV | 0.040004 | 6.4814e-08
| 10 | 0.3 | 0.2 | 0 | 10 | 0.3 | 20 | 10 | 0 | 50.0417802E+00 | DC220MV | DC220MV | 0.040005 | 4.3282e-08
| 10 | 0.4 | 0.2 | 0.1 | 10 | 0.3 | 20 | 10 | 0 | 49.9707247E+00 | DC220MV | DC220MV | 0.030005 | 3.3781e-08
| 10 | 0.4 | 0.2 | 0 | 10 | 0.4 | 20 | 10 | 0 | 49.9903839E+00 | DC220MV | DC220MV | 0.030005 | 4.149e-08
| 10 | 0.5 | 0.2 | 0.1 | 10 | 0.4 | 20 | 10 | 0 | 50.0350529E+00 | DC220MV | DC220MV | 0.020005 | 5.0905e-08
| 10 | 0.5 | 0.2 | 0 | 10 | 0.5 | 20 | 10 | 0 | 49.9981086E+00 | DC220MV | DC220MV | 0.020005 | 3.9469e-08
| 10 | 0.6 | 0.2 | 0.1 | 10 | 0.5 | 20 | 10 | 0 | 49.9734151E+00 | DC220MV | DC220MV | 0.010005 | 3.5281e-08
| 10 | 0.6 | 0.2 | 0 | 10 | 0.6 | 20 | 10 | 0 | 50.0092753E+00 | DC220MV | DC220MV | 0.010006 | 3.0983e-08
| 10 | 0.7 | 0.2 | 0.1 | 10 | 0.6 | 20 | 10 | 0 | 49.9634418E+00 | DC220MV | DC220MV | 5.4446e-06 | 5.5812e-08
| 10 | 0.7 | 0.2 | 0 | 10 | 0.7 | 20 | 10 | 0 | 50.0123144E+00 | DC220MV | DC220MV | 5.4484e-06 | 8.7916e-08
| 10 | 0.8 | 0.2 | 0.1 | 10 | 0.7 | 20 | 10 | 0 | 49.9921569E+00 | DC220MV | DC220MV | 4.7549e-06 | 6.0539e-08
| 10 | 0.8 | 0.2 | 0 | 10 | 0.8 | 20 | 10 | 0 | 49.9707859E+00 | DC220MV | DC220MV | 4.8223e-06 | 5.8733e-08
| 10 | 0.9 | 0.2 | 0.1 | 10 | 0.8 | 20 | 10 | 0 | 50.0001604E+00 | DC220MV | DC220MV | 0.010004 | 4.5554e-08
| 10 | 0.9 | 0.2 | 0 | 10 | 0.9 | 20 | 10 | 0 | 49.9567874E+00 | DC220MV | DC220MV | -0.0099949 | 2.9958e-08
| 10 | 1 | 0.2 | 0.1 | 10 | 0.9 | 20 | 10 | 0 | 50.0067839E+00 | DC220MV | DC220MV | 4.6438e-06 | 6.5219e-08
| 10 | 1 | 0.2 | 0 | 10 | 1 | 20 | 10 | 0 | 49.9751774E+00 | DC220MV | DC220MV | 5.0585e-06 | 2.973e-08
| 10 | 1 | 0.2 | 0.1 | 10 | 0.9 | 20 | 10 | 0 | 49.9859945E+00 | DC220MV | DC220MV | 4.5116e-06 | 4.2838e-08
| 10 | 0.9 | 0.2 | 0 | 10 | 0.9 | 20 | 10 | 0 | 50.0301172E+00 | DC220MV | DC220MV | 4.9348e-06 | 3.9575e-08
| 10 | 0.9 | 0.2 | 0.1 | 10 | 0.8 | 20 | 10 | 0 | 50.0593676E+00 | DC220MV | DC220MV | -0.0099951 | 3.8644e-08
| 10 | 0.8 | 0.2 | 0 | 10 | 0.8 | 20 | 10 | 0 | 50.0273924E+00 | DC220MV | DC220MV | -0.0099946 | 6.0552e-08
| 10 | 0.8 | 0.2 | 0.1 | 10 | 0.7 | 20 | 10 | 0 | 50.0151697E+00 | DC220MV | DC220MV | -0.019995 | 7.5968e-08
| 10 | 0.7 | 0.2 | 0 | 10 | 0.7 | 20 | 10 | 0 | 49.9966182E+00 | DC220MV | DC220MV | -0.019994 | 3.544e-08
| 10 | 0.7 | 0.2 | 0.1 | 10 | 0.6 | 20 | 10 | 0 | 49.9998443E+00 | DC220MV | DC220MV | -0.029994 | 4.0054e-08
| 10 | 0.6 | 0.2 | 0 | 10 | 0.6 | 20 | 10 | 0 | 50.0308634E+00 | DC220MV | DC220MV | -0.029994 | 3.9743e-08
| 10 | 0.6 | 0.2 | 0.1 | 10 | 0.5 | 20 | 10 | 0 | 49.9736742E+00 | DC220MV | DC220MV | -0.039994 | 5.294e-08
| 10 | 0.5 | 0.2 | 0 | 10 | 0.5 | 20 | 10 | 0 | 50.0024577E+00 | DC220MV | DC220MV | -0.039993 | 3.5819e-08
| 10 | 0.5 | 0.2 | 0.1 | 10 | 0.4 | 20 | 10 | 0 | 50.0028256E+00 | DC220MV | DC220MV | -0.049993 | 4.8468e-08
| 10 | 0.4 | 0.2 | 0 | 10 | 0.4 | 20 | 10 | 0 | 50.0306981E+00 | DC220MV | DC220MV | -0.049993 | 3.4955e-08
| 10 | 0.4 | 0.2 | 0.1 | 10 | 0.3 | 20 | 10 | 0 | 49.9751397E+00 | DC220MV | DC220MV | -0.059993 | 4.5685e-08
| 10 | 0.3 | 0.2 | 0 | 10 | 0.3 | 20 | 10 | 0 | 50.0084212E+00 | DC220MV | DC220MV | -0.059993 | 3.6657e-08
| 10 | 0.3 | 0.2 | 0.1 | 10 | 0.2 | 20 | 10 | 0 | 49.9834489E+00 | DC220MV | DC220MV | -0.069993 | 5.076e-08
| 10 | 0.2 | 0.2 | 0 | 10 | 0.2 | 20 | 10 | 0 | 49.9227751E+00 | DC220MV | DC220MV | -0.069992 | 4.5716e-08
| 10 | 0.2 | 0.2 | 0.1 | 10 | 0.1 | 20 | 10 | 0 | 49.9635784E+00 | DC220MV | DC220MV | -0.079993 | 5.4041e-08
| 10 | 0.1 | 0.2 | 0 | 10 | 0.1 | 20 | 10 | 0 | 49.9812806E+00 | DC220MV | DC220MV | -0.079992 | 4.7165e-08
| 10 | 0.1 | 0.2 | 0.1 | 10 | 0 | 20 | 10 | 0 | 49.9246373E+00 | DC220MV | DC220MV | -0.089992 | 4.5394e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0779657E+00 | DC220MV | DC220MV | -0.089992 | 2.9758e-08
| 10 | 0.1 | 0.2 | 0.1 | 10 | 0 | 20 | 10 | 0 | 49.9783961E+00 | DC220MV | DC220MV | -0.099992 | 4.8841e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9276849E+00 | DC220MV | DC220MV | -0.089992 | 3.9385e-08
| 10 | 0.1 | 0.2 | 0 | 10 | 0.1 | 20 | 10 | 0 | 49.9834819E+00 | DC220MV | DC220MV | -0.089992 | 7.4541e-08
| 10 | -0.1 | 0.2 | 0 | 10 | -0.1 | 20 | 10 | 0 | 50.0069538E+00 | DC220MV | DC220MV | 0.080002 | 4.9294e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9680582E+00 | DC220MV | DC220MV | 0.080003 | 4.2433e-08
| 10 | -0.1 | 0.2 | -0.1 | 10 | 0 | 20 | 10 | 0 | 50.0065008E+00 | DC220MV | DC220MV | 0.070003 | 5.5564e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0455414E+00 | DC220MV | DC220MV | 0.070003 | 4.4479e-08
| 10 | -0.1 | 0.2 | -0.1 | 10 | 0 | 20 | 10 | 0 | 50.0172890E+00 | DC220MV | DC220MV | 0.060003 | 3.8506e-08
| 10 | -0.1 | 0.2 | 0 | 10 | -0.1 | 20 | 10 | 0 | 50.0313923E+00 | DC220MV | DC220MV | 0.060004 | 5.6544e-08
| 10 | -0.2 | 0.2 | -0.1 | 10 | -0.1 | 20 | 10 | 0 | 50.0098180E+00 | DC220MV | DC220MV | 0.050004 | 4.5149e-08
| 10 | -0.2 | 0.2 | 0 | 10 | -0.2 | 20 | 10 | 0 | 49.9871212E+00 | DC220MV | DC220MV | 0.050004 | 5.6065e-08
| 10 | -0.3 | 0.2 | -0.1 | 10 | -0.2 | 20 | 10 | 0 | 49.9888091E+00 | DC220MV | DC220MV | 0.040004 | 6.4814e-08
| 10 | -0.3 | 0.2 | 0 | 10 | -0.3 | 20 | 10 | 0 | 50.0417802E+00 | DC220MV | DC220MV | 0.040005 | 4.3282e-08
| 10 | -0.4 | 0.2 | -0.1 | 10 | -0.3 | 20 | 10 | 0 | 49.9707247E+00 | DC220MV | DC220MV | 0.030005 | 3.3781e-08
| 10 | -0.4 | 0.2 | 0 | 10 | -0.4 | 20 | 10 | 0 | 49.9903839E+00 | DC220MV | DC220MV | 0.030005 | 4.149e-08
| 10 | -0.5 | 0.2 | -0.1 | 10 | -0.4 | 20 | 10 | 0 | 50.0350529E+00 | DC220MV | DC220MV | 0.020005 | 5.0905e-08
| 10 | -0.5 | 0.2 | 0 | 10 | -0.5 | 20 | 10 | 0 | 49.9981086E+00 | DC220MV | DC220MV | 0.020005 | 3.9469e-08
| 10 | -0.6 | 0.2 | -0.1 | 10 | -0.5 | 20 | 10 | 0 | 49.9734151E+00 | DC220MV | DC220MV | 0.010005 | 3.5281e-08
| 10 | -0.6 | 0.2 | 0 | 10 | -0.6 | 20 | 10 | 0 | 50.0092753E+00 | DC220MV | DC220MV | 0.010006 | 3.0983e-08
| 10 | -0.7 | 0.2 | -0.1 | 10 | -0.6 | 20 | 10 | 0 | 49.9634418E+00 | DC220MV | DC220MV | 5.4446e-06 | 5.5812e-08
| 10 | -0.7 | 0.2 | 0 | 10 | -0.7 | 20 | 10 | 0 | 50.0123144E+00 | DC220MV | DC220MV | 5.4484e-06 | 8.7916e-08
| 10 | -0.8 | 0.2 | -0.1 | 10 | -0.7 | 20 | 10 | 0 | 49.9921569E+00 | DC220MV | DC220MV | 4.7549e-06 | 6.0539e-08
| 10 | -0.8 | 0.2 | 0 | 10 | -0.8 | 20 | 10 | 0 | 49.9707859E+00 | DC220MV | DC220MV | 4.8223e-06 | 5.8733e-08
| 10 | -0.9 | 0.2 | -0.1 | 10 | -0.8 | 20 | 10 | 0 | 50.0001604E+00 | DC220MV | DC220MV | 0.010004 | 4.5554e-08
| 10 | -0.9 | 0.2 | 0 | 10 | -0.9 | 20 | 10 | 0 | 49.9567874E+00 | DC220MV | DC220MV | -0.0099949 | 2.9958e-08
| 10 | -1 | 0.2 | -0.1 | 10 | -0.9 | 20 | 10 | 0 | 50.0067839E+00 | DC220MV | DC220MV | 4.6438e-06 | 6.5219e-08
| 10 | -1 | 0.2 | 0 | 10 | -1 | 20 | 10 | 0 | 49.9751774E+00 | DC220MV | DC220MV | 5.0585e-06 | 2.973e-08
| 10 | -1 | 0.2 | -0.1 | 10 | -0.9 | 20 | 10 | 0 | 49.9859945E+00 | DC220MV | DC220MV | 4.5116e-06 | 4.2838e-08
| 10 | -0.9 | 0.2 | 0 | 10 | -0.9 | 20 | 10 | 0 | 50.0301172E+00 | DC220MV | DC220MV | 4.9348e-06 | 3.9575e-08
| 10 | -0.9 | 0.2 | -0.1 | 10 | -0.8 | 20 | 10 | 0 | 50.0593676E+00 | DC220MV | DC220MV | -0.0099951 | 3.8644e-08
| 10 | -0.8 | 0.2 | 0 | 10 | -0.8 | 20 | 10 | 0 | 50.0273924E+00 | DC220MV | DC220MV | -0.0099946 | 6.0552e-08
| 10 | -0.8 | 0.2 | -0.1 | 10 | -0.7 | 20 | 10 | 0 | 50.0151697E+00 | DC220MV | DC220MV | -0.019995 | 7.5968e-08
| 10 | -0.7 | 0.2 | 0 | 10 | -0.7 | 20 | 10 | 0 | 49.9966182E+00 | DC220MV | DC220MV | -0.019994 | 3.544e-08
| 10 | -0.7 | 0.2 | -0.1 | 10 | -0.6 | 20 | 10 | 0 | 49.9998443E+00 | DC220MV | DC220MV | -0.029994 | 4.0054e-08
| 10 | -0.6 | 0.2 | 0 | 10 | -0.6 | 20 | 10 | 0 | 50.0308634E+00 | DC220MV | DC220MV | -0.029994 | 3.9743e-08
| 10 | -0.6 | 0.2 | -0.1 | 10 | -0.5 | 20 | 10 | 0 | 49.9736742E+00 | DC220MV | DC220MV | -0.039994 | 5.294e-08
| 10 | -0.5 | 0.2 | 0 | 10 | -0.5 | 20 | 10 | 0 | 50.0024577E+00 | DC220MV | DC220MV | -0.039993 | 3.5819e-08
| 10 | -0.5 | 0.2 | -0.1 | 10 | -0.4 | 20 | 10 | 0 | 50.0028256E+00 | DC220MV | DC220MV | -0.049993 | 4.8468e-08
| 10 | -0.4 | 0.2 | 0 | 10 | -0.4 | 20 | 10 | 0 | 50.0306981E+00 | DC220MV | DC220MV | -0.049993 | 3.4955e-08
| 10 | -0.4 | 0.2 | -0.1 | 10 | -0.3 | 20 | 10 | 0 | 49.9751397E+00 | DC220MV | DC220MV | -0.059993 | 4.5685e-08
| 10 | -0.3 | 0.2 | 0 | 10 | -0.3 | 20 | 10 | 0 | 50.0084212E+00 | DC220MV | DC220MV | -0.059993 | 3.6657e-08
| 10 | -0.3 | 0.2 | -0.1 | 10 | -0.2 | 20 | 10 | 0 | 49.9834489E+00 | DC220MV | DC220MV | -0.069993 | 5.076e-08
| 10 | -0.2 | 0.2 | 0 | 10 | -0.2 | 20 | 10 | 0 | 49.9227751E+00 | DC220MV | DC220MV | -0.069992 | 4.5716e-08
| 10 | -0.2 | 0.2 | -0.1 | 10 | -0.1 | 20 | 10 | 0 | 49.9635784E+00 | DC220MV | DC220MV | -0.079993 | 5.4041e-08
| 10 | -0.1 | 0.2 | 0 | 10 | -0.1 | 20 | 10 | 0 | 49.9812806E+00 | DC220MV | DC220MV | -0.079992 | 4.7165e-08
| 10 | -0.1 | 0.2 | -0.1 | 10 | 0 | 20 | 10 | 0 | 49.9246373E+00 | DC220MV | DC220MV | -0.089992 | 4.5394e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0779657E+00 | DC220MV | DC220MV | -0.089992 | 2.9758e-08
| 10 | -0.1 | 0.2 | -0.1 | 10 | 0 | 20 | 10 | 0 | 49.9783961E+00 | DC220MV | DC220MV | -0.099992 | 4.8841e-08
| 10 | 0 | 0.2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9276849E+00 | DC220MV | DC220MV | -0.089992 | 3.9385e-08
| 10 | -0.1 | 0.2 | 0 | 10 | -0.1 | 20 | 10 | 0 | 49.9834819E+00 | DC220MV | DC220MV | -0.089992 | 7.4541e-08
| 10 | 1 | 2 | 0 | 10 | 1 | 20 | 10 | 0 | 49.9816624E+00 | DC220MV | DC2_2V | -0.079992 | 3.2109e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9695751E+00 | DC220MV | DC2_2V | -0.079993 | 5.3007e-08
| 10 | 1 | 2 | 1 | 10 | 0 | 20 | 10 | 0 | 50.0398430E+00 | DC220MV | DC2_2V | -0.069993 | 4.2062e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0250739E+00 | DC220MV | DC2_2V | -0.069993 | 3.84e-08
| 10 | 1 | 2 | 1 | 10 | 0 | 20 | 10 | 0 | 50.0171096E+00 | DC220MV | DC2_2V | -0.059993 | 4.2363e-08
| 10 | 1 | 2 | 0 | 10 | 1 | 20 | 10 | 0 | 49.9844105E+00 | DC220MV | DC2_2V | -0.059993 | 4.3653e-08
| 10 | 2 | 2 | 1 | 10 | 1 | 20 | 10 | 0 | 50.0200316E+00 | DC220MV | DC2_2V | -0.049993 | 6.0246e-08
| 10 | 2 | 2 | 0 | 10 | 2 | 20 | 10 | 0 | 50.0360873E+00 | DC220MV | DC2_2V | -0.049993 | 3.7076e-08
| 10 | 3 | 2 | 1 | 10 | 2 | 20 | 10 | 0 | 49.9919258E+00 | DC220MV | DC2_2V | -0.039993 | 4.2491e-08
| 10 | 3 | 2 | 0 | 10 | 3 | 20 | 10 | 0 | 50.0050854E+00 | DC220MV | DC2_2V | -0.039994 | 4.7936e-08
| 10 | 4 | 2 | 1 | 10 | 3 | 20 | 10 | 0 | 50.0059394E+00 | DC220MV | DC2_2V | -0.029994 | 6.0326e-08
| 10 | 4 | 2 | 0 | 10 | 4 | 20 | 10 | 0 | 50.0113469E+00 | DC220MV | DC2_2V | -0.029994 | 2.7836e-08
| 10 | 5 | 2 | 1 | 10 | 4 | 20 | 10 | 0 | 49.9915014E+00 | DC220MV | DC2_2V | -0.019994 | 4.6542e-08
| 10 | 5 | 2 | 0 | 10 | 5 | 20 | 10 | 0 | 50.0096103E+00 | DC220MV | DC2_2V | -0.019995 | 4.1285e-08
| 10 | 5 | 2 | 1 | 10 | 4 | 20 | 10 | 0 | 50.0684095E+00 | DC220MV | DC2_2V | -0.0099947 | 4.8006e-08
| 10 | 4 | 2 | 0 | 10 | 4 | 20 | 10 | 0 | 50.0477106E+00 | DC220MV | DC2_2V | -0.0099951 | 4.321e-08
| 10 | 4 | 2 | 1 | 10 | 3 | 20 | 10 | 0 | 49.9717565E+00 | DC220MV | DC2_2V | 4.9641e-06 | 3.9357e-08
| 10 | 3 | 2 | 0 | 10 | 3 | 20 | 10 | 0 | 50.0544838E+00 | DC220MV | DC2_2V | 4.5697e-06 | 7.0682e-08
| 10 | 3 | 2 | 1 | 10 | 2 | 20 | 10 | 0 | 50.0214290E+00 | DC220MV | DC2_2V | 5.163e-06 | 4.253e-08
| 10 | 2 | 2 | 0 | 10 | 2 | 20 | 10 | 0 | 49.9219569E+00 | DC220MV | DC2_2V | 4.6397e-06 | 3.1806e-08
| 10 | 2 | 2 | 1 | 10 | 1 | 20 | 10 | 0 | 50.0338814E+00 | DC220MV | DC2_2V | -0.009995 | 4.4132e-08
| 10 | 1 | 2 | 0 | 10 | 1 | 20 | 10 | 0 | 50.1087360E+00 | DC220MV | DC2_2V | 0.1 | 3.0148e-08
| 10 | 1 | 2 | 1 | 10 | 0 | 20 | 10 | 0 | 50.0456029E+00 | DC220MV | DC2_2V | 4.6034e-06 | 3.6841e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9975662E+00 | DC220MV | DC2_2V | 3.8322e-07 | 3.3318e-08
| 10 | 1 | 2 | 1 | 10 | 0 | 20 | 10 | 0 | 50.0492891E+00 | DC220MV | DC2_2V | 4.4662e-06 | 4.7154e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9829162E+00 | DC220MV | DC2_2V | 0.010005 | 5.2663e-08
| 10 | 1 | 2 | 0 | 10 | 1 | 20 | 10 | 0 | 49.9883046E+00 | DC220MV | DC2_2V | 5.6183e-06 | 3.6452e-08
| 10 | -1 | 2 | 0 | 10 | -1 | 20 | 10 | 0 | 49.9816624E+00 | DC220MV | DC2_2V | -0.079992 | 3.2109e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9695751E+00 | DC220MV | DC2_2V | -0.079993 | 5.3007e-08
| 10 | -1 | 2 | -1 | 10 | 0 | 20 | 10 | 0 | 50.0398430E+00 | DC220MV | DC2_2V | -0.069993 | 4.2062e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0250739E+00 | DC220MV | DC2_2V | -0.069993 | 3.84e-08
| 10 | -1 | 2 | -1 | 10 | 0 | 20 | 10 | 0 | 50.0171096E+00 | DC220MV | DC2_2V | -0.059993 | 4.2363e-08
| 10 | -2 | 2 | 0 | 10 | -2 | 20 | 10 | 0 | 49.9844105E+00 | DC220MV | DC2_2V | -0.059993 | 4.3653e-08
| 10 | -2 | 2 | -1 | 10 | -1 | 20 | 10 | 0 | 50.0200316E+00 | DC220MV | DC2_2V | -0.049993 | 6.0246e-08
| 10 | -2 | 2 | 0 | 10 | -2 | 20 | 10 | 0 | 50.0360873E+00 | DC220MV | DC2_2V | -0.049993 | 3.7076e-08
| 10 | -3 | 2 | -1 | 10 | -2 | 20 | 10 | 0 | 49.9919258E+00 | DC220MV | DC2_2V | -0.039993 | 4.2491e-08
| 10 | -3 | 2 | 0 | 10 | -3 | 20 | 10 | 0 | 50.0050854E+00 | DC220MV | DC2_2V | -0.039994 | 4.7936e-08
| 10 | -4 | 2 | -1 | 10 | -3 | 20 | 10 | 0 | 50.0059394E+00 | DC220MV | DC2_2V | -0.029994 | 6.0326e-08
| 10 | -4 | 2 | 0 | 10 | -4 | 20 | 10 | 0 | 50.0113469E+00 | DC220MV | DC2_2V | -0.029994 | 2.7836e-08
| 10 | -5 | 2 | -1 | 10 | -4 | 20 | 10 | 0 | 49.9915014E+00 | DC220MV | DC2_2V | -0.019994 | 4.6542e-08
| 10 | -5 | 2 | 0 | 10 | -5 | 20 | 10 | 0 | 50.0096103E+00 | DC220MV | DC2_2V | -0.019995 | 4.1285e-08
| 10 | -5 | 2 | -1 | 10 | -4 | 20 | 10 | 0 | 50.0684095E+00 | DC220MV | DC2_2V | -0.0099947 | 4.8006e-08
| 10 | -4 | 2 | 0 | 10 | -4 | 20 | 10 | 0 | 50.0477106E+00 | DC220MV | DC2_2V | -0.0099951 | 4.321e-08
| 10 | -4 | 2 | -1 | 10 | -3 | 20 | 10 | 0 | 49.9717565E+00 | DC220MV | DC2_2V | 4.9641e-06 | 3.9357e-08
| 10 | -3 | 2 | 0 | 10 | -3 | 20 | 10 | 0 | 50.0544838E+00 | DC220MV | DC2_2V | 4.5697e-06 | 7.0682e-08
| 10 | -3 | 2 | -1 | 10 | -2 | 20 | 10 | 0 | 50.0214290E+00 | DC220MV | DC2_2V | 5.163e-06 | 4.253e-08
| 10 | -2 | 2 | 0 | 10 | -2 | 20 | 10 | 0 | 49.9219569E+00 | DC220MV | DC2_2V | 4.6397e-06 | 3.1806e-08
| 10 | -2 | 2 | -1 | 10 | -1 | 20 | 10 | 0 | 50.0338814E+00 | DC220MV | DC2_2V | -0.009995 | 4.4132e-08
| 10 | -1 | 2 | 0 | 10 | -1 | 20 | 10 | 0 | 50.1087360E+00 | DC220MV | DC2_2V | 0.1 | 3.0148e-08
| 10 | -1 | 2 | -1 | 10 | 0 | 20 | 10 | 0 | 50.0456029E+00 | DC220MV | DC2_2V | 4.6034e-06 | 3.6841e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9975662E+00 | DC220MV | DC2_2V | 3.8322e-07 | 3.3318e-08
| 10 | -1 | 2 | -1 | 10 | 0 | 20 | 10 | 0 | 50.0492891E+00 | DC220MV | DC2_2V | 4.4662e-06 | 4.7154e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9829162E+00 | DC220MV | DC2_2V | 0.010005 | 5.2663e-08
| 10 | -1 | 2 | 0 | 10 | -1 | 20 | 10 | 0 | 49.9883046E+00 | DC220MV | DC2_2V | 5.6183e-06 | 3.6452e-08
| 10 | 1 | 2 | 0 | 10 | 1 | 20 | 10 | 0 | 49.9992547E+00 | DC220MV | DC2_2V | 5.1692e-06 | 4.3858e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9908602E+00 | DC220MV | DC2_2V | 5.5943e-06 | 3.7356e-08
| 10 | 1 | 2 | 1 | 10 | 0 | 20 | 10 | 0 | 49.9872721E+00 | DC220MV | DC2_2V | 5.2605e-06 | 3.5703e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0022878E+00 | DC220MV | DC2_2V | 0.010005 | 3.5711e-08
| 10 | 1 | 2 | 1 | 10 | 0 | 20 | 10 | 0 | 49.9996226E+00 | DC220MV | DC2_2V | 0.010005 | 3.7356e-08
| 10 | 1 | 2 | 0 | 10 | 1 | 20 | 10 | 0 | 50.0800194E+00 | DC220MV | DC2_2V | 0.020005 | 4.5103e-08
| 10 | 2 | 2 | 1 | 10 | 1 | 20 | 10 | 0 | 49.9637951E+00 | DC220MV | DC2_2V | 0.020005 | 5.1348e-08
| 10 | 2 | 2 | 0 | 10 | 2 | 20 | 10 | 0 | 49.9802531E+00 | DC220MV | DC2_2V | 0.030005 | 4.6575e-08
| 10 | 3 | 2 | 1 | 10 | 2 | 20 | 10 | 0 | 49.9646712E+00 | DC220MV | DC2_2V | 0.030004 | 3.8575e-08
| 10 | 3 | 2 | 0 | 10 | 3 | 20 | 10 | 0 | 49.9537362E+00 | DC220MV | DC2_2V | 0.040004 | 4.146e-08
| 10 | 4 | 2 | 1 | 10 | 3 | 20 | 10 | 0 | 49.9582519E+00 | DC220MV | DC2_2V | 0.040004 | 5.3093e-08
| 10 | 4 | 2 | 0 | 10 | 4 | 20 | 10 | 0 | 50.0248756E+00 | DC220MV | DC2_2V | 0.050004 | 4.4105e-08
| 10 | 5 | 2 | 1 | 10 | 4 | 20 | 10 | 0 | 50.0470206E+00 | DC220MV | DC2_2V | 0.050003 | 4.989e-08
| 10 | 5 | 2 | 0 | 10 | 5 | 20 | 10 | 0 | 50.0172276E+00 | DC220MV | DC2_2V | 0.060003 | 4.2866e-08
| 10 | 6 | 2 | 1 | 10 | 5 | 20 | 10 | 0 | 49.9859143E+00 | DC220MV | DC220MV | 0.060003 | 4.4228e-08
| 10 | 6 | 2 | 0 | 10 | 6 | 20 | 10 | 0 | 50.0120690E+00 | DC220MV | DC220MV | 0.070003 | 7.9555e-08
| 10 | 7 | 2 | 1 | 10 | 6 | 20 | 10 | 0 | 50.0356528E+00 | DC220MV | DC220MV | 0.070002 | 4.1837e-08
| 10 | 7 | 2 | 0 | 10 | 7 | 20 | 10 | 0 | 49.9932603E+00 | DC220MV | DC220MV | 0.080002 | 7.1368e-08
| 10 | 8 | 2 | 1 | 10 | 7 | 20 | 10 | 0 | 50.0500075E+00 | DC220MV | DC220MV | 0.080002 | 5.1504e-08
| 10 | 8 | 2 | 0 | 10 | 8 | 20 | 10 | 0 | 50.0102144E+00 | DC220MV | DC220MV | 0.090002 | 4.66e-08
| 10 | 9 | 2 | 1 | 10 | 8 | 20 | 10 | 0 | 49.9887808E+00 | DC220MV | DC2_2V | 0.090002 | 3.2379e-08
| 10 | 9 | 2 | 0 | 10 | 9 | 20 | 10 | 0 | 50.0200505E+00 | DC220MV | DC2_2V | 0.1 | 5.4103e-08
| 10 | 10 | 2 | 1 | 10 | 9 | 20 | 10 | 0 | 49.9935952E+00 | DC220MV | DC2_2V | 0.090002 | 5.163e-08
| 10 | 10 | 2 | 0 | 10 | 10 | 20 | 10 | 0 | 50.0334138E+00 | DC220MV | DC2_2V | 0.090002 | 5.37e-07
| 10 | 10 | 2 | 1 | 10 | 9 | 20 | 10 | 0 | 50.0612684E+00 | DC220MV | DC2_2V | 0.080002 | 4.9294e-08
| 10 | 9 | 2 | 0 | 10 | 9 | 20 | 10 | 0 | 50.0152783E+00 | DC220MV | DC2_2V | 0.080003 | 4.2433e-08
| 10 | 9 | 2 | 1 | 10 | 8 | 20 | 10 | 0 | 49.9900256E+00 | DC220MV | DC2_2V | 0.070003 | 5.5564e-08
| 10 | 8 | 2 | 0 | 10 | 8 | 20 | 10 | 0 | 49.9713042E+00 | DC220MV | DC2_2V | 0.070003 | 4.4479e-08
| 10 | 8 | 2 | 1 | 10 | 7 | 20 | 10 | 0 | 49.9393284E+00 | DC220MV | DC2_2V | 0.060003 | 3.8506e-08
| 10 | 7 | 2 | 0 | 10 | 7 | 20 | 10 | 0 | 49.9646194E+00 | DC220MV | DC2_2V | 0.060004 | 5.6544e-08
| 10 | 7 | 2 | 1 | 10 | 6 | 20 | 10 | 0 | 49.9501535E+00 | DC220MV | DC2_2V | 0.050004 | 4.5149e-08
| 10 | 6 | 2 | 0 | 10 | 6 | 20 | 10 | 0 | 50.0500501E+00 | DC220MV | DC2_2V | 0.050004 | 5.6065e-08
| 10 | 6 | 2 | 1 | 10 | 5 | 20 | 10 | 0 | 49.9714738E+00 | DC220MV | DC2_2V | 0.040004 | 6.4814e-08
| 10 | 5 | 2 | 0 | 10 | 5 | 20 | 10 | 0 | 50.0005236E+00 | DC220MV | DC2_2V | 0.040005 | 4.3282e-08
| 10 | 5 | 2 | 1 | 10 | 4 | 20 | 10 | 0 | 50.0048354E+00 | DC220MV | DC2_2V | 0.030005 | 3.3781e-08
| 10 | 4 | 2 | 0 | 10 | 4 | 20 | 10 | 0 | 50.0362243E+00 | DC220MV | DC2_2V | 0.030005 | 4.149e-08
| 10 | 4 | 2 | 1 | 10 | 3 | 20 | 10 | 0 | 49.9836091E+00 | DC220MV | DC2_2V | 0.020005 | 5.0905e-08
| 10 | 3 | 2 | 0 | 10 | 3 | 20 | 10 | 0 | 50.0367298E+00 | DC220MV | DC2_2V | 0.020005 | 3.9469e-08
| 10 | 3 | 2 | 1 | 10 | 2 | 20 | 10 | 0 | 49.9842738E+00 | DC220MV | DC2_2V | 0.010005 | 3.5281e-08
| 10 | 2 | 2 | 0 | 10 | 2 | 20 | 10 | 0 | 49.9949250E+00 | DC220MV | DC2_2V | 0.010006 | 3.0983e-08
| 10 | 2 | 2 | 1 | 10 | 1 | 20 | 10 | 0 | 49.9986793E+00 | DC220MV | DC2_2V | 5.4446e-06 | 5.5812e-08
| 10 | 1 | 2 | 0 | 10 | 1 | 20 | 10 | 0 | 49.9566037E+00 | DC220MV | DC2_2V | 5.4484e-06 | 8.7916e-08
| 10 | 1 | 2 | 1 | 10 | 0 | 20 | 10 | 0 | 50.1014460E+00 | DC220MV | DC2_2V | 4.7549e-06 | 6.0539e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9959437E+00 | DC220MV | DC2_2V | 4.8223e-06 | 5.8733e-08
| 10 | 1 | 2 | 1 | 10 | 0 | 20 | 10 | 0 | 50.0459715E+00 | DC220MV | DC2_2V | 0.010004 | 4.5554e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 49.9684774E+00 | DC220MV | DC2_2V | -0.0099949 | 2.9958e-08
| 10 | 1 | 2 | 0 | 10 | 1 | 20 | 10 | 0 | 49.9719874E+00 | DC220MV | DC2_2V | 4.6438e-06 | 6.5219e-08
| 10 | -1 | 2 | 0 | 10 | -1 | 20 | 10 | 0 | 49.9837600E+00 | DC220MV | DC2_2V | 5.0585e-06 | 2.973e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0258577E+00 | DC220MV | DC2_2V | 4.5116e-06 | 4.2838e-08
| 10 | -1 | 2 | -1 | 10 | 0 | 20 | 10 | 0 | 49.9915910E+00 | DC220MV | DC2_2V | 4.9348e-06 | 3.9575e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0515247E+00 | DC220MV | DC2_2V | -0.0099951 | 3.8644e-08
| 10 | -1 | 2 | -1 | 10 | 0 | 20 | 10 | 0 | 50.0354119E+00 | DC220MV | DC2_2V | -0.0099946 | 6.0552e-08
| 10 | -1 | 2 | 0 | 10 | -1 | 20 | 10 | 0 | 50.0273499E+00 | DC220MV | DC2_2V | -0.019995 | 7.5968e-08
| 10 | -2 | 2 | -1 | 10 | -1 | 20 | 10 | 0 | 49.9984812E+00 | DC220MV | DC2_2V | -0.019994 | 3.544e-08
| 10 | -2 | 2 | 0 | 10 | -2 | 20 | 10 | 0 | 49.9898417E+00 | DC220MV | DC2_2V | -0.029994 | 4.0054e-08
| 10 | -3 | 2 | -1 | 10 | -2 | 20 | 10 | 0 | 49.9354655E+00 | DC220MV | DC2_2V | -0.029994 | 3.9743e-08
| 10 | -3 | 2 | 0 | 10 | -3 | 20 | 10 | 0 | 49.9777363E+00 | DC220MV | DC2_2V | -0.039994 | 5.294e-08
| 10 | -4 | 2 | -1 | 10 | -3 | 20 | 10 | 0 | 50.0357094E+00 | DC220MV | DC2_2V | -0.039993 | 3.5819e-08
| 10 | -4 | 2 | 0 | 10 | -4 | 20 | 10 | 0 | 49.9519283E+00 | DC220MV | DC2_2V | -0.049993 | 4.8468e-08
| 10 | -5 | 2 | -1 | 10 | -4 | 20 | 10 | 0 | 49.9450228E+00 | DC220MV | DC2_2V | -0.049993 | 3.4955e-08
| 10 | -5 | 2 | 0 | 10 | -5 | 20 | 10 | 0 | 50.0212165E+00 | DC220MV | DC2_2V | -0.059993 | 4.5685e-08
| 10 | -6 | 2 | -1 | 10 | -5 | 20 | 10 | 0 | 49.9804463E+00 | DC220MV | DC220MV | -0.059993 | 3.6657e-08
| 10 | -6 | 2 | 0 | 10 | -6 | 20 | 10 | 0 | 49.9891109E+00 | DC220MV | DC220MV | -0.069993 | 5.076e-08
| 10 | -7 | 2 | -1 | 10 | -6 | 20 | 10 | 0 | 49.9998066E+00 | DC220MV | DC220MV | -0.069992 | 4.5716e-08
| 10 | -7 | 2 | 0 | 10 | -7 | 20 | 10 | 0 | 50.0422385E+00 | DC2_2V | DC2_2V | -0.079993 | 5.4041e-08
| 10 | -8 | 2 | -1 | 10 | -7 | 20 | 10 | 0 | 49.9700039E+00 | DC2_2V | DC2_2V | -0.079992 | 4.7165e-08
| 10 | -8 | 2 | 0 | 10 | -8 | 20 | 10 | 0 | 50.0165526E+00 | DC2_2V | DC2_2V | -0.089992 | 4.5394e-08
| 10 | -9 | 2 | -1 | 10 | -8 | 20 | 10 | 0 | 49.9795602E+00 | DC2_2V | DC11V | -0.089992 | 2.9758e-08
| 10 | -9 | 2 | 0 | 10 | -9 | 20 | 10 | 0 | 49.8893445E+00 | DC2_2V | DC11V | -0.099992 | 4.8841e-08
| 10 | -10 | 2 | -1 | 10 | -9 | 20 | 10 | 0 | 49.9553088E+00 | DC2_2V | DC11V | -0.089992 | 3.9385e-08
| 10 | -10 | 2 | 0 | 10 | -10 | 20 | 10 | 0 | 50.0452768E+00 | DC2_2V | DC11V | -0.089992 | 7.4541e-08
| 10 | -10 | 2 | -1 | 10 | -9 | 20 | 10 | 0 | 49.9977501E+00 | DC2_2V | DC11V | -0.079992 | 3.2109e-08
| 10 | -9 | 2 | 0 | 10 | -9 | 20 | 10 | 0 | 50.0547486E+00 | DC2_2V | DC11V | -0.079993 | 5.3007e-08
| 10 | -9 | 2 | -1 | 10 | -8 | 20 | 10 | 0 | 50.0525741E+00 | DC2_2V | DC11V | -0.069993 | 4.2062e-08
| 10 | -8 | 2 | 0 | 10 | -8 | 20 | 10 | 0 | 49.9871212E+00 | DC2_2V | DC11V | -0.069993 | 3.84e-08
| 10 | -8 | 2 | -1 | 10 | -7 | 20 | 10 | 0 | 49.9714785E+00 | DC2_2V | DC11V | -0.059993 | 4.2363e-08
| 10 | -7 | 2 | 0 | 10 | -7 | 20 | 10 | 0 | 50.0065527E+00 | DC2_2V | DC11V | -0.059993 | 4.3653e-08
| 10 | -7 | 2 | -1 | 10 | -6 | 20 | 10 | 0 | 49.9839957E+00 | DC2_2V | DC11V | -0.049993 | 6.0246e-08
| 10 | -6 | 2 | 0 | 10 | -6 | 20 | 10 | 0 | 49.9931142E+00 | DC2_2V | DC11V | -0.049993 | 3.7076e-08
| 10 | -6 | 2 | -1 | 10 | -5 | 20 | 10 | 0 | 49.9825250E+00 | DC2_2V | DC11V | -0.039993 | 4.2491e-08
| 10 | -5 | 2 | 0 | 10 | -5 | 20 | 10 | 0 | 50.0452579E+00 | DC2_2V | DC11V | -0.039994 | 4.7936e-08
| 10 | -5 | 2 | -1 | 10 | -4 | 20 | 10 | 0 | 49.8609679E+00 | DC2_2V | DC11V | -0.029994 | 6.0326e-08
| 10 | -4 | 2 | 0 | 10 | -4 | 20 | 10 | 0 | 49.9730900E+00 | DC2_2V | DC11V | -0.029994 | 2.7836e-08
| 10 | -4 | 2 | -1 | 10 | -3 | 20 | 10 | 0 | 50.0250172E+00 | DC2_2V | DC11V | -0.019994 | 4.6542e-08
| 10 | -3 | 2 | 0 | 10 | -3 | 20 | 10 | 0 | 50.0696488E+00 | DC2_2V | DC11V | -0.019995 | 4.1285e-08
| 10 | -3 | 2 | -1 | 10 | -2 | 20 | 10 | 0 | 50.0135508E+00 | DC2_2V | DC11V | -0.0099947 | 4.8006e-08
| 10 | -2 | 2 | 0 | 10 | -2 | 20 | 10 | 0 | 49.9570887E+00 | DC2_2V | DC11V | -0.0099951 | 4.321e-08
| 10 | -2 | 2 | -1 | 10 | -1 | 20 | 10 | 0 | 49.9908696E+00 | DC2_2V | DC11V | 4.9641e-06 | 3.9357e-08
| 10 | -1 | 2 | 0 | 10 | -1 | 20 | 10 | 0 | 49.9654861E+00 | DC2_2V | DC11V | 4.5697e-06 | 7.0682e-08
| 10 | -1 | 2 | -1 | 10 | 0 | 20 | 10 | 0 | 50.0630369E+00 | DC2_2V | DC11V | 5.163e-06 | 4.253e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0308728E+00 | DC2_2V | DC11V | 4.6397e-06 | 3.1806e-08
| 10 | -1 | 2 | -1 | 10 | 0 | 20 | 10 | 0 | 50.0105777E+00 | DC2_2V | DC11V | -0.009995 | 4.4132e-08
| 10 | 0 | 2 | 0 | 10 | 0 | 20 | 10 | 0 | 50.0179546E+00 | DC2_2V | DC11V | 0.1 | 3.0148e-08
| 10 | -1 | 2 | 0 | 10 | -1 | 20 | 10 | 0 | 49.9165076E+00 | DC2_2V | DC11V | 4.6034e-06 | 3.6841e-08
| 100 | 5 | 10 | 0 | 100 | 5 | 20 | 10 | 0 | 49.9685198E+00 | DC2_2V | DC11V | 3.8322e-07 | 3.3318e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 49.9781322E+00 | DC2_2V | DC11V | 4.4662e-06 | 4.7154e-08
| 100 | 5 | 10 | 5 | 100 | 0 | 20 | 10 | 0 | 49.9635549E+00 | DC2_2V | DC11V | 0.010005 | 5.2663e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 49.9867064E+00 | DC2_2V | DC11V | 5.6183e-06 | 3.6452e-08
| 100 | 5 | 10 | 5 | 100 | 0 | 20 | 10 | 0 | 49.9604323E+00 | DC2_2V | DC11V | 5.1692e-06 | 4.3858e-08
| 100 | 5 | 10 | 0 | 100 | 5 | 20 | 10 | 0 | 49.9561470E+00 | DC2_2V | DC11V | 5.5943e-06 | 3.7356e-08
| 100 | 10 | 10 | 5 | 100 | 5 | 20 | 10 | 0 | 50.0151980E+00 | DC2_2V | DC11V | 5.2605e-06 | 3.5703e-08
| 100 | 10 | 10 | 0 | 100 | 10 | 20 | 10 | 0 | 50.0420353E+00 | DC2_2V | DC11V | 0.010005 | 3.5711e-08
| 100 | 15 | 10 | 5 | 100 | 10 | 20 | 10 | 0 | 49.9588546E+00 | DC2_2V | DC11V | 0.010005 | 3.7356e-08
| 100 | 15 | 10 | 0 | 100 | 15 | 20 | 10 | 0 | 49.9365852E+00 | DC2_2V | DC11V | 0.020005 | 4.5103e-08
| 100 | 20 | 10 | 5 | 100 | 15 | 20 | 10 | 0 | 49.9921333E+00 | DC2_2V | DC11V | 0.020005 | 5.1348e-08
| 100 | 20 | 10 | 0 | 100 | 20 | 20 | 10 | 0 | 50.0283699E+00 | DC2_2V | DC11V | 0.030005 | 4.6575e-08
| 100 | 25 | 10 | 5 | 100 | 20 | 20 | 10 | 0 | 50.0458392E+00 | DC2_2V | DC11V | 0.030004 | 3.8575e-08
| 100 | 25 | 10 | 0 | 100 | 25 | 20 | 10 | 0 | 50.0737454E+00 | DC2_2V | DC11V | 0.040004 | 4.146e-08
| 100 | 30 | 10 | 5 | 100 | 25 | 20 | 10 | 0 | 50.0049297E+00 | DC2_2V | DC2_2V | 0.040004 | 5.3093e-08
| 100 | 30 | 10 | 0 | 100 | 30 | 20 | 10 | 0 | 49.9964012E+00 | DC2_2V | DC2_2V | 0.050004 | 4.4105e-08
| 100 | 35 | 10 | 5 | 100 | 30 | 20 | 10 | 0 | 49.9888138E+00 | DC2_2V | DC2_2V | 0.050003 | 4.989e-08
| 100 | 35 | 10 | 0 | 100 | 35 | 20 | 10 | 0 | 49.9775714E+00 | DC2_2V | DC2_2V | 0.060003 | 4.2866e-08
| 100 | 40 | 10 | 5 | 100 | 35 | 20 | 10 | 0 | 50.0598120E+00 | DC2_2V | DC2_2V | 0.060003 | 4.4228e-08
| 100 | 40 | 10 | 0 | 100 | 40 | 20 | 10 | 0 | 49.9908837E+00 | DC2_2V | DC2_2V | 0.070003 | 7.9555e-08
| 100 | 45 | 10 | 5 | 100 | 40 | 20 | 10 | 0 | 50.0086666E+00 | DC2_2V | DC11V | 0.070002 | 4.1837e-08
| 100 | 45 | 10 | 0 | 100 | 45 | 20 | 10 | 0 | 49.9911761E+00 | DC2_2V | DC11V | 0.080002 | 7.1368e-08
| 100 | 50 | 10 | 5 | 100 | 45 | 20 | 10 | 0 | 50.0189883E+00 | DC2_2V | DC11V | 0.080002 | 5.1504e-08
| 100 | 50 | 10 | 0 | 100 | 50 | 20 | 10 | 0 | 50.0652170E+00 | DC2_2V | DC11V | 0.090002 | 4.66e-08
| 100 | 50 | 10 | 5 | 100 | 45 | 20 | 10 | 0 | 50.0313168E+00 | DC2_2V | DC11V | 0.090002 | 3.2379e-08
| 100 | 45 | 10 | 0 | 100 | 45 | 20 | 10 | 0 | 49.9904169E+00 | DC2_2V | DC11V | 0.1 | 5.4103e-08
| 100 | 45 | 10 | 5 | 100 | 40 | 20 | 10 | 0 | 49.9616662E+00 | DC2_2V | DC11V | 0.090002 | 5.163e-08
| 100 | 40 | 10 | 0 | 100 | 40 | 20 | 10 | 0 | 49.9760822E+00 | DC2_2V | DC11V | 0.090002 | 5.37e-07
| 100 | 40 | 10 | 5 | 100 | 35 | 20 | 10 | 0 | 50.0089497E+00 | DC2_2V | DC11V | 0.080002 | 4.9294e-08
| 100 | 35 | 10 | 0 | 100 | 35 | 20 | 10 | 0 | 50.0098227E+00 | DC2_2V | DC11V | 0.080003 | 4.2433e-08
| 100 | 35 | 10 | 5 | 100 | 30 | 20 | 10 | 0 | 50.0548715E+00 | DC2_2V | DC11V | 0.070003 | 5.5564e-08
| 100 | 30 | 10 | 0 | 100 | 30 | 20 | 10 | 0 | 50.0107523E+00 | DC2_2V | DC11V | 0.070003 | 4.4479e-08
| 100 | 30 | 10 | 5 | 100 | 25 | 20 | 10 | 0 | 50.0186626E+00 | DC2_2V | DC11V | 0.060003 | 3.8506e-08
| 100 | 25 | 10 | 0 | 100 | 25 | 20 | 10 | 0 | 50.0054346E+00 | DC2_2V | DC11V | 0.060004 | 5.6544e-08
| 100 | 25 | 10 | 5 | 100 | 20 | 20 | 10 | 0 | 49.9804275E+00 | DC2_2V | DC11V | 0.050004 | 4.5149e-08
| 100 | 20 | 10 | 0 | 100 | 20 | 20 | 10 | 0 | 50.0244601E+00 | DC2_2V | DC11V | 0.050004 | 5.6065e-08
| 100 | 20 | 10 | 5 | 100 | 15 | 20 | 10 | 0 | 49.9897568E+00 | DC2_2V | DC11V | 0.040004 | 6.4814e-08
| 100 | 15 | 10 | 0 | 100 | 15 | 20 | 10 | 0 | 50.0130742E+00 | DC2_2V | DC11V | 0.040005 | 4.3282e-08
| 100 | 15 | 10 | 5 | 100 | 10 | 20 | 10 | 0 | 49.9879746E+00 | DC2_2V | DC11V | 0.030005 | 3.3781e-08
| 100 | 10 | 10 | 0 | 100 | 10 | 20 | 10 | 0 | 49.9643792E+00 | DC2_2V | DC11V | 0.030005 | 4.149e-08
| 100 | 10 | 10 | 5 | 100 | 5 | 20 | 10 | 0 | 49.9486471E+00 | DC2_2V | DC11V | 0.020005 | 5.0905e-08
| 100 | 5 | 10 | 0 | 100 | 5 | 20 | 10 | 0 | 49.9932603E+00 | DC2_2V | DC11V | 0.020005 | 3.9469e-08
| 100 | 5 | 10 | 0 | 100 | 5 | 20 | 10 | 0 | 50.0016038E+00 | DC2_2V | DC11V | 0.010005 | 3.5281e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 50.0230484E+00 | DC2_2V | DC11V | 0.010006 | 3.0983e-08
| 100 | 5 | 10 | 0 | 100 | 5 | 20 | 10 | 0 | 49.9626506E+00 | DC2_2V | DC11V | 5.4446e-06 | 5.5812e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 49.9980991E+00 | DC2_2V | DC11V | 5.4484e-06 | 8.7916e-08
| 100 | 5 | 10 | 0 | 100 | 5 | 20 | 10 | 0 | 49.9965851E+00 | DC2_2V | DC11V | 4.7549e-06 | 6.0539e-08
| 100 | 10 | 10 | 0 | 100 | 10 | 20 | 10 | 0 | 49.9741407E+00 | DC2_2V | DC11V | 4.8223e-06 | 5.8733e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 50.0042975E+00 | DC2_2V | DC11V | 0.010004 | 4.5554e-08
| 100 | 10 | 10 | 10 | 100 | 0 | 20 | 10 | 0 | 49.9474892E+00 | DC2_2V | DC11V | -0.0099949 | 2.9958e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 50.0383738E+00 | DC2_2V | DC11V | 4.6438e-06 | 6.5219e-08
| 100 | 10 | 10 | 10 | 100 | 0 | 20 | 10 | 0 | 49.9884791E+00 | DC2_2V | DC11V | 5.0585e-06 | 2.973e-08
| 100 | 10 | 10 | 0 | 100 | 10 | 20 | 10 | 0 | 50.0185021E+00 | DC2_2V | DC11V | 4.5116e-06 | 4.2838e-08
| 100 | 20 | 10 | 10 | 100 | 10 | 20 | 10 | 0 | 50.0553773E+00 | DC2_2V | DC11V | 4.9348e-06 | 3.9575e-08
| 100 | 20 | 10 | 0 | 100 | 20 | 20 | 10 | 0 | 49.9801353E+00 | DC2_2V | DC11V | -0.0099951 | 3.8644e-08
| 100 | 30 | 10 | 10 | 100 | 20 | 20 | 10 | 0 | 49.9973822E+00 | DC2_2V | DC11V | -0.0099946 | 6.0552e-08
| 100 | 30 | 10 | 0 | 100 | 30 | 20 | 10 | 0 | 50.0116820E+00 | DC2_2V | DC11V | -0.019995 | 7.5968e-08
| 100 | 40 | 10 | 10 | 100 | 30 | 20 | 10 | 0 | 50.0381470E+00 | DC2_2V | DC11V | -0.019994 | 3.544e-08
| 100 | 40 | 10 | 0 | 100 | 40 | 20 | 10 | 0 | 49.9937932E+00 | DC2_2V | DC11V | -0.029994 | 4.0054e-08
| 100 | 50 | 10 | 10 | 100 | 40 | 20 | 10 | 0 | 50.0414069E+00 | DC2_2V | DC11V | -0.029994 | 3.9743e-08
| 100 | 50 | 10 | 0 | 100 | 50 | 20 | 10 | 0 | 50.0424417E+00 | DC2_2V | DC11V | -0.039994 | 5.294e-08
| 100 | 60 | 10 | 10 | 100 | 50 | 20 | 10 | 0 | 49.9540752E+00 | DC2_2V | DC2_2V | -0.039993 | 3.5819e-08
| 100 | 60 | 10 | 0 | 100 | 60 | 20 | 10 | 0 | 49.9572441E+00 | DC2_2V | DC2_2V | -0.049993 | 4.8468e-08
| 100 | 70 | 10 | 10 | 100 | 60 | 20 | 10 | 0 | 50.0156511E+00 | DC2_2V | DC2_2V | -0.049993 | 3.4955e-08
| 100 | 70 | 10 | 0 | 100 | 70 | 20 | 10 | 0 | 50.0131308E+00 | DC220MV | DC220MV | -0.059993 | 4.5685e-08
| 100 | 80 | 10 | 10 | 100 | 70 | 20 | 10 | 0 | 49.9719827E+00 | DC220MV | DC220MV | -0.059993 | 3.6657e-08
| 100 | 80 | 10 | 0 | 100 | 80 | 20 | 10 | 0 | 50.0056421E+00 | DC220MV | DC220MV | -0.069993 | 5.076e-08
| 100 | 90 | 10 | 10 | 100 | 80 | 20 | 10 | 0 | 49.9625988E+00 | DC220MV | DC220MV | -0.069992 | 4.5716e-08
| 100 | 90 | 10 | 0 | 100 | 90 | 20 | 10 | 0 | 50.0264432E+00 | DC220MV | DC220MV | -0.079993 | 5.4041e-08
| 100 | 100 | 10 | 10 | 100 | 90 | 20 | 10 | 0 | 50.0096339E+00 | DC220MV | DC220MV | -0.079992 | 4.7165e-08
| 100 | 100 | 10 | 0 | 100 | 100 | 20 | 10 | 0 | 49.9681288E+00 | DC220MV | DC220MV | -0.089992 | 4.5394e-08
| 100 | 100 | 10 | 10 | 100 | 90 | 20 | 10 | 0 | 49.9975426E+00 | DC220MV | DC220MV | -0.089992 | 2.9758e-08
| 100 | 90 | 10 | 0 | 100 | 90 | 20 | 10 | 0 | 50.0329037E+00 | DC220MV | DC220MV | -0.099992 | 4.8841e-08
| 100 | 90 | 10 | 10 | 100 | 80 | 20 | 10 | 0 | 50.0029530E+00 | DC220MV | DC220MV | -0.089992 | 3.9385e-08
| 100 | 80 | 10 | 0 | 100 | 80 | 20 | 10 | 0 | 49.9937083E+00 | DC220MV | DC220MV | -0.089992 | 7.4541e-08
| 100 | 80 | 10 | 10 | 100 | 70 | 20 | 10 | 0 | 50.0149054E+00 | DC220MV | DC220MV | -0.079992 | 3.2109e-08
| 100 | 70 | 10 | 0 | 100 | 70 | 20 | 10 | 0 | 49.9931330E+00 | DC220MV | DC220MV | -0.079993 | 5.3007e-08
| 100 | 70 | 10 | 10 | 100 | 60 | 20 | 10 | 0 | 49.9917466E+00 | DC220MV | DC220MV | -0.069993 | 4.2062e-08
| 100 | 60 | 10 | 0 | 100 | 60 | 20 | 10 | 0 | 50.0594385E+00 | DC220MV | DC220MV | -0.069993 | 3.84e-08
| 100 | 60 | 10 | 10 | 100 | 50 | 20 | 10 | 0 | 49.9829162E+00 | DC220MV | DC220MV | -0.059993 | 4.2363e-08
| 100 | 50 | 10 | 0 | 100 | 50 | 20 | 10 | 0 | 49.9492356E+00 | DC220MV | DC220MV | -0.059993 | 4.3653e-08
| 100 | 50 | 10 | 10 | 100 | 40 | 20 | 10 | 0 | 50.0236433E+00 | DC220MV | DC220MV | -0.049993 | 6.0246e-08
| 100 | 40 | 10 | 0 | 100 | 40 | 20 | 10 | 0 | 49.9746355E+00 | DC220MV | DC220MV | -0.049993 | 3.7076e-08
| 100 | 40 | 10 | 10 | 100 | 30 | 20 | 10 | 0 | 49.9932698E+00 | DC220MV | DC220MV | -0.039993 | 4.2491e-08
| 100 | 30 | 10 | 0 | 100 | 30 | 20 | 10 | 0 | 49.9617510E+00 | DC220MV | DC220MV | -0.039994 | 4.7936e-08
| 100 | 30 | 10 | 10 | 100 | 20 | 20 | 10 | 0 | 49.9962880E+00 | DC220MV | DC220MV | -0.029994 | 6.0326e-08
| 100 | 20 | 10 | 0 | 100 | 20 | 20 | 10 | 0 | 49.9896012E+00 | DC220MV | DC220MV | -0.029994 | 2.7836e-08
| 100 | 20 | 10 | 10 | 100 | 10 | 20 | 10 | 0 | 50.0031417E+00 | DC220MV | DC220MV | -0.019994 | 4.6542e-08
| 100 | 10 | 10 | 0 | 100 | 10 | 20 | 10 | 0 | 49.9789805E+00 | DC220MV | DC220MV | -0.019995 | 4.1285e-08
| 100 | 10 | 10 | 10 | 100 | 0 | 20 | 10 | 0 | 50.0171238E+00 | DC220MV | DC220MV | -0.0099947 | 4.8006e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 49.9998915E+00 | DC220MV | DC220MV | -0.0099951 | 4.321e-08
| 100 | 10 | 10 | 10 | 100 | 0 | 20 | 10 | 0 | 49.9831943E+00 | DC220MV | DC220MV | 4.9641e-06 | 3.9357e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 49.9900492E+00 | DC220MV | DC220MV | 4.5697e-06 | 7.0682e-08
| 100 | 10 | 10 | 0 | 100 | 10 | 20 | 10 | 0 | 49.9985614E+00 | DC220MV | DC220MV | 5.163e-06 | 4.253e-08
| 100 | -10 | 10 | 0 | 100 | -10 | 20 | 10 | 0 | 50.0033021E+00 | DC220MV | DC220MV | 4.6397e-06 | 3.1806e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 49.9917985E+00 | DC220MV | DC220MV | -0.009995 | 4.4132e-08
| 100 | -10 | 10 | -10 | 100 | 0 | 20 | 10 | 0 | 49.9843775E+00 | DC220MV | DC220MV | 0.1 | 3.0148e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 50.0249511E+00 | DC220MV | DC220MV | 4.6034e-06 | 3.6841e-08
| 100 | -10 | 10 | -10 | 100 | 0 | 20 | 10 | 0 | 49.9716387E+00 | DC220MV | DC220MV | 3.8322e-07 | 3.3318e-08
| 100 | -10 | 10 | 0 | 100 | -10 | 20 | 10 | 0 | 49.9675352E+00 | DC220MV | DC220MV | 4.4662e-06 | 4.7154e-08
| 100 | -20 | 10 | -10 | 100 | -10 | 20 | 10 | 0 | 50.0280393E+00 | DC220MV | DC220MV | 0.010005 | 5.2663e-08
| 100 | -20 | 10 | 0 | 100 | -20 | 20 | 10 | 0 | 50.0554009E+00 | DC220MV | DC220MV | 5.6183e-06 | 3.6452e-08
| 100 | -30 | 10 | -10 | 100 | -20 | 20 | 10 | 0 | 49.9716906E+00 | DC220MV | DC220MV | 5.1692e-06 | 4.3858e-08
| 100 | -30 | 10 | 0 | 100 | -30 | 20 | 10 | 0 | 49.9874135E+00 | DC220MV | DC220MV | 5.5943e-06 | 3.7356e-08
| 100 | -40 | 10 | -10 | 100 | -30 | 20 | 10 | 0 | 49.9689156E+00 | DC220MV | DC220MV | 5.2605e-06 | 3.5703e-08
| 100 | -40 | 10 | 0 | 100 | -40 | 20 | 10 | 0 | 49.9946185E+00 | DC220MV | DC220MV | 0.010005 | 3.5711e-08
| 100 | -50 | 10 | -10 | 100 | -40 | 20 | 10 | 0 | 49.9880877E+00 | DC220MV | DC220MV | 0.010005 | 3.7356e-08
| 100 | -50 | 10 | 0 | 100 | -50 | 20 | 10 | 0 | 50.0087987E+00 | DC220MV | DC220MV | 0.020005 | 4.5103e-08
| 100 | -60 | 10 | -10 | 100 | -50 | 20 | 10 | 0 | 49.9794990E+00 | DC220MV | DC220MV | 0.020005 | 5.1348e-08
| 100 | -60 | 10 | 0 | 100 | -60 | 20 | 10 | 0 | 49.9480446E+00 | DC220MV | DC220MV | 0.030005 | 4.6575e-08
| 100 | -70 | 10 | -10 | 100 | -60 | 20 | 10 | 0 | 50.0089497E+00 | DC220MV | DC220MV | 0.030004 | 3.8575e-08
| 100 | -70 | 10 | 0 | 100 | -70 | 20 | 10 | 0 | 49.9811958E+00 | DC220MV | DC220MV | 0.040004 | 4.146e-08
| 100 | -80 | 10 | -10 | 100 | -70 | 20 | 10 | 0 | 49.9946138E+00 | DC220MV | DC220MV | 0.040004 | 5.3093e-08
| 100 | -80 | 10 | 0 | 100 | -80 | 20 | 10 | 0 | 50.0182567E+00 | DC220MV | DC220MV | 0.050004 | 4.4105e-08
| 100 | -90 | 10 | -10 | 100 | -80 | 20 | 10 | 0 | 50.0069538E+00 | DC220MV | DC220MV | 0.050003 | 4.989e-08
| 100 | -90 | 10 | 0 | 100 | -90 | 20 | 10 | 0 | 49.9680582E+00 | DC220MV | DC220MV | 0.060003 | 4.2866e-08
| 100 | -100 | 10 | -10 | 100 | -90 | 20 | 10 | 0 | 50.0065008E+00 | DC220MV | DC220MV | 0.060003 | 4.4228e-08
| 100 | -100 | 10 | 0 | 100 | -100 | 20 | 10 | 0 | 50.0455414E+00 | DC220MV | DC220MV | 0.070003 | 7.9555e-08
| 100 | -100 | 10 | -10 | 100 | -90 | 20 | 10 | 0 | 50.0172890E+00 | DC220MV | DC220MV | 0.070002 | 4.1837e-08
| 100 | -90 | 10 | 0 | 100 | -90 | 20 | 10 | 0 | 50.0313923E+00 | DC220MV | DC220MV | 0.080002 | 7.1368e-08
| 100 | -90 | 10 | -10 | 100 | -80 | 20 | 10 | 0 | 50.0098180E+00 | DC220MV | DC220MV | 0.080002 | 5.1504e-08
| 100 | -80 | 10 | 0 | 100 | -80 | 20 | 10 | 0 | 49.9871212E+00 | DC220MV | DC220MV | 0.090002 | 4.66e-08
| 100 | -80 | 10 | -10 | 100 | -70 | 20 | 10 | 0 | 49.9888091E+00 | DC220MV | DC220MV | 0.090002 | 3.2379e-08
| 100 | -70 | 10 | 0 | 100 | -70 | 20 | 10 | 0 | 50.0417802E+00 | DC220MV | DC220MV | 0.1 | 5.4103e-08
| 100 | -70 | 10 | -10 | 100 | -60 | 20 | 10 | 0 | 49.9707247E+00 | DC220MV | DC220MV | 0.090002 | 5.163e-08
| 100 | -60 | 10 | 0 | 100 | -60 | 20 | 10 | 0 | 49.9903839E+00 | DC220MV | DC220MV | 0.090002 | 5.37e-07
| 100 | -60 | 10 | -10 | 100 | -50 | 20 | 10 | 0 | 50.0350529E+00 | DC220MV | DC220MV | 0.080002 | 4.9294e-08
| 100 | -50 | 10 | 0 | 100 | -50 | 20 | 10 | 0 | 49.9981086E+00 | DC220MV | DC220MV | 0.080003 | 4.2433e-08
| 100 | -50 | 10 | -10 | 100 | -40 | 20 | 10 | 0 | 49.9734151E+00 | DC220MV | DC220MV | 0.070003 | 5.5564e-08
| 100 | -40 | 10 | 0 | 100 | -40 | 20 | 10 | 0 | 50.0092753E+00 | DC220MV | DC220MV | 0.070003 | 4.4479e-08
| 100 | -40 | 10 | -10 | 100 | -30 | 20 | 10 | 0 | 49.9634418E+00 | DC220MV | DC220MV | 0.060003 | 3.8506e-08
| 100 | -30 | 10 | 0 | 100 | -30 | 20 | 10 | 0 | 50.0123144E+00 | DC220MV | DC220MV | 0.060004 | 5.6544e-08
| 100 | -30 | 10 | -10 | 100 | -20 | 20 | 10 | 0 | 49.9921569E+00 | DC220MV | DC220MV | 0.050004 | 4.5149e-08
| 100 | -20 | 10 | 0 | 100 | -20 | 20 | 10 | 0 | 49.9707859E+00 | DC220MV | DC220MV | 0.050004 | 5.6065e-08
| 100 | -20 | 10 | -10 | 100 | -10 | 20 | 10 | 0 | 50.0001604E+00 | DC220MV | DC220MV | 0.040004 | 6.4814e-08
| 100 | -10 | 10 | 0 | 100 | -10 | 20 | 10 | 0 | 49.9567874E+00 | DC220MV | DC220MV | 0.040005 | 4.3282e-08
| 100 | -10 | 10 | -10 | 100 | 0 | 20 | 10 | 0 | 50.0067839E+00 | DC220MV | DC220MV | 0.030005 | 3.3781e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 49.9751774E+00 | DC220MV | DC220MV | 0.030005 | 4.149e-08
| 100 | -10 | 10 | -10 | 100 | 0 | 20 | 10 | 0 | 49.9859945E+00 | DC220MV | DC220MV | 0.020005 | 5.0905e-08
| 100 | 0 | 10 | 0 | 100 | 0 | 20 | 10 | 0 | 50.0301172E+00 | DC220MV | DC220MV | 0.020005 | 3.9469e-08
| 100 | -10 | 10 | 0 | 100 | -10 | 20 | 10 | 0 | 50.0593676E+00 | DC220MV | DC220MV | 0.010005 | 3.5281e-08