## Sheet
instruments: | Meter: HP | S: FL5700 | X: FL5730
Start Row | 823.0 | Stop Row | 1009.0
| X Range | X Settings (V) | S Range | S Settings (V) | DVM Range | Nominal reading | #Readings | Delay (S) | DVM pause | DVM status | S status | X status | Mean | STD | start time | end time | readings...
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9535715E+00 | DC220MV |  | -1.0000023357000001 | 1.9534792038259174e-07 | 1517345663.76 | 1517345679.191 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0427063E+00 | DC220MV |  | -8.2218296582499996e-08 | 2.5669755682787186e-08 | 1517345694.055 | 1517345709.63 | -8.5803e-08 | -8.1102e-08 | -5.8182e-08 | -5.2305e-08 | -6.3177e-08 | -9.2268e-08 | -9.6088e-08 | -6.0239e-08 | -1.3238e-07 | -1.2297e-07 | -1.0226e-07 | -7.7722e-08 | -1.0799e-07 | -4.7309e-08 | -3.8053e-08 | -7.0817e-08 | -9.4325e-08 | -5.3333e-08 | -9.8586e-08 | -1.0946e-07
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0002123E+00 | DC2_2V |  | 3.6134875827000006e-06 | 7.4247543390603461e-08 | 1517345724.476 | 1517345739.986 | 3.6166e-06 | 3.5994e-06 | 3.6408e-06 | 3.6341e-06 | 3.6265e-06 | 3.591e-06 | 3.8415e-06 | 3.5495e-06 | 3.5498e-06 | 3.6973e-06 | 3.5564e-06 | 3.6527e-06 | 3.5002e-06 | 3.5773e-06 | 3.5836e-06 | 3.5963e-06 | 3.5919e-06 | 3.6982e-06 | 3.6589e-06 | 3.5077e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0120265E+00 | DC220MV |  | -8.5740726350499996e-08 | 2.4364197857589055e-08 | 1517345754.836 | 1517345770.391 | -2.2499e-08 | -5.3969e-08 | -6.9851e-08 | -7.3087e-08 | -6.7646e-08 | -9.2204e-08 | -1.347e-07 | -7.8969e-08 | -8.9263e-08 | -1.1309e-07 | -6.8528e-08 | -9.4704e-08 | -9.9851e-08 | -1.075e-07 | -9.8968e-08 | -7.8086e-08 | -7.6616e-08 | -7.4998e-08 | -9.9998e-08 | -1.2029e-07
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0043164E+00 | DC2_2V |  | 3.6383633989999997e-06 | 7.4422464056002899e-08 | 1517345785.25 | 1517345800.771 | 3.703e-06 | 3.7251e-06 | 3.6311e-06 | 3.6405e-06 | 3.6698e-06 | 3.6643e-06 | 3.5812e-06 | 3.5374e-06 | 3.6723e-06 | 3.82e-06 | 3.5201e-06 | 3.7155e-06 | 3.6492e-06 | 3.563e-06 | 3.6495e-06 | 3.656e-06 | 3.5073e-06 | 3.654e-06 | 3.5605e-06 | 3.6474e-06
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9816954E+00 | DC220MV |  | -1.00000236385 | 7.9610473554376901e-08 | 1517345815.626 | 1517345831.042 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -2.0 | 2.0 | -1.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9829539E+00 | DC2_2V |  | -0.99999909190499969 | 1.0025286767189055e-07 | 1517345845.903 | 1517345861.361 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -2.0 | 2.0 | 0.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 49.9874796E+00 | DC220MV |  | -2.0000058254999997 | 1.3849061337267953e-07 | 1517345876.222 | 1517345891.651 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -3.0 | 2.0 | -1.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 50.0017076E+00 | DC2_2V |  | -2.0000043452000003 | 2.0370508098711781e-07 | 1517345906.519 | 1517345921.95 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -3.0 | 2.0 | 0.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 49.9583696E+00 | DC220MV |  | -3.0000104454500001 | 1.6639245026230336e-07 | 1517345936.807 | 1517345952.239 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -4.0 | 2.0 | -1.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 49.9852308E+00 | DC2_2V |  | -3.0000070303499995 | 1.5645455412282203e-07 | 1517345967.088 | 1517345982.525 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -4.0 | 2.0 | 0.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9736036E+00 | DC220MV |  | -4.0000134441999986 | 1.9953059917721649e-07 | 1517345997.388 | 1517346012.827 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -5.0 | 2.0 | -1.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 50.0132252E+00 | DC2_2V |  | -4.0000102880499995 | 2.099632051234827e-07 | 1517346027.673 | 1517346043.108 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -5.0 | 2.0 | 0.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 50.0131072E+00 | DC220MV |  | -5.0000169304500002 | 1.9906367708209015e-07 | 1517346057.967 | 1517346073.392 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -6.0 | 2.0 | -1.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 50.0361015E+00 | DC2_2V |  | -5.0000136568000002 | 2.0449342285442103e-07 | 1517346088.245 | 1517346103.684 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -6.0 | 2.0 | 0.0 | 10 | -6.0 | 20.0 | 10.0 | 0.0 | 50.0145703E+00 | DC220MV |  | -6.0000203579499995 | 2.5858586088466859e-07 | 1517346118.526 | 1517346133.964 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6
| 20 | -7.0 | 2.0 | -1.0 | 10 | -6.0 | 20.0 | 10.0 | 0.0 | 49.9735753E+00 | DC2_2V |  | -6.0000169586999998 | 2.7688067833709047e-07 | 1517346148.824 | 1517346164.241 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6
| 20 | -7.0 | 2.0 | 0.0 | 10 | -7.0 | 20.0 | 10.0 | 0.0 | 49.9613978E+00 | DC220MV |  | -7.0000236946000003 | 2.1226031183114342e-07 | 1517346179.079 | 1517346194.526 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7
| 20 | -8.0 | 2.0 | -1.0 | 10 | -7.0 | 20.0 | 10.0 | 0.0 | 50.0560722E+00 | DC2_2V |  | -7.0000204888000015 | 2.0641816771755593e-07 | 1517346209.385 | 1517346224.835 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7
| 20 | -8.0 | 2.0 | 0.0 | 10 | -8.0 | 20.0 | 10.0 | 0.0 | 49.9813277E+00 | DC220MV |  | -8.000027309550001 | 2.9614024975176213e-07 | 1517346239.682 | 1517346255.082 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8
| 20 | -9.0 | 2.0 | -1.0 | 10 | -8.0 | 20.0 | 10.0 | 0.0 | 50.0227556E+00 | DC2_2V |  | -8.0000242855499994 | 3.5726915260971461e-07 | 1517346269.917 | 1517346285.362 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8
| 20 | -9.0 | 2.0 | 0.0 | 10 | -9.0 | 20.0 | 10.0 | 0.0 | 49.9221121E+00 | DC220MV |  | -9.0000311811000007 | 5.1254754892725572e-07 | 1517346300.218 | 1517346315.642 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9
| 20 | -10.0 | 2.0 | -1.0 | 10 | -9.0 | 20.0 | 10.0 | 0.0 | 50.0045381E+00 | DC2_2V |  | -9.0000284135499999 | 4.4289225275427356e-07 | 1517346330.489 | 1517346345.957 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9
| 20 | -10.0 | 2.0 | 0.0 | 10 | -10.0 | 20.0 | 10.0 | 0.0 | 49.9689815E+00 | DC220MV |  | -10.000034982499999 | 2.5664907958234749e-07 | 1517346360.801 | 1517346376.209 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10
| 20 | -10.0 | 2.0 | -1.0 | 10 | -9.0 | 20.0 | 10.0 | 0.0 | 49.9924021E+00 | DC2_2V |  | -9.0000283749500021 | 3.2709470713204278e-07 | 1517346391.046 | 1517346406.492 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9
| 20 | -9.0 | 2.0 | 0.0 | 10 | -9.0 | 20.0 | 10.0 | 0.0 | 49.9949109E+00 | DC220MV |  | -9.000032663199999 | 3.0063908598954003e-07 | 1517346421.367 | 1517346436.802 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9
| 20 | -9.0 | 2.0 | -1.0 | 10 | -8.0 | 20.0 | 10.0 | 0.0 | 49.9338612E+00 | DC2_2V |  | -8.0000257691000005 | 3.5425427314870656e-07 | 1517346451.631 | 1517346467.056 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8
| 20 | -8.0 | 2.0 | 0.0 | 10 | -8.0 | 20.0 | 10.0 | 0.0 | 50.0089969E+00 | DC220MV |  | -8.0000287166500001 | 2.8492284493736587e-07 | 1517346481.884 | 1517346497.337 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8
| 20 | -8.0 | 2.0 | -1.0 | 10 | -7.0 | 20.0 | 10.0 | 0.0 | 49.9995802E+00 | DC2_2V |  | -7.0000219565999995 | 1.6606938309583565e-07 | 1517346512.196 | 1517346527.624 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7
| 20 | -7.0 | 2.0 | 0.0 | 10 | -7.0 | 20.0 | 10.0 | 0.0 | 49.9633382E+00 | DC220MV |  | -7.0000251137999996 | 3.0964166385372145e-07 | 1517346542.484 | 1517346557.92 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7
| 20 | -7.0 | 2.0 | -1.0 | 10 | -6.0 | 20.0 | 10.0 | 0.0 | 49.9884555E+00 | DC2_2V |  | -6.0000183481500002 | 2.3498665397755056e-07 | 1517346572.779 | 1517346588.222 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6
| 20 | -6.0 | 2.0 | 0.0 | 10 | -6.0 | 20.0 | 10.0 | 0.0 | 49.9886960E+00 | DC220MV |  | -6.0000219784499995 | 2.6731189178694465e-07 | 1517346603.08 | 1517346618.52 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6
| 20 | -6.0 | 2.0 | -1.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 50.0164016E+00 | DC2_2V |  | -5.0000151424999997 | 2.2743097849554484e-07 | 1517346633.364 | 1517346648.795 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -5.0 | 2.0 | 0.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 50.0247245E+00 | DC220MV |  | -5.0000183269999994 | 2.6226227335142032e-07 | 1517346663.618 | 1517346679.034 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -5.0 | 2.0 | -1.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9333672E+00 | DC2_2V |  | -4.0000115929 | 1.994188807304574e-07 | 1517346693.887 | 1517346709.306 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -4.0 | 2.0 | 0.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9592361E+00 | DC220MV |  | -4.0000148688000001 | 3.2168363342964815e-07 | 1517346724.143 | 1517346739.606 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -4.0 | 2.0 | -1.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 50.0547817E+00 | DC2_2V |  | -3.0000080929000004 | 2.2866829683490418e-07 | 1517346754.467 | 1517346769.902 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -3.0 | 2.0 | 0.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 49.9920249E+00 | DC220MV |  | -3.0000113059500002 | 2.4849395867447437e-07 | 1517346784.753 | 1517346800.146 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -3.0 | 2.0 | -1.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 50.0290735E+00 | DC2_2V |  | -2.0000045011999998 | 2.7182799704692674e-07 | 1517346814.988 | 1517346830.423 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -2.0 | 2.0 | 0.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 49.9994764E+00 | DC220MV |  | -2.0000063064999996 | 1.2672470160727269e-07 | 1517346845.245 | 1517346860.65 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -2.0 | 2.0 | -1.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9805076E+00 | DC2_2V |  | -0.9999996684000001 | 9.4141356477245312e-08 | 1517346875.521 | 1517346890.949 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 50.0471057E+00 | DC220MV |  | -1.0000033442 | 1.7621736578090778e-07 | 1517346905.831 | 1517346921.268 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9770247E+00 | DC2_2V |  | 3.23580806925e-06 | 4.6090787113476329e-08 | 1517346936.128 | 1517346951.63 | 3.2947e-06 | 3.2489e-06 | 3.1977e-06 | 3.1395e-06 | 3.2526e-06 | 3.303e-06 | 3.2239e-06 | 3.2735e-06 | 3.1464e-06 | 3.208e-06 | 3.2818e-06 | 3.2448e-06 | 3.2838e-06 | 3.2473e-06 | 3.2964e-06 | 3.2023e-06 | 3.2514e-06 | 3.2137e-06 | 3.1843e-06 | 3.222e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0308114E+00 | DC220MV |  | -3.2375953064499994e-07 | 4.381228810464312e-08 | 1517346966.532 | 1517346982.074 | -3.0194e-07 | -3.0988e-07 | -3.2827e-07 | -3.3268e-07 | -3.5724e-07 | -2.3061e-07 | -2.8165e-07 | -3.3239e-07 | -2.7341e-07 | -2.662e-07 | -2.9871e-07 | -2.8165e-07 | -3.7489e-07 | -3.518e-07 | -3.4077e-07 | -3.0709e-07 | -3.4136e-07 | -4.1151e-07 | -3.8577e-07 | -3.6739e-07
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9696599E+00 | DC2_2V |  | 3.4067471956500009e-06 | 7.0216041402931589e-08 | 1517346996.92 | 1517347012.432 | 3.3695e-06 | 3.4621e-06 | 3.3224e-06 | 3.4649e-06 | 3.4519e-06 | 3.4379e-06 | 3.5472e-06 | 3.4669e-06 | 3.3789e-06 | 3.4129e-06 | 3.3183e-06 | 3.3823e-06 | 3.3835e-06 | 3.3436e-06 | 3.4591e-06 | 3.3919e-06 | 3.2693e-06 | 3.5189e-06 | 3.4371e-06 | 3.3164e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0178790E+00 | DC220MV |  | -3.56241330145e-07 | 4.0758549847929478e-08 | 1517347027.27 | 1517347042.826 | -3.4376e-07 | -4.2078e-07 | -4.0828e-07 | -3.0908e-07 | -3.2965e-07 | -3.2333e-07 | -3.783e-07 | -3.8198e-07 | -2.7528e-07 | -3.6743e-07 | -3.3509e-07 | -3.736e-07 | -3.6596e-07 | -3.4185e-07 | -4.4341e-07 | -3.711e-07 | -3.37e-07 | -3.3877e-07 | -3.8638e-07 | -2.9379e-07
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 50.0249275E+00 | DC220MV |  | -1.0000030245000002 | 6.0924133136714773e-08 | 1517347057.682 | 1517347073.098 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 100 | 5.0 | 10 | 0.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 49.9652553E+00 | DC11V |  | 4.9999974464499983 | 8.0708620262690077e-06 | 1517347088.64 | 1517347103.738 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9721570E+00 | DC11V |  | -5.8647250024100002e-06 | 7.8205700195666361e-06 | 1517347118.569 | 1517347133.727 | -6.3283e-06 | 5.151e-06 | -5.151e-06 | -2.3547e-06 | -1.2215e-05 | -6.4755e-06 | 2.6491e-06 | -7.0642e-06 | -5.8868e-06 | -7.6528e-06 | -1.3687e-05 | 4.1208e-06 | -2.2076e-05 | 0 | 6.7698e-06 | -1.5306e-05 | -7.3585e-07 | -1.457e-05 | 4.4151e-07 | -1.6925e-05
| 100 | 5.0 | 10 | 5.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0219861E+00 | DC11V |  | -2.7598976391500003e-05 | 5.9603300771590576e-06 | 1517347148.561 | 1517347163.724 | -2.9431e-05 | -3.0461e-05 | -2.8107e-05 | -2.8548e-05 | -3.3404e-05 | -2.8107e-05 | -1.8983e-05 | -2.6194e-05 | -2.3839e-05 | -3.1344e-05 | -1.9572e-05 | -2.796e-05 | -2.016e-05 | -1.5451e-05 | -2.4428e-05 | -2.899e-05 | -2.5752e-05 | -3.9143e-05 | -3.6495e-05 | -3.5612e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9883612E+00 | DC11V |  | -1.3083708401615002e-05 | 1.079401347662859e-05 | 1517347178.55 | 1517347193.7 | -2.3414e-05 | -4.4177e-07 | -5.5957e-06 | -8.9826e-06 | -1.3695e-05 | -9.8662e-06 | -7.5101e-06 | -3.2249e-05 | -8.8354e-06 | -1.3106e-05 | 8.3936e-06 | -1.9585e-05 | -1.7376e-05 | -3.7403e-05 | 3.0924e-06 | -9.5716e-06 | -1.1928e-05 | -2.5181e-05 | -1.7818e-05 | -1.0602e-05
| 100 | 5.0 | 10 | 5.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9522391E+00 | DC11V |  | -3.3469967163499998e-05 | 8.0392429021088073e-06 | 1517347208.531 | 1517347223.699 | -3.4132e-05 | -2.707e-05 | -2.7217e-05 | -2.3392e-05 | -4.1341e-05 | -2.5011e-05 | -5.3258e-05 | -3.7369e-05 | -2.913e-05 | -3.119e-05 | -3.7663e-05 | -2.604e-05 | -3.9281e-05 | -3.472e-05 | -3.4426e-05 | -3.472e-05 | -3.1337e-05 | -4.7226e-05 | -1.8243e-05 | -3.6633e-05
| 100 | 5.0 | 10 | 0.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0169255E+00 | DC11V |  | 4.9999776807499998 | 7.1263711794039668e-06 | 1517347238.538 | 1517347253.683 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 100 | 10.0 | 10 | 5.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0472852E+00 | DC11V |  | 4.9999616191499996 | 1.2117009359915803e-05 | 1517347268.539 | 1517347283.648 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 4.9999 | 4.9999 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 4.9999 | 5 | 5 | 4.9999
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9347833E+00 | DC11V |  | 9.9999764082000002 | 1.010969031952397e-05 | 1517347298.49 | 1517347313.592 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
| 100 | 15.0 | 10 | 5.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 50.0269721E+00 | DC11V |  | 9.9999604612999988 | 9.0668097868523056e-06 | 1517347328.457 | 1517347343.614 | 9.9999 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 9.9999 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
| 100 | 15.0 | 10 | 0.0 | 100 | 15.0 | 20.0 | 10.0 | 0.0 | 49.9698390E+00 | DC11V |  | 14.999967610499999 | 8.681043989857386e-06 | 1517347358.451 | 1517347373.564 | 15 | 15 | 15 | 15 | 15 | 15 | 14.9999 | 15 | 15 | 14.9999 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15
| 100 | 20.0 | 10 | 5.0 | 100 | 15.0 | 20.0 | 10.0 | 0.0 | 50.0595851E+00 | DC11V |  | 14.999952027500001 | 6.7278405712276385e-06 | 1517347388.419 | 1517347403.547 | 14.9999 | 15 | 15 | 15 | 14.9999 | 15 | 15 | 14.9999 | 14.9999 | 14.9999 | 15 | 15 | 14.9999 | 15 | 15 | 15 | 15 | 14.9999 | 15 | 15
| 100 | 20.0 | 10 | 0.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9812617E+00 | DC11V |  | 19.999969643 | 7.9371739933560308e-06 | 1517347418.384 | 1517347433.471 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
| 100 | 25.0 | 10 | 5.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9626176E+00 | DC11V |  | 19.999947666499999 | 7.3609987603696512e-06 | 1517347448.322 | 1517347463.447 | 20 | 19.9999 | 20 | 19.9999 | 19.9999 | 20 | 20 | 20 | 19.9999 | 19.9999 | 20 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 20 | 19.9999 | 19.9999 | 20
| 100 | 25.0 | 10 | 0.0 | 100 | 25.0 | 20.0 | 10.0 | 0.0 | 50.0912937E+00 | DC11V |  | 24.999957129499997 | 8.2766040593294798e-06 | 1517347478.301 | 1517347493.433 | 24.9999 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 24.9999 | 24.9999 | 25 | 25 | 24.9999 | 25 | 25 | 25 | 25 | 25 | 25
| 100 | 30.0 | 10 | 5.0 | 100 | 25.0 | 20.0 | 10.0 | 0.0 | 49.9708519E+00 | DC11V |  | 24.999943979000001 | 1.0842101687025547e-05 | 1517347508.276 | 1517347523.36 | 24.9999 | 25 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 25 | 25 | 25 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 25 | 24.9999
| 100 | 30.0 | 10 | 0.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9985802E+00 | DC11V |  | 29.999954918999997 | 8.5821092393411973e-06 | 1517347538.194 | 1517347553.33 | 29.9999 | 30 | 30 | 30 | 30 | 30 | 30 | 29.9999 | 29.9999 | 30 | 30 | 30 | 30 | 29.9999 | 30 | 30 | 30 | 30 | 30 | 29.9999
| 100 | 35.0 | 10 | 5.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9845849E+00 | DC11V |  | 29.999937700000004 | 8.1488312047768901e-06 | 1517347568.182 | 1517347583.305 | 29.9999 | 30 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 30 | 29.9999 | 29.9999
| 100 | 35.0 | 10 | 0.0 | 100 | 35.0 | 20.0 | 10.0 | 0.0 | 49.9987029E+00 | DC11V |  | 34.999949118500005 | 1.0116845494427362e-05 | 1517347598.151 | 1517347613.274 | 35 | 35 | 35 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 35 | 34.9999 | 35 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 35 | 34.9999 | 34.9999 | 35 | 35
| 100 | 40.0 | 10 | 5.0 | 100 | 35.0 | 20.0 | 10.0 | 0.0 | 49.9811628E+00 | DC11V |  | 34.999931410000002 | 8.5317770729525072e-06 | 1517347628.134 | 1517347643.253 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999
| 100 | 40.0 | 10 | 0.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 50.0337775E+00 | DC11V |  | 39.999948184499999 | 8.2600705062332579e-06 | 1517347658.095 | 1517347673.214 | 39.9999 | 40 | 40 | 40 | 39.9999 | 40 | 39.9999 | 40 | 40 | 39.9999 | 40 | 39.9999 | 40 | 40 | 40 | 39.9999 | 40 | 39.9999 | 39.9999 | 39.9999
| 100 | 45.0 | 10 | 5.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9427825E+00 | DC11V |  | 39.999932121499995 | 7.7171265865545278e-06 | 1517347688.036 | 1517347703.12 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 40 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 45.0 | 10 | 0.0 | 100 | 45.0 | 20.0 | 10.0 | 0.0 | 49.9739711E+00 | DC11V |  | 44.999941652499999 | 1.1081701977894581e-05 | 1517347717.982 | 1517347733.11 | 45 | 44.9999 | 44.9999 | 44.9999 | 45 | 45 | 44.9999 | 45 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 45 | 45 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999
| 100 | 50.0 | 10 | 5.0 | 100 | 45.0 | 20.0 | 10.0 | 0.0 | 49.9611246E+00 | DC11V |  | 44.999909686999999 | 4.6598614797590763e-05 | 1517347747.95 | 1517347763.075 | 44.9997 | 44.9998 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999
| 100 | 50.0 | 10 | 0.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 49.9937979E+00 | DC11V |  | 49.999937294499993 | 1.1333869363861764e-05 | 1517347777.911 | 1517347793.059 | 49.9999 | 49.9999 | 49.9999 | 50 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 50 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 50.0 | 10 | 5.0 | 100 | 45.0 | 20.0 | 10.0 | 0.0 | 49.8667523E+00 | DC11V |  | 44.999922614999996 | 1.173301687506908e-05 | 1517347807.883 | 1517347823.005 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999
| 100 | 45.0 | 10 | 0.0 | 100 | 45.0 | 20.0 | 10.0 | 0.0 | 49.9966182E+00 | DC11V |  | 44.999935117999989 | 8.3138195795567648e-06 | 1517347837.853 | 1517347853.036 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999
| 100 | 45.0 | 10 | 5.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9543954E+00 | DC11V |  | 39.999923301999999 | 8.8652002795091186e-06 | 1517347867.862 | 1517347882.974 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 40.0 | 10 | 0.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 50.0153302E+00 | DC11V |  | 39.9999359595 | 7.6696574080225854e-06 | 1517347897.807 | 1517347912.955 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 40 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 40.0 | 10 | 5.0 | 100 | 35.0 | 20.0 | 10.0 | 0.0 | 50.0506550E+00 | DC11V |  | 34.999922031499992 | 9.69858148151664e-06 | 1517347927.796 | 1517347942.898 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999
| 100 | 35.0 | 10 | 0.0 | 100 | 35.0 | 20.0 | 10.0 | 0.0 | 49.9924351E+00 | DC11V |  | 34.999942032499995 | 8.5107842615467126e-06 | 1517347957.735 | 1517347972.819 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 35 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 35 | 34.9999 | 35 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999
| 100 | 35.0 | 10 | 5.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 50.0581808E+00 | DC11V |  | 29.999930402499995 | 7.8069148033255539e-06 | 1517347987.694 | 1517348002.797 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 30.0 | 10 | 0.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 50.0002453E+00 | DC11V |  | 29.999939945499996 | 9.2587118299881948e-06 | 1517348017.667 | 1517348032.755 | 30 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 30 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 30.0 | 10 | 5.0 | 100 | 25.0 | 20.0 | 10.0 | 0.0 | 50.0159343E+00 | DC11V |  | 24.999928117500001 | 6.594913096474005e-06 | 1517348047.584 | 1517348062.707 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999
| 100 | 25.0 | 10 | 0.0 | 100 | 25.0 | 20.0 | 10.0 | 0.0 | 49.9942224E+00 | DC11V |  | 24.999939248999997 | 8.3700327958461927e-06 | 1517348077.528 | 1517348092.638 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 25 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 25 | 24.9999 | 24.9999 | 24.9999
| 100 | 25.0 | 10 | 5.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9963493E+00 | DC11V |  | 19.999927646499998 | 8.7272265211362241e-06 | 1517348107.484 | 1517348122.596 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 20.0 | 10 | 0.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9747345E+00 | DC11V |  | 19.999946375499999 | 1.0937658101784424e-05 | 1517348137.426 | 1517348152.542 | 19.9999 | 19.9999 | 20 | 20 | 20 | 19.9999 | 20 | 19.9999 | 20 | 19.9999 | 19.9999 | 20 | 19.9999 | 19.9999 | 19.9999 | 20 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 20.0 | 10 | 5.0 | 100 | 15.0 | 20.0 | 10.0 | 0.0 | 50.0210277E+00 | DC11V |  | 14.999930642999999 | 8.648358861538573e-06 | 1517348167.372 | 1517348182.502 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999
| 100 | 15.0 | 10 | 0.0 | 100 | 15.0 | 20.0 | 10.0 | 0.0 | 50.0507732E+00 | DC11V |  | 14.999947914 | 8.7752882574674882e-06 | 1517348197.358 | 1517348212.462 | 14.9999 | 15 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 15 | 15 | 15 | 14.9999 | 14.9999 | 14.9999 | 15 | 14.9999 | 15 | 14.9999 | 14.9999 | 15 | 14.9999 | 14.9999
| 100 | 15.0 | 10 | 5.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9360065E+00 | DC11V |  | 9.9999328325499981 | 1.0603724046393098e-05 | 1517348227.296 | 1517348242.395 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9507326E+00 | DC11V |  | 9.9999490723000015 | 7.6626601130726398e-06 | 1517348257.253 | 1517348272.415 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 10 | 10 | 9.9999 | 9.9999 | 10 | 10 | 9.9999 | 9.9999 | 10 | 10 | 10 | 9.9999 | 10 | 10 | 9.9999
| 100 | 10.0 | 10 | 5.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0150328E+00 | DC11V |  | 4.9999351797499987 | 8.3924178272558436e-06 | 1517348287.253 | 1517348302.391 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999
| 100 | 5.0 | 10 | 0.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0431835E+00 | DC11V |  | 4.9999513155499997 | 8.4838663501067324e-06 | 1517348317.246 | 1517348332.362 | 4.9999 | 4.9999 | 4.9999 | 5 | 5 | 4.9999 | 5 | 5 | 5 | 4.9999 | 4.9999 | 5 | 4.9999 | 5 | 4.9999 | 5 | 4.9999 | 5 | 5 | 4.9999
| 100 | 5.0 | 10 | 5 | 100 | 0 | 20.0 | 10.0 | 0.0 | 49.9954815E+00 | DC11V |  | -6.3374074356000002e-05 | 8.0000642854400758e-06 | 1517348347.212 | 1517348362.354 | -5.774e-05 | -7.2764e-05 | -6.1128e-05 | -6.481e-05 | -6.0981e-05 | -6.4221e-05 | -6.4368e-05 | -7.8951e-05 | -6.3632e-05 | -4.3158e-05 | -6.5547e-05 | -6.157e-05 | -7.247e-05 | -5.5972e-05 | -6.3043e-05 | -7.2175e-05 | -7.1439e-05 | -6.319e-05 | -6.1128e-05 | -4.9197e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0022029E+00 | DC11V |  | -4.5001983773499992e-05 | 7.660907334306229e-06 | 1517348377.208 | 1517348392.377 | -4.9871e-05 | -5.046e-05 | -4.487e-05 | -5.0166e-05 | -4.7812e-05 | -4.84e-05 | -4.8106e-05 | -3.0452e-05 | -5.5756e-05 | -3.7955e-05 | -4.5458e-05 | -5.1931e-05 | -5.5168e-05 | -4.7076e-05 | -2.4715e-05 | -3.8397e-05 | -4.7665e-05 | -4.3104e-05 | -4.5752e-05 | -3.6925e-05
| 100 | 5.0 | 10 | 5 | 100 | 0 | 20.0 | 10.0 | 0.0 | 49.9619865E+00 | DC11V |  | -5.8489505081500009e-05 | 1.0493529786468223e-05 | 1517348407.211 | 1517348422.372 | -6.3427e-05 | -5.5186e-05 | -6.843e-05 | -7.0491e-05 | -6.6223e-05 | -6.1808e-05 | -5.0035e-05 | -6.5487e-05 | -6.0336e-05 | -5.2831e-05 | -3.6496e-05 | -4.4737e-05 | -4.3266e-05 | -5.6952e-05 | -7.1079e-05 | -5.1654e-05 | -5.6216e-05 | -5.9159e-05 | -5.3861e-05 | -8.2116e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0406840E+00 | DC11V |  | -4.8645839535500004e-05 | 8.581794292850337e-06 | 1517348437.214 | 1517348452.394 | -4.3094e-05 | -4.6329e-05 | -4.5447e-05 | -3.5593e-05 | -4.5594e-05 | -4.4564e-05 | -5.8831e-05 | -5.6919e-05 | -4.2505e-05 | -4.8536e-05 | -6.3537e-05 | -4.074e-05 | -6.339e-05 | -3.6769e-05 | -5.7066e-05 | -4.0299e-05 | -5.5595e-05 | -5.3095e-05 | -5.6625e-05 | -3.8387e-05
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 100 | 5.0 | 10 | 0 | 100 | 0 | 20.0 | 10.0 | 0.0 | 49.9844435E+00 | DC11V |  | 4.9999477634999989 | 1.2375429824810497e-05 | 1517348467.244 | 1517348482.354 | 4.9999 | 5 | 5 | 4.9999 | 4.9999 | 4.9999 | 5 | 4.9999 | 4.9999 | 5 | 5 | 4.9999 | 5 | 4.9999 | 4.9999 | 5 | 4.9999 | 4.9999 | 5 | 5
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9523756E+00 | DC11V |  | 9.9999464672000009 | 8.6576612696353324e-06 | 1517348497.197 | 1517348512.338 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 10 | 10 | 10 | 10 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 10
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0227792E+00 | DC11V |  | -4.8770367976000001e-05 | 9.2866988082571087e-06 | 1517348527.164 | 1517348542.344 | -3.8828e-05 | -4.0593e-05 | -5.4271e-05 | -5.9713e-05 | -4.7947e-05 | -4.2946e-05 | -3.4563e-05 | -6.2801e-05 | -5.1035e-05 | -3.0739e-05 | -4.7653e-05 | -4.3093e-05 | -5.383e-05 | -6.7214e-05 | -4.2946e-05 | -4.2064e-05 | -5.633e-05 | -5.8536e-05 | -5.2506e-05 | -4.78e-05
| 100 | 10.0 | 10 | 10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0186249E+00 | DC11V |  | -7.7944236623e-05 | 7.797796425876598e-06 | 1517348557.171 | 1517348572.317 | -7.9078e-05 | -8.7325e-05 | -8.1582e-05 | -6.9065e-05 | -7.9078e-05 | -6.5825e-05 | -8.4527e-05 | -7.3041e-05 | -8.3054e-05 | -8.8944e-05 | -8.7177e-05 | -6.2291e-05 | -6.8034e-05 | -7.2599e-05 | -7.9667e-05 | -8.3643e-05 | -8.5558e-05 | -8.4379e-05 | -6.9801e-05 | -7.4219e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9722748E+00 | DC11V |  | -5.0398015019500001e-05 | 1.0152106061542561e-05 | 1517348587.154 | 1517348602.307 | -4.7991e-05 | -5.7854e-05 | -3.6214e-05 | -5.3438e-05 | -4.9905e-05 | -3.7981e-05 | -3.7392e-05 | -5.4174e-05 | -3.7097e-05 | -4.5341e-05 | -6.4037e-05 | -3.9747e-05 | -6.1093e-05 | -5.4321e-05 | -5.5057e-05 | -6.1535e-05 | -6.9778e-05 | -4.6519e-05 | -3.7392e-05 | -6.1093e-05
| 100 | 10.0 | 10 | 10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0374620E+00 | DC11V |  | -7.5561768193499986e-05 | 8.7961521806184403e-06 | 1517348617.162 | 1517348632.336 | -7.5025e-05 | -7.4731e-05 | -7.1641e-05 | -7.5466e-05 | -7.8261e-05 | -7.6055e-05 | -6.7817e-05 | -8.3557e-05 | -8.3704e-05 | -9.3119e-05 | -6.7669e-05 | -6.8111e-05 | -7.0612e-05 | -5.2812e-05 | -6.8993e-05 | -7.9879e-05 | -7.2377e-05 | -7.6349e-05 | -8.5764e-05 | -8.9294e-05
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9621136E+00 | DC11V |  | 9.9999482232000023 | 9.8442195147888799e-06 | 1517348647.183 | 1517348662.288 | 9.9999 | 10 | 9.9999 | 9.9999 | 10 | 9.9999 | 10 | 9.9999 | 10 | 10 | 10 | 9.9999 | 9.9999 | 10 | 9.9999 | 10 | 10 | 9.9999 | 10 | 9.9999
| 100 | 20.0 | 10 | 10.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 50.0190214E+00 | DC11V |  | 9.9999111051499998 | 9.5779394510898239e-06 | 1517348677.131 | 1517348692.272 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 20.0 | 10 | 0.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9625375E+00 | DC11V |  | 19.999936392999999 | 8.6556571674428688e-06 | 1517348707.091 | 1517348722.181 | 19.9999 | 20 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 20 | 19.9999 | 19.9999
| 100 | 30.0 | 10 | 10.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9973916E+00 | DC11V |  | 19.999910339500001 | 9.1797753104445917e-06 | 1517348737.043 | 1517348752.176 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 30.0 | 10 | 0.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9957079E+00 | DC11V |  | 29.999929641000001 | 6.7468747580167499e-06 | 1517348767.053 | 1517348782.177 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 40.0 | 10 | 10.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 50.0026747E+00 | DC11V |  | 29.999012282000002 | 0.0012352475929608309 | 1517348797.021 | 1517348812.131 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9992 | 29.9989 | 29.9986 | 29.999 | 29.9954 | 29.9976 | 29.9975 | 29.9974 | 29.9977
| 100 | 40.0 | 10 | 0.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 50.0237188E+00 | DC11V |  | 39.998206078500004 | 0.0012520037359778752 | 1517348826.988 | 1517348842.101 | 39.9992 | 39.9983 | 39.9979 | 39.9977 | 39.9974 | 39.9977 | 39.9987 | 39.9985 | 39.998 | 39.9982 | 39.9981 | 39.998 | 39.9959 | 39.9952 | 39.9967 | 39.9988 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 50.0 | 10 | 10.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9830341E+00 | DC11V |  | 39.999895685999995 | 8.45722850525924e-06 | 1517348856.944 | 1517348872.05 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 50.0 | 10 | 0.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 49.9981840E+00 | DC11V |  | 49.999921264499996 | 8.4715432332832645e-06 | 1517348886.882 | 1517348902.003 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 60.0 | 10 | 10.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 49.9627448E+00 | DC11V |  | 49.9998867125 | 7.728736556131287e-06 | 1517348916.834 | 1517348931.945 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 60.0 | 10 | 0.0 | 100 | 60.0 | 20.0 | 10.0 | 0.0 | 50.0877339E+00 | DC11V |  | 59.999913534999997 | 8.8702060291779048e-06 | 1517348946.807 | 1517348961.941 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999
| 100 | 70.0 | 10 | 10.0 | 100 | 60.0 | 20.0 | 10.0 | 0.0 | 50.0113752E+00 | DC11V |  | 59.999878976500007 | 8.2929097880700261e-06 | 1517348976.777 | 1517348991.862 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999
| 100 | 70.0 | 10 | 0.0 | 100 | 70.0 | 20.0 | 10.0 | 0.0 | 49.9648266E+00 | DC11V |  | 69.999899496999987 | 8.2439232762636535e-06 | 1517349006.704 | 1517349021.82 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999
| 100 | 80.0 | 10 | 10.0 | 100 | 70.0 | 20.0 | 10.0 | 0.0 | 50.0015755E+00 | DC11V |  | 69.999864513999995 | 9.2543489235968399e-06 | 1517349036.654 | 1517349051.775 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999
| 100 | 80.0 | 10 | 0.0 | 100 | 80.0 | 20.0 | 10.0 | 0.0 | 50.0328470E+00 | DC11V |  | 79.999883340499991 | 9.2283386777817138e-06 | 1517349066.61 | 1517349081.718 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999
| 100 | 90.0 | 10 | 10.0 | 100 | 80.0 | 20.0 | 10.0 | 0.0 | 49.9972407E+00 | DC11V |  | 79.999854066500006 | 8.2004501556544865e-06 | 1517349096.55 | 1517349111.656 | 79.9998 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9998 | 79.9998 | 79.9999 | 79.9998 | 79.9999 | 79.9998 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9998 | 79.9999 | 79.9999 | 79.9999
| 100 | 90.0 | 10 | 0.0 | 100 | 90.0 | 20.0 | 10.0 | 0.0 | 50.0219530E+00 | DC11V |  | 89.999874187999993 | 8.1889911459299594e-06 | 1517349126.482 | 1517349141.605 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999
| 100 | 100.0 | 10 | 10.0 | 100 | 90.0 | 20.0 | 10.0 | 0.0 | 49.9654531E+00 | DC11V |  | 89.999842440000009 | 8.621804335117195e-06 | 1517349156.443 | 1517349171.552 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9999 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9999 | 89.9998 | 89.9998
| 100 | 100.0 | 10 | 0.0 | 100 | 100.0 | 20.0 | 10.0 | 0.0 | 50.0133621E+00 | DC11V |  | 99.999859686000008 | 9.5045617477887993e-06 | 1517349186.39 | 1517349201.532 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9998 | 99.9999 | 99.9999 | 99.9999 | 99.9998 | 99.9999 | 99.9998 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999
| 100 | 100.0 | 10 | 10.0 | 100 | 90.0 | 20.0 | 10.0 | 0.0 | 50.0012076E+00 | DC11V |  | 89.999841089 | 1.0045175409710058e-05 | 1517349216.358 | 1517349231.462 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9998
| 100 | 90.0 | 10 | 0.0 | 100 | 90.0 | 20.0 | 10.0 | 0.0 | 50.0253855E+00 | DC11V |  | 89.999875953 | 7.8219333299982297e-06 | 1517349246.282 | 1517349261.406 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999
| 100 | 90.0 | 10 | 10.0 | 100 | 80.0 | 20.0 | 10.0 | 0.0 | 49.9690899E+00 | DC11V |  | 79.999856968999993 | 8.8813320509792709e-06 | 1517349276.228 | 1517349291.322 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9998 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9998 | 79.9999 | 79.9999 | 79.9999 | 79.9998 | 79.9999
| 100 | 80.0 | 10 | 0.0 | 100 | 80.0 | 20.0 | 10.0 | 0.0 | 49.9661550E+00 | DC11V |  | 79.999882414499993 | 6.8515403187283919e-06 | 1517349306.148 | 1517349321.284 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999
| 100 | 80.0 | 10 | 10.0 | 100 | 70.0 | 20.0 | 10.0 | 0.0 | 50.0396919E+00 | DC11V |  | 69.999867130000013 | 8.4285319009003062e-06 | 1517349336.145 | 1517349351.277 | 69.9999 | 69.9999 | 69.9998 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999
| 100 | 70.0 | 10 | 0.0 | 100 | 70.0 | 20.0 | 10.0 | 0.0 | 49.9934537E+00 | DC11V |  | 69.999897540000006 | 8.8043591466009637e-06 | 1517349366.131 | 1517349381.234 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999
| 100 | 70.0 | 10 | 10.0 | 100 | 60.0 | 20.0 | 10.0 | 0.0 | 50.0003962E+00 | DC11V |  | 59.999871251499997 | 7.9380912541382647e-06 | 1517349396.092 | 1517349411.209 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999
| 100 | 60.0 | 10 | 0.0 | 100 | 60.0 | 20.0 | 10.0 | 0.0 | 49.9964201E+00 | DC11V |  | 59.999908356500001 | 9.1197611126410454e-06 | 1517349426.054 | 1517349441.163 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999
| 100 | 60.0 | 10 | 10.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 50.0459667E+00 | DC11V |  | 49.999886320000002 | 9.7386580184260267e-06 | 1517349456.003 | 1517349471.114 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 50.0 | 10 | 0.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 49.9709179E+00 | DC11V |  | 49.999916320499992 | 1.0007244613292332e-05 | 1517349485.969 | 1517349501.06 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 50.0 | 10 | 10.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9876964E+00 | DC11V |  | 39.999894089000001 | 7.5318542868307727e-06 | 1517349515.883 | 1517349531.006 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 40.0 | 10 | 0.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 50.0096481E+00 | DC11V |  | 39.999921989000001 | 7.5124782187648827e-06 | 1517349545.864 | 1517349560.989 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 40.0 | 10 | 10.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 50.0500737E+00 | DC11V |  | 29.999906384500001 | 7.716568197732922e-06 | 1517349575.816 | 1517349590.922 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 30.0 | 10 | 0.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 50.0564410E+00 | DC11V |  | 29.999927460000002 | 1.1597066871148246e-05 | 1517349605.763 | 1517349620.871 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 30.0 | 10 | 10.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9571029E+00 | DC11V |  | 19.999909648999999 | 8.5399091916583438e-06 | 1517349635.696 | 1517349650.781 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 20.0 | 10 | 0.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9631734E+00 | DC11V |  | 19.999940588999998 | 1.0281973983778019e-05 | 1517349665.623 | 1517349680.758 | 19.9999 | 19.9999 | 19.9999 | 20 | 19.9999 | 19.9999 | 19.9999 | 20 | 19.9999 | 20 | 19.9999 | 19.9999 | 20 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 20.0 | 10 | 10.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9225964E+00 | DC11V |  | 9.9999227755000017 | 6.3787735538145009e-06 | 1517349695.612 | 1517349710.733 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9956608E+00 | DC11V |  | 9.9999500048000005 | 8.2504545122726491e-06 | 1517349725.573 | 1517349740.718 | 10 | 10 | 9.9999 | 9.9999 | 9.9999 | 10 | 9.9999 | 9.9999 | 10 | 10 | 10 | 9.9999 | 9.9999 | 10 | 10 | 9.9999 | 9.9999 | 9.9999 | 10 | 10
| 100 | 10.0 | 10 | 10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9649397E+00 | DC11V |  | -7.3656762237499994e-05 | 8.9746455913005327e-06 | 1517349755.547 | 1517349770.705 | -7.1537e-05 | -7.6395e-05 | -7.5953e-05 | -8.5226e-05 | -6.4472e-05 | -8.5521e-05 | -8.8612e-05 | -7.8603e-05 | -7.139e-05 | -6.5944e-05 | -6.668e-05 | -6.4913e-05 | -8.1399e-05 | -7.8455e-05 | -6.4913e-05 | -5.7554e-05 | -6.6532e-05 | -8.7876e-05 | -7.7867e-05 | -6.3294e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9634042E+00 | DC11V |  | -6.9480325975500003e-05 | 2.9328814114297394e-05 | 1517349785.54 | 1517349800.712 | -7.5166e-05 | -5.163e-05 | -6.1339e-05 | -6.8693e-05 | -6.6193e-05 | -0.0001165 | -0.00013592 | -0.000109 | -0.00011179 | -0.00010929 | -4.9571e-05 | -4.4129e-05 | -4.6041e-05 | -5.0895e-05 | -5.9573e-05 | -3.2508e-05 | -4.8983e-05 | -6.4281e-05 | -3.7951e-05 | -5.0159e-05
| 100 | 10.0 | 10 | 10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0039154E+00 | DC11V |  | -7.1604137767000012e-05 | 8.3020327689190288e-06 | 1517349815.546 | 1517349830.708 | -9.2862e-05 | -6.9169e-05 | -8.6829e-05 | -6.858e-05 | -7.0935e-05 | -5.8425e-05 | -6.7255e-05 | -7.1376e-05 | -7.8146e-05 | -7.8293e-05 | -7.2995e-05 | -5.9161e-05 | -7.4761e-05 | -6.1222e-05 | -7.8146e-05 | -6.8433e-05 | -6.5195e-05 | -7.2848e-05 | -7.1229e-05 | -6.6225e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9894079E+00 | DC11V |  | -5.1767897453999998e-05 | 8.8566307776006075e-06 | 1517349845.541 | 1517349860.686 | -4.4619e-05 | -3.9612e-05 | -3.7698e-05 | -5.2129e-05 | -7.4512e-05 | -5.8755e-05 | -4.8742e-05 | -5.3454e-05 | -4.241e-05 | -4.5502e-05 | -5.5221e-05 | -6.3467e-05 | -5.154e-05 | -4.3882e-05 | -5.0067e-05 | -4.4177e-05 | -5.9639e-05 | -5.9786e-05 | -5.9933e-05 | -5.0214e-05
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9862915E+00 | DC11V |  | 9.9999461400499996 | 6.2863760664017882e-06 | 1517349875.517 | 1517349890.637 | 9.9999 | 9.9999 | 9.9999 | 10 | 9.9999 | 9.9999 | 10 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 10 | 10 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | -10.0 | 10 | 0.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 49.9902472E+00 | DC11V |  | -10.000046057500001 | 9.0465456804414806e-06 | 1517349905.482 | 1517349920.585 | -10.0001 | -10 | -10.0001 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10.0001 | -10.0001 | -10 | -10 | -10 | -10.0001 | -10.0001 | -10 | -10 | -10
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9798855E+00 | DC11V |  | -4.8043584632999998e-05 | 9.0096574970128908e-06 | 1517349935.424 | 1517349950.577 | -5.7855e-05 | -3.9012e-05 | -5.6972e-05 | -4.387e-05 | -3.6657e-05 | -4.6078e-05 | -5.7708e-05 | -3.5479e-05 | -3.1946e-05 | -5.0495e-05 | -5.0347e-05 | -4.4312e-05 | -6.0211e-05 | -5.1378e-05 | -5.6089e-05 | -5.8886e-05 | -3.3271e-05 | -4.8728e-05 | -4.3576e-05 | -5.8003e-05
| 100 | -10.0 | 10 | -10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0289554E+00 | DC11V |  | -2.5286183334150004e-05 | 8.2039585640591955e-06 | 1517349965.432 | 1517349980.583 | -2.6354e-05 | -3.0034e-05 | -2.5029e-06 | -2.7679e-05 | -2.1348e-05 | -2.3262e-05 | -2.0612e-05 | -3.4598e-05 | -2.9298e-05 | -2.3703e-05 | -3.6071e-05 | -3.4009e-05 | -1.4134e-05 | -1.6048e-05 | -2.6795e-05 | -3.0918e-05 | -3.5776e-05 | -2.9151e-05 | -1.6048e-05 | -2.7384e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9997075E+00 | DC11V |  | -5.2595433298499992e-05 | 6.6648989245087765e-06 | 1517349995.419 | 1517350010.557 | -4.8611e-05 | -5.686e-05 | -5.1852e-05 | -4.5075e-05 | -4.5665e-05 | -4.4192e-05 | -5.082e-05 | -3.6679e-05 | -5.2441e-05 | -5.7449e-05 | -5.5534e-05 | -4.8464e-05 | -5.3914e-05 | -5.8038e-05 | -5.5534e-05 | -5.1557e-05 | -5.6565e-05 | -5.5092e-05 | -6.9234e-05 | -5.8333e-05
| 100 | -10.0 | 10 | -10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9755450E+00 | DC11V |  | -3.0041754950500002e-05 | 9.5314668935711046e-06 | 1517350025.413 | 1517350040.567 | -2.2373e-05 | -1.9282e-05 | -3.8123e-05 | -1.8546e-05 | -3.2382e-05 | -1.413e-05 | -4.3421e-05 | -2.3109e-05 | -4.0183e-05 | -4.3274e-05 | -2.046e-05 | -1.6927e-05 | -3.2824e-05 | -2.4287e-05 | -4.3863e-05 | -4.1508e-05 | -3.3854e-05 | -3.4884e-05 | -3.1205e-05 | -2.62e-05
| 100 | -10.0 | 10 | 0.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 50.0615663E+00 | DC11V |  | -10.000051256499999 | 1.1104743254484706e-05 | 1517350055.407 | 1517350070.5 | -10.0001 | -10 | -10 | -10 | -10 | -10.0001 | -10.0001 | -10 | -10.0001 | -10 | -10.0001 | -10 | -10 | -10.0001 | -10.0001 | -10 | -10 | -10 | -10 | -10.0001
| 100 | -20.0 | 10 | -10.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 50.0510710E+00 | DC11V |  | -10.000015377999999 | 9.3173019698003059e-06 | 1517350085.333 | 1517350101.247 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10
| 100 | -20.0 | 10 | 0.0 | 100 | -20.0 | 20.0 | 10.0 | 0.0 | 50.0135886E+00 | DC11V |  | -20.000041367999998 | 9.0280953696248414e-06 | 1517350116.1 | 1517350132.015 | -20 | -20 | -20 | -20 | -20 | -20.0001 | -20 | -20 | -20 | -20 | -20 | -20.0001 | -20 | -20 | -20 | -20.0001 | -20 | -20.0001 | -20 | -20
| 100 | -30.0 | 10 | -10.0 | 100 | -20.0 | 20.0 | 10.0 | 0.0 | 49.9789381E+00 | DC11V |  | -20.000000209 | 9.4334659060474908e-06 | 1517350146.869 | 1517350161.954 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20
| 100 | -30.0 | 10 | 0.0 | 100 | -30.0 | 20.0 | 10.0 | 0.0 | 49.9971700E+00 | DC11V |  | -30.000025799500001 | 1.0156219510514767e-05 | 1517350176.804 | 1517350191.918 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30
| 100 | -40.0 | 10 | -10.0 | 100 | -30.0 | 20.0 | 10.0 | 0.0 | 50.0199938E+00 | DC11V |  | -29.999988487999996 | 8.6494147777250054e-06 | 1517350206.796 | 1517350221.914 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30
| 100 | -40.0 | 10 | 0.0 | 100 | -40.0 | 20.0 | 10.0 | 0.0 | 50.0011227E+00 | DC11V |  | -40.000014154999995 | 8.994991662203165e-06 | 1517350236.736 | 1517350251.834 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40
| 100 | -50.0 | 10 | -10.0 | 100 | -40.0 | 20.0 | 10.0 | 0.0 | 50.0055572E+00 | DC11V |  | -39.999981740499997 | 8.4861925948700648e-06 | 1517350266.673 | 1517350281.793 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40
| 100 | -50.0 | 10 | 0.0 | 100 | -50.0 | 20.0 | 10.0 | 0.0 | 50.0093838E+00 | DC11V |  | -50.000010075000006 | 9.0863680861153096e-06 | 1517350296.616 | 1517350311.721 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50
| 100 | -60.0 | 10 | -10.0 | 100 | -50.0 | 20.0 | 10.0 | 0.0 | 50.0051043E+00 | DC11V |  | -49.999969113999995 | 6.7775418843145023e-06 | 1517350326.545 | 1517350341.647 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50
| 100 | -60.0 | 10 | 0.0 | 100 | -60.0 | 20.0 | 10.0 | 0.0 | 50.0343396E+00 | DC11V |  | -60.000000918499993 | 7.4153160912712183e-06 | 1517350356.474 | 1517350371.567 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60
| 100 | -70.0 | 10 | -10.0 | 100 | -60.0 | 20.0 | 10.0 | 0.0 | 50.0140322E+00 | DC11V |  | -59.999959199999999 | 8.8523556184490637e-06 | 1517350386.436 | 1517350401.533 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -59.9999 | -59.9999 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -59.9999 | -60
| 100 | -70.0 | 10 | 0.0 | 100 | -70.0 | 20.0 | 10.0 | 0.0 | 49.9985944E+00 | DC11V |  | -69.999986299 | 8.1298289642179256e-06 | 1517350416.374 | 1517350431.492 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70
| 100 | -80.0 | 10 | -10.0 | 100 | -70.0 | 20.0 | 10.0 | 0.0 | 50.0034200E+00 | DC11V |  | -69.999951492999998 | 8.0592599519985461e-06 | 1517350446.343 | 1517350461.466 | -70 | -70 | -69.9999 | -70 | -70 | -69.9999 | -69.9999 | -69.9999 | -70 | -70 | -70 | -70 | -69.9999 | -70 | -69.9999 | -69.9999 | -70 | -70 | -70 | -69.9999
| 100 | -80.0 | 10 | 0.0 | 100 | -80.0 | 20.0 | 10.0 | 0.0 | 50.0338436E+00 | DC11V |  | -79.999975826500005 | 8.891038902658884e-06 | 1517350476.309 | 1517350491.413 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80
| 100 | -90.0 | 10 | -10.0 | 100 | -80.0 | 20.0 | 10.0 | 0.0 | 49.9905395E+00 | DC11V |  | -79.99994141900001 | 7.4214809170523535e-06 | 1517350506.235 | 1517350521.339 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -80 | -79.9999 | -79.9999
| 100 | -90.0 | 10 | 0.0 | 100 | -90.0 | 20.0 | 10.0 | 0.0 | 49.9647372E+00 | DC11V |  | -89.999968022999994 | 8.8752414612627044e-06 | 1517350536.176 | 1517350551.294 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -89.9999 | -90 | -90 | -90 | -90
| 100 | -100.0 | 10 | -10.0 | 100 | -90.0 | 20.0 | 10.0 | 0.0 | 49.9586286E+00 | DC11V |  | -89.999928179000008 | 8.7384345844584618e-06 | 1517350566.129 | 1517350581.269 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -90 | -89.9999
| 100 | -100.0 | 10 | 0.0 | 100 | -100.0 | 20.0 | 10.0 | 0.0 | 50.0181339E+00 | DC11V |  | -99.999953720000008 | 9.9806733228616261e-06 | 1517350596.117 | 1517350611.245 | -99.9999 | -99.9999 | -100 | -100 | -100 | -100 | -99.9999 | -100 | -100 | -100 | -100 | -99.9999 | -100 | -99.9999 | -100 | -100 | -99.9999 | -100 | -100 | -99.9999
| 100 | -100.0 | 10 | -10.0 | 100 | -90.0 | 20.0 | 10.0 | 0.0 | 50.0250644E+00 | DC11V |  | -89.999928045499985 | 8.1219692666249803e-06 | 1517350626.076 | 1517350641.191 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999 | -89.9999
| 100 | -90.0 | 10 | 0.0 | 100 | -90.0 | 20.0 | 10.0 | 0.0 | 49.9930387E+00 | DC11V |  | -89.999966757999999 | 8.1951977387128908e-06 | 1517350656.026 | 1517350671.131 | -90 | -90 | -90 | -90 | -90 | -90 | -89.9999 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90
| 100 | -90.0 | 10 | -10.0 | 100 | -80.0 | 20.0 | 10.0 | 0.0 | 49.9911242E+00 | DC11V |  | -79.999938320500021 | 8.5219073409190803e-06 | 1517350686.074 | 1517350701.188 | -79.9999 | -79.9999 | -79.9999 | -80 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999 | -79.9999
| 100 | -80.0 | 10 | 0.0 | 100 | -80.0 | 20.0 | 10.0 | 0.0 | 50.0174825E+00 | DC11V |  | -79.9999792765 | 9.2022172747159338e-06 | 1517350716.009 | 1517350731.124 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80
| 100 | -80.0 | 10 | -10.0 | 100 | -70.0 | 20.0 | 10.0 | 0.0 | 49.9624198E+00 | DC11V |  | -69.999947742499998 | 9.1265562379509346e-06 | 1517350745.943 | 1517350761.059 | -69.9999 | -69.9999 | -69.9999 | -69.9999 | -69.9999 | -69.9999 | -69.9999 | -70 | -70 | -69.9999 | -69.9999 | -70 | -70 | -70 | -69.9999 | -70 | -69.9999 | -69.9999 | -69.9999 | -69.9999
| 100 | -70.0 | 10 | 0.0 | 100 | -70.0 | 20.0 | 10.0 | 0.0 | 50.0363330E+00 | DC11V |  | -69.999987272499979 | 9.7035199164486947e-06 | 1517350775.908 | 1517350791.028 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70
| 100 | -70.0 | 10 | -10.0 | 100 | -60.0 | 20.0 | 10.0 | 0.0 | 49.9723973E+00 | DC11V |  | -59.999959953999998 | 8.6528691205211891e-06 | 1517350806.002 | 1517350821.111 | -59.9999 | -60 | -59.9999 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -59.9999 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60
| 100 | -60.0 | 10 | 0.0 | 100 | -60.0 | 20.0 | 10.0 | 0.0 | 49.9902896E+00 | DC11V |  | -60.000001660500004 | 8.1136264855614735e-06 | 1517350835.975 | 1517350851.099 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60
| 100 | -60.0 | 10 | -10.0 | 100 | -50.0 | 20.0 | 10.0 | 0.0 | 50.0044532E+00 | DC11V |  | -49.999970945000008 | 8.7808937480637425e-06 | 1517350865.923 | 1517350881.035 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50
| 100 | -50.0 | 10 | 0.0 | 100 | -50.0 | 20.0 | 10.0 | 0.0 | 50.0086996E+00 | DC11V |  | -50.000007991000004 | 1.1885953853358934e-05 | 1517350895.876 | 1517350910.967 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50
| 100 | -50.0 | 10 | -10.0 | 100 | -40.0 | 20.0 | 10.0 | 0.0 | 49.9865131E+00 | DC11V |  | -39.999983331499998 | 7.3206477000766206e-06 | 1517350925.826 | 1517350940.941 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40
| 100 | -40.0 | 10 | 0.0 | 100 | -40.0 | 20.0 | 10.0 | 0.0 | 50.0187004E+00 | DC11V |  | -40.000014951500006 | 6.9632092282337367e-06 | 1517350955.789 | 1517350970.894 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40
| 100 | -40.0 | 10 | -10.0 | 100 | -30.0 | 20.0 | 10.0 | 0.0 | 50.0410951E+00 | DC11V |  | -29.999993089499998 | 9.3988033680216705e-06 | 1517350985.75 | 1517351000.85 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30
| 100 | -30.0 | 10 | 0.0 | 100 | -30.0 | 20.0 | 10.0 | 0.0 | 49.9907564E+00 | DC11V |  | -30.000023611500001 | 1.2234550778581589e-05 | 1517351015.698 | 1517351030.781 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30.0001 | -30 | -30 | -30 | -30
| 100 | -30.0 | 10 | -10.0 | 100 | -20.0 | 20.0 | 10.0 | 0.0 | 50.0173173E+00 | DC11V |  | -20.000000401499999 | 8.5965593550839325e-06 | 1517351045.614 | 1517351060.72 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20
| 100 | -20.0 | 10 | 0.0 | 100 | -20.0 | 20.0 | 10.0 | 0.0 | 50.0157314E+00 | DC11V |  | -20.000034658000004 | 9.8426071750190268e-06 | 1517351075.541 | 1517351090.628 | -20 | -20 | -20 | -20.0001 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20
| 100 | -20.0 | 10 | -10.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 50.0159626E+00 | DC11V |  | -10.000008586 | 8.9958959029194605e-06 | 1517351105.472 | 1517351120.577 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10
| 100 | -10.0 | 10 | 0.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 50.0331635E+00 | DC11V |  | -10.0000451455 | 6.4565311699567981e-06 | 1517351135.397 | 1517351150.482 | -10 | -10 | -10.0001 | -10 | -10 | -10.0001 | -10 | -10 | -10 | -10 | -10 | -10.0001 | -10 | -10 | -10 | -10 | -10 | -10 | -10.0001 | -10
| 100 | -10.0 | 10 | -10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0041277E+00 | DC11V |  | -1.9721247693859997e-05 | 8.6995632350424993e-06 | 1517351165.322 | 1517351180.465 | -2.1356e-05 | -1.841e-05 | -1.9441e-05 | -3.2108e-05 | -2.6658e-05 | -2.5333e-05 | -1.0457e-05 | -2.5627e-05 | -1.2666e-05 | -2.4596e-05 | -1.6054e-05 | -3.2108e-05 | -1.9736e-05 | -8.6897e-06 | -1.4286e-05 | -1.679e-05 | -1.8263e-05 | -1.2814e-05 | -7.3642e-07 | -3.8294e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9574560E+00 | DC11V |  | -4.7844661540500004e-05 | 8.0422236290709719e-06 | 1517351195.317 | 1517351210.466 | -4.1079e-05 | -5.3005e-05 | -4.5202e-05 | -5.065e-05 | -4.5791e-05 | -6.287e-05 | -3.4453e-05 | -4.8441e-05 | -3.8723e-05 | -5.2564e-05 | -4.2993e-05 | -5.065e-05 | -5.1386e-05 | -4.3582e-05 | -3.1214e-05 | -5.8748e-05 | -5.7422e-05 | -3.9901e-05 | -5.4478e-05 | -5.3742e-05
| 100 | -10.0 | 10 | -10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0233789E+00 | DC11V |  | -2.580279056085e-05 | 9.8733422158045597e-06 | 1517351225.303 | 1517351240.485 | -2.9119e-05 | -3.4119e-05 | -3.206e-05 | -1.4118e-05 | -6.3239e-06 | -3.4855e-05 | -3.7649e-05 | -2.2795e-05 | -2.7943e-06 | -2.6766e-05 | -2.2648e-05 | -3.7208e-05 | -2.4119e-05 | -3.3678e-05 | -3.2208e-05 | -3.6472e-05 | -1.6618e-05 | -2.0295e-05 | -3.2355e-05 | -1.9854e-05
| 100 | 0 | 10 | 0.0 | 100 | 0 | 20.0 | 10.0 | 0.0 | 50.0074823E+00 | DC11V |  | -5.0693977991999997e-05 | 1.0641327068481506e-05 | 1517351255.326 | 1517351270.486 | -4.7243e-05 | -6.3138e-05 | -6.3138e-05 | -4.6507e-05 | -7.7266e-05 | -3.6646e-05 | -4.7831e-05 | -4.636e-05 | -5.4454e-05 | -5.1658e-05 | -5.2688e-05 | -4.6213e-05 | -5.7986e-05 | -4.1503e-05 | -5.416e-05 | -6.4315e-05 | -4.5918e-05 | -3.1348e-05 | -3.5763e-05 | -4.9745e-05