## Sheet
instruments: | Meter: HP | S: FL5700 | X: FL5730
Start Row | 9 | Stop Row | 961
| X Range | X Settings (V) | S Range | S Settings (V) | DVM Range | Nominal reading | #Readings | Delay (S) | DVM pause | DVM status | S status | X status | Mean | STD | start time | end time | readings...
| 0.2 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0313309E+00 | DC220MV |  | 0.010000378488500002 | 4.853672622614333e-08 | 1517776045.339 | 1517776060.826 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
| 0.1 | 0.0001 | 0.2 | 0.0 | 0.1 | 0.0001 | 20.0 | 10.0 | 0.0 | 49.9662681E+00 | DC220MV |  | 0.00010035457753499999 | 6.774701217413328e-08 | 1517776075.972 | 1517776091.479 | 0.00010059 | 0.00010036 | 0.00010034 | 0.00010042 | 0.00010037 | 0.00010039 | 0.0001003 | 0.00010033 | 0.00010032 | 0.0001004 | 0.00010031 | 0.00010034 | 0.0001004 | 0.00010037 | 0.00010031 | 0.00010028 | 0.00010035 | 0.00010027 | 0.00010031 | 0.00010033
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0632639E+00 | DC220MV |  | 3.0938836745499995e-07 | 5.2078868661881508e-08 | 1517776106.327 | 1517776121.894 | 3.2904e-07 | 2.767e-07 | 3.3698e-07 | 2.9963e-07 | 3.9579e-07 | 3.1772e-07 | 2.6229e-07 | 1.9025e-07 | 2.7979e-07 | 2.3891e-07 | 3.1022e-07 | 3.0316e-07 | 3.3918e-07 | 2.1789e-07 | 3.7447e-07 | 3.5021e-07 | 3.361e-07 | 2.9993e-07 | 3.6653e-07 | 3.63e-07
| 0.1 | 0.0001 | 0.2 | 0.0001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0323606E+00 | DC220MV |  | 2.5499539999999998e-07 | 2.7878532965403897e-08 | 1517776136.744 | 1517776152.278 | 2.6308e-07 | 2.7088e-07 | 2.0139e-07 | 3.1828e-07 | 2.3996e-07 | 2.8619e-07 | 2.248e-07 | 2.538e-07 | 2.4247e-07 | 2.7486e-07 | 2.7721e-07 | 2.3452e-07 | 2.7014e-07 | 2.3555e-07 | 2.9473e-07 | 2.329e-07 | 2.5925e-07 | 2.6249e-07 | 2.1082e-07 | 2.4659e-07
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0005236E+00 | DC220MV |  | 2.4467315006500001e-07 | 3.5690713791262673e-08 | 1517776167.137 | 1517776182.678 | 2.6991e-07 | 2.3386e-07 | 2.4637e-07 | 2.2782e-07 | 2.7595e-07 | 2.1575e-07 | 2.3238e-07 | 2.3386e-07 | 2.4975e-07 | 1.7337e-07 | 2.4239e-07 | 2.8802e-07 | 2.5917e-07 | 2.3651e-07 | 2.5829e-07 | 2.3415e-07 | 2.1104e-07 | 3.5454e-07 | 2.0486e-07 | 2.4548e-07
| 0.1 | 0.0001 | 0.2 | 0.0001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0040711E+00 | DC220MV |  | 2.1615544595999998e-07 | 3.6989396154683875e-08 | 1517776197.534 | 1517776213.089 | 2.327e-07 | 2.3006e-07 | 1.5548e-07 | 2.0034e-07 | 2.7463e-07 | 2.3138e-07 | 2.2005e-07 | 1.8166e-07 | 2.7521e-07 | 1.5533e-07 | 2.1608e-07 | 2.1388e-07 | 1.7769e-07 | 2.1211e-07 | 2.1152e-07 | 2.7757e-07 | 2.6771e-07 | 1.6239e-07 | 2.1947e-07 | 2.0784e-07
| 0.1 | 0.0001 | 0.2 | 0.0 | 0.1 | 0.0001 | 20.0 | 10.0 | 0.0 | 50.0293002E+00 | DC220MV |  | 0.00010017927469500001 | 4.6533719330281755e-08 | 1517776227.969 | 1517776243.493 | 0.00010014 | 0.00010026 | 0.00010012 | 0.00010017 | 0.00010023 | 0.00010008 | 0.00010021 | 0.0001002 | 0.00010016 | 0.00010025 | 0.00010019 | 0.0001002 | 0.00010021 | 0.00010015 | 0.00010023 | 0.0001001 | 0.00010019 | 0.00010016 | 0.0001002 | 0.00010016
| 0.1 | 0.0002 | 0.2 | 0.0001 | 0.1 | 0.0001 | 20.0 | 10.0 | 0.0 | 49.9928642E+00 | DC220MV |  | 0.00010016549620500001 | 4.8622582411668333e-08 | 1517776258.344 | 1517776273.857 | 0.00010019 | 0.00010019 | 0.0001003 | 0.00010007 | 0.00010013 | 0.00010021 | 0.00010019 | 0.00010016 | 0.00010017 | 0.00010014 | 0.00010013 | 0.00010015 | 0.00010011 | 0.00010021 | 0.00010017 | 0.00010023 | 0.00010013 | 0.00010015 | 0.00010012 | 0.00010015
| 0.1 | 0.0002 | 0.2 | 0.0 | 0.1 | 0.0002 | 20.0 | 10.0 | 0.0 | 50.0129609E+00 | DC220MV |  | 0.00020012367580500004 | 3.8651109272070534e-08 | 1517776288.712 | 1517776304.23 | 0.0002001 | 0.0002001 | 0.00020013 | 0.0002001 | 0.00020008 | 0.00020022 | 0.00020018 | 0.00020015 | 0.00020015 | 0.00020008 | 0.00020014 | 0.00020011 | 0.00020005 | 0.0002001 | 0.0002001 | 0.00020013 | 0.00020016 | 0.00020017 | 0.00020013 | 0.0002001
| 0.1 | 0.0003 | 0.2 | 0.0001 | 0.1 | 0.0002 | 20.0 | 10.0 | 0.0 | 49.9775243E+00 | DC220MV |  | 0.00020010332238000001 | 2.9017655779759038e-08 | 1517776319.087 | 1517776334.61 | 0.00020016 | 0.00020007 | 0.00020012 | 0.00020011 | 0.00020006 | 0.00020012 | 0.0002001 | 0.00020011 | 0.00020012 | 0.00020013 | 0.00020009 | 0.00020012 | 0.00020011 | 0.00020006 | 0.00020013 | 0.00020005 | 0.00020006 | 0.00020013 | 0.00020007 | 0.00020011
| 0.1 | 0.0003 | 0.2 | 0.0 | 0.1 | 0.0003 | 20.0 | 10.0 | 0.0 | 50.0185635E+00 | DC220MV |  | 0.00030008927380499996 | 3.4790117896068012e-08 | 1517776349.47 | 1517776365.004 | 0.00030013 | 0.00030008 | 0.00030007 | 0.00030007 | 0.00030005 | 0.00030005 | 0.00030019 | 0.0003001 | 0.00030006 | 0.00030012 | 0.00030013 | 0.00030004 | 0.00030009 | 0.00030007 | 0.0003001 | 0.0003001 | 0.00030007 | 0.00030008 | 0.00030009 | 0.00030009
| 0.1 | 0.0004 | 0.2 | 0.0001 | 0.1 | 0.0003 | 20.0 | 10.0 | 0.0 | 49.9762094E+00 | DC220MV |  | 0.000300078424555 | 3.9018872275657237e-08 | 1517776379.848 | 1517776395.358 | 0.00030001 | 0.00030001 | 0.00030012 | 0.00030004 | 0.00030004 | 0.0003001 | 0.00030007 | 0.00030008 | 0.00030009 | 0.0003001 | 0.00030009 | 0.00030016 | 0.00030005 | 0.00030012 | 0.00030008 | 0.00030012 | 0.00030009 | 0.00030009 | 0.00030008 | 0.00030003
| 0.1 | 0.0004 | 0.2 | 0.0 | 0.1 | 0.0004 | 20.0 | 10.0 | 0.0 | 50.0553064E+00 | DC220MV |  | 0.00040005476513 | 4.9497914061213797e-08 | 1517776410.223 | 1517776425.768 | 0.0004001 | 0.00040007 | 0.0004001 | 0.00040005 | 0.00040006 | 0.00040012 | 0.0004 | 0.00040007 | 0.00040005 | 0.00040007 | 0.00040006 | 0.00040009 | 0.00040001 | 0.00040015 | 0.00040002 | 0.00040008 | 0.00040006 | 0.00039994 | 0.0004 | 0.00039998
| 0.1 | 0.0005 | 0.2 | 0.0001 | 0.1 | 0.0004 | 20.0 | 10.0 | 0.0 | 49.9830246E+00 | DC220MV |  | 0.00040005412939500003 | 4.0464980473347103e-08 | 1517776440.623 | 1517776456.127 | 0.00040013 | 0.00040007 | 0.00040007 | 0.00040005 | 0.00040007 | 0.00040001 | 0.00040005 | 0.00040003 | 0.00040001 | 0.00040013 | 0.00040001 | 0.00040007 | 0.00040008 | 0.00040011 | 0.00040007 | 0.00039999 | 0.00040007 | 0.00039999 | 0.00040001 | 0.00040007
| 0.1 | 0.0005 | 0.2 | 0.0 | 0.1 | 0.0005 | 20.0 | 10.0 | 0.0 | 50.0488826E+00 | DC220MV |  | 0.00050004765017999989 | 4.5801264485659674e-08 | 1517776470.977 | 1517776486.515 | 0.00050001 | 0.00050004 | 0.00050011 | 0.00050006 | 0.00050008 | 0.0005 | 0.00050008 | 0.00050009 | 0.00050008 | 0.00050001 | 0.00050009 | 0.00050009 | 0.00050008 | 0.00050004 | 0.00049993 | 0.00050009 | 0.00050005 | 0.00049996 | 0.00050002 | 0.00050003
| 0.1 | 0.0006 | 0.2 | 0.0001 | 0.1 | 0.0004999999999999999 | 20.0 | 10.0 | 0.0 | 49.9760351E+00 | DC220MV |  | 0.00050002623149500002 | 4.4446273694235255e-08 | 1517776501.353 | 1517776516.849 | 0.00050005 | 0.00050004 | 0.00050002 | 0.00050007 | 0.00050008 | 0.0005 | 0.00049999 | 0.00049995 | 0.00050008 | 0.00050003 | 0.00050003 | 0.00050003 | 0.00050001 | 0.00050003 | 0.00049998 | 0.00050009 | 0.00050004 | 0.00050006 | 0.00050002 | 0.0004999
| 0.1 | 0.0006 | 0.2 | 0.0 | 0.1 | 0.0006 | 20.0 | 10.0 | 0.0 | 50.0074256E+00 | DC220MV |  | 0.00060001752180500003 | 3.4764319099533034e-08 | 1517776531.708 | 1517776547.253 | 0.00060004 | 0.00059996 | 0.00059997 | 0.00060001 | 0.00060005 | 0.00060002 | 0.00060004 | 0.00060001 | 0.00060009 | 0.00059999 | 0.00059997 | 0.00060002 | 0.00059998 | 0.00060003 | 0.00060004 | 0.00060007 | 0.00060002 | 0.00059999 | 0.00059999 | 0.00060006
| 0.1 | 0.0007 | 0.2 | 0.0001 | 0.1 | 0.0006 | 20.0 | 10.0 | 0.0 | 49.9890543E+00 | DC220MV |  | 0.00060002568232 | 3.2611807700878811e-08 | 1517776562.075 | 1517776577.588 | 0.00060002 | 0.00060006 | 0.00060007 | 0.00060004 | 0.00060004 | 0.0006 | 0.00059997 | 0.00059999 | 0.00060002 | 0.00060006 | 0.00059999 | 0.00060003 | 0.00060009 | 0.00060002 | 0.00059998 | 0.00060001 | 0.00060007 | 0.00060001 | 0.00060003 | 0.0006
| 0.1 | 0.0007 | 0.2 | 0.0 | 0.1 | 0.0007 | 20.0 | 10.0 | 0.0 | 49.9967974E+00 | DC220MV |  | 0.00070000189120499989 | 3.9145755593358091e-08 | 1517776592.442 | 1517776607.974 | 0.00070008 | 0.00070001 | 0.00070001 | 0.00069996 | 0.00069993 | 0.00069999 | 0.00069998 | 0.00070004 | 0.00070002 | 0.0007 | 0.00070001 | 0.00069997 | 0.00070001 | 0.00070001 | 0.00069998 | 0.0007001 | 0.00069996 | 0.00069999 | 0.00069996 | 0.00070003
| 0.1 | 0.0008 | 0.2 | 0.0001 | 0.1 | 0.0007 | 20.0 | 10.0 | 0.0 | 50.0688257E+00 | DC220MV |  | 0.00070000940223999991 | 3.4637014573450092e-08 | 1517776622.848 | 1517776638.377 | 0.00070006 | 0.00069999 | 0.00070002 | 0.00070004 | 0.00069999 | 0.00070005 | 0.0007 | 0.00070002 | 0.00069998 | 0.0007 | 0.00070001 | 0.00070006 | 0.00069998 | 0.00070001 | 0.00070002 | 0.00069989 | 0.0007 | 0.00070001 | 0.00070003 | 0.00070003
| 0.1 | 0.0008 | 0.2 | 0.0 | 0.1 | 0.0008 | 20.0 | 10.0 | 0.0 | 50.0126070E+00 | DC220MV |  | 0.00079998750573999995 | 4.5677001716022755e-08 | 1517776653.229 | 1517776668.734 | 0.00080001 | 0.00079999 | 0.00080002 | 0.00080012 | 0.0008 | 0.00079999 | 0.00079997 | 0.00080002 | 0.00079993 | 0.00079996 | 0.00080001 | 0.00079998 | 0.00079993 | 0.00079997 | 0.0008 | 0.00079992 | 0.00079996 | 0.00079992 | 0.00080003 | 0.00080001
| 0.1 | 0.0009 | 0.2 | 0.0001 | 0.1 | 0.0007999999999999999 | 20.0 | 10.0 | 0.0 | 50.0336594E+00 | DC220MV |  | 0.00079996429940000003 | 3.9018612342236816e-08 | 1517776683.584 | 1517776699.103 | 0.00079991 | 0.00079998 | 0.00079993 | 0.00079997 | 0.00080001 | 0.00079996 | 0.00079995 | 0.00079999 | 0.00079988 | 0.00079993 | 0.00080001 | 0.00079993 | 0.0008 | 0.00080001 | 0.00079997 | 0.00079999 | 0.0007999 | 0.00079997 | 0.0008 | 0.0008
| 0.1 | 0.0009 | 0.2 | 0.0 | 0.1 | 0.0009 | 20.0 | 10.0 | 0.0 | 50.0161467E+00 | DC220MV |  | 0.00089995843870499992 | 4.2050803697280008e-08 | 1517776713.95 | 1517776729.468 | 0.00089999 | 0.00089993 | 0.00089999 | 0.00089999 | 0.00089997 | 0.00089996 | 0.00089991 | 0.00090002 | 0.00089998 | 0.00089991 | 0.0009 | 0.00089991 | 0.0009 | 0.0008999 | 0.00089999 | 0.00089996 | 0.00089997 | 0.00089988 | 0.0008999 | 0.0009
| 0.1 | 0.001 | 0.2 | 0.0001 | 0.1 | 0.0009 | 20.0 | 10.0 | 0.0 | 50.0312695E+00 | DC220MV |  | 0.00089997020185499999 | 3.6497372742682808e-08 | 1517776744.296 | 1517776759.824 | 0.00089996 | 0.00090001 | 0.00089997 | 0.00089997 | 0.00090005 | 0.00089995 | 0.00089996 | 0.0008999 | 0.00089993 | 0.00089993 | 0.0009 | 0.0009 | 0.00089994 | 0.00090001 | 0.00089999 | 0.00089998 | 0.00089999 | 0.00089995 | 0.0009 | 0.00089991
| 0.1 | 0.001 | 0.2 | 0.0 | 0.1 | 0.001 | 20.0 | 10.0 | 0.0 | 49.9593209E+00 | DC220MV |  | 0.00099662196571499999 | 5.5045867910616554e-06 | 1517776774.668 | 1517776790.188 | 0.00099868 | 0.00099793 | 0.00099852 | 0.00099769 | 0.00099871 | 0.00099088 | 0.00098044 | 0.0009841 | 0.00099261 | 0.00099322 | 0.001 | 0.00099993 | 0.00099996 | 0.00099996 | 0.00099997 | 0.00099998 | 0.00099997 | 0.00099999 | 0.00099996 | 0.00099993
| 0.1 | 0.001 | 0.2 | 0.0001 | 0.1 | 0.0009 | 20.0 | 10.0 | 0.0 | 50.0201543E+00 | DC220MV |  | 0.00089994083547000005 | 4.9917959303988711e-08 | 1517776805.053 | 1517776820.606 | 0.00089999 | 0.00089998 | 0.00089995 | 0.00089992 | 0.00089978 | 0.00089996 | 0.0008999 | 0.00089993 | 0.00089996 | 0.0009 | 0.00089996 | 0.00089996 | 0.00089995 | 0.00089992 | 0.00089992 | 0.00089992 | 0.00089992 | 0.00089994 | 0.00089992 | 0.00090003
| 0.1 | 0.0009 | 0.2 | 0.0 | 0.1 | 0.0009 | 20.0 | 10.0 | 0.0 | 49.9943827E+00 | DC220MV |  | 0.00089995372983499994 | 3.7228286785966662e-08 | 1517776835.46 | 1517776850.966 | 0.00089998 | 0.00089997 | 0.00089997 | 0.00089991 | 0.00089998 | 0.00089994 | 0.00089995 | 0.00089989 | 0.00089993 | 0.0008999 | 0.00089992 | 0.00090001 | 0.00089997 | 0.00089999 | 0.00089999 | 0.0008999 | 0.00089994 | 0.00089998 | 0.00090002 | 0.00089996
| 0.1 | 0.0009 | 0.2 | 0.0001 | 0.1 | 0.0007999999999999999 | 20.0 | 10.0 | 0.0 | 50.0019246E+00 | DC220MV |  | 0.00079995451526999992 | 6.1286948362941423e-08 | 1517776865.822 | 1517776881.359 | 0.00079992 | 0.00080001 | 0.00079992 | 0.00079994 | 0.00079998 | 0.00079994 | 0.00079992 | 0.00079999 | 0.00079996 | 0.00079993 | 0.00079997 | 0.00079995 | 0.00079998 | 0.0007999 | 0.00080016 | 0.00079991 | 0.00079988 | 0.00080001 | 0.00079986 | 0.00079995
| 0.1 | 0.0008 | 0.2 | 0.0 | 0.1 | 0.0008 | 20.0 | 10.0 | 0.0 | 49.9987076E+00 | DC220MV |  | 0.00079993740025999997 | 2.707772608064984e-08 | 1517776896.201 | 1517776911.729 | 0.00079995 | 0.00079997 | 0.00079991 | 0.00079994 | 0.0007999 | 0.00079995 | 0.0007999 | 0.00079997 | 0.00079994 | 0.0008 | 0.00079996 | 0.00079995 | 0.00079995 | 0.00079993 | 0.0007999 | 0.00079993 | 0.00079996 | 0.0007999 | 0.00079991 | 0.00079993
| 0.1 | 0.0008 | 0.2 | 0.0001 | 0.1 | 0.0007 | 20.0 | 10.0 | 0.0 | 49.9925954E+00 | DC220MV |  | 0.00069992874583999995 | 4.0863199848371339e-08 | 1517776926.592 | 1517776942.128 | 0.00069991 | 0.00069996 | 0.0007 | 0.00069992 | 0.00069992 | 0.00069984 | 0.00069991 | 0.00069995 | 0.00069986 | 0.00069992 | 0.00069994 | 0.00069997 | 0.00069988 | 0.00069993 | 0.00069993 | 0.00069993 | 0.00070001 | 0.00069998 | 0.0006999 | 0.00069993
| 0.1 | 0.0007 | 0.2 | 0.0 | 0.1 | 0.0007 | 20.0 | 10.0 | 0.0 | 50.0353741E+00 | DC220MV |  | 0.00069991437031500007 | 4.2018796792347602e-08 | 1517776956.974 | 1517776972.517 | 0.00069987 | 0.00069995 | 0.00069993 | 0.00069995 | 0.00069991 | 0.00069989 | 0.00069992 | 0.00069982 | 0.00069984 | 0.00069995 | 0.00069992 | 0.00069991 | 0.00069993 | 0.00069994 | 0.00069987 | 0.0007 | 0.0006999 | 0.00069991 | 0.00069998 | 0.0006999
| 0.1 | 0.0007 | 0.2 | 0.0001 | 0.1 | 0.0006 | 20.0 | 10.0 | 0.0 | 50.0157550E+00 | DC220MV |  | 0.00059990424703000003 | 4.6420166181136469e-08 | 1517776987.38 | 1517777002.907 | 0.00059992 | 0.00059994 | 0.00059994 | 0.00059993 | 0.0005999 | 0.00059975 | 0.00059993 | 0.00059986 | 0.00059989 | 0.00059989 | 0.0005999 | 0.00059989 | 0.00059992 | 0.00059989 | 0.0005999 | 0.00059989 | 0.00059999 | 0.00059989 | 0.00059992 | 0.00059997
| 0.1 | 0.0006 | 0.2 | 0.0 | 0.1 | 0.0006 | 20.0 | 10.0 | 0.0 | 50.0132771E+00 | DC220MV |  | 0.00059993784861999994 | 3.7768242747927249e-08 | 1517777017.765 | 1517777033.291 | 0.00059993 | 0.00059991 | 0.0006 | 0.00059994 | 0.00059994 | 0.00059994 | 0.00059989 | 0.00059993 | 0.00059996 | 0.00059997 | 0.00059989 | 0.00059985 | 0.00059997 | 0.00059995 | 0.00059992 | 0.00059997 | 0.00059989 | 0.00059997 | 0.0006 | 0.00059992
| 0.1 | 0.0006 | 0.2 | 0.0001 | 0.1 | 0.0004999999999999999 | 20.0 | 10.0 | 0.0 | 50.0137632E+00 | DC220MV |  | 0.00049991611742000004 | 4.1550440928239934e-08 | 1517777048.152 | 1517777063.685 | 0.00049994 | 0.00049992 | 0.0005 | 0.00049995 | 0.0004999 | 0.00049992 | 0.00049988 | 0.00049993 | 0.00049988 | 0.00049992 | 0.0005 | 0.00049987 | 0.00049989 | 0.00049994 | 0.00049989 | 0.00049994 | 0.00049996 | 0.00049989 | 0.00049987 | 0.00049983
| 0.1 | 0.0005 | 0.2 | 0.0 | 0.1 | 0.0005 | 20.0 | 10.0 | 0.0 | 49.9867394E+00 | DC220MV |  | 0.00049989783748500009 | 3.2909197701983479e-08 | 1517777078.539 | 1517777094.074 | 0.00049994 | 0.00049997 | 0.00049986 | 0.00049991 | 0.00049989 | 0.00049988 | 0.00049992 | 0.00049993 | 0.00049986 | 0.00049991 | 0.00049985 | 0.00049987 | 0.00049987 | 0.00049987 | 0.0004999 | 0.00049988 | 0.0004999 | 0.0004999 | 0.00049996 | 0.0004999
| 0.1 | 0.0005 | 0.2 | 0.0001 | 0.1 | 0.0004 | 20.0 | 10.0 | 0.0 | 49.9704373E+00 | DC220MV |  | 0.00039991011093500002 | 4.0475240867462456e-08 | 1517777108.933 | 1517777124.471 | 0.00039995 | 0.00039991 | 0.00039991 | 0.00039988 | 0.00039994 | 0.00039988 | 0.00039991 | 0.0003999 | 0.00039994 | 0.00039986 | 0.00039988 | 0.00039994 | 0.00039988 | 0.00039987 | 0.0003999 | 0.00039982 | 0.00039995 | 0.00039998 | 0.0003999 | 0.00039999
| 0.1 | 0.0004 | 0.2 | 0.0 | 0.1 | 0.0004 | 20.0 | 10.0 | 0.0 | 50.0013774E+00 | DC220MV |  | 0.00039989124587999999 | 4.2567291503750217e-08 | 1517777139.314 | 1517777154.864 | 0.00039997 | 0.00039999 | 0.00039981 | 0.00039996 | 0.00039984 | 0.00039988 | 0.0003999 | 0.00039987 | 0.00039989 | 0.00039985 | 0.00039987 | 0.0003999 | 0.00039991 | 0.00039988 | 0.00039992 | 0.0003999 | 0.00039991 | 0.00039988 | 0.00039985 | 0.00039987
| 0.1 | 0.0004 | 0.2 | 0.0001 | 0.1 | 0.0003 | 20.0 | 10.0 | 0.0 | 49.9858153E+00 | DC220MV |  | 0.00029988710353500006 | 4.0311907340083587e-08 | 1517777169.74 | 1517777185.261 | 0.0002998 | 0.00029988 | 0.00029991 | 0.00029995 | 0.00029989 | 0.00029985 | 0.00029988 | 0.00029993 | 0.00029991 | 0.00029988 | 0.00029986 | 0.00029987 | 0.00029992 | 0.00029979 | 0.00029992 | 0.00029991 | 0.00029988 | 0.00029992 | 0.00029992 | 0.0002999
| 0.1 | 0.0003 | 0.2 | 0.0 | 0.1 | 0.0003 | 20.0 | 10.0 | 0.0 | 49.9823223E+00 | DC220MV |  | 0.00029988565370000001 | 2.4767784587379899e-08 | 1517777200.111 | 1517777215.642 | 0.00029989 | 0.00029991 | 0.00029989 | 0.00029985 | 0.00029989 | 0.00029988 | 0.00029989 | 0.00029993 | 0.00029986 | 0.0002999 | 0.00029983 | 0.0002999 | 0.00029986 | 0.0002999 | 0.00029986 | 0.00029987 | 0.00029989 | 0.00029991 | 0.00029992 | 0.00029986
| 0.1 | 0.0003 | 0.2 | 0.0001 | 0.1 | 0.0002 | 20.0 | 10.0 | 0.0 | 49.9975945E+00 | DC220MV |  | 0.00019988967309000002 | 5.3666392895853873e-08 | 1517777230.494 | 1517777246.026 | 0.00019989 | 0.00019989 | 0.00019987 | 0.00020005 | 0.00019985 | 0.00019979 | 0.00019988 | 0.00019993 | 0.00019981 | 0.00019986 | 0.00019986 | 0.00019992 | 0.00019988 | 0.00019987 | 0.00019989 | 0.00019993 | 0.00019994 | 0.00019993 | 0.00019984 | 0.00019989
| 0.1 | 0.0002 | 0.2 | 0.0 | 0.1 | 0.0002 | 20.0 | 10.0 | 0.0 | 49.9860416E+00 | DC220MV |  | 0.00019986517213000003 | 4.936901526897584e-08 | 1517777260.884 | 1517777276.419 | 0.00019982 | 0.00019999 | 0.00019988 | 0.00019989 | 0.00019989 | 0.00019997 | 0.00019986 | 0.00019987 | 0.00019988 | 0.00019988 | 0.00019983 | 0.00019984 | 0.00019982 | 0.0001998 | 0.00019985 | 0.00019989 | 0.00019984 | 0.00019984 | 0.00019989 | 0.00019978
| 0.1 | 0.0002 | 0.2 | 0.0001 | 0.1 | 0.0001 | 20.0 | 10.0 | 0.0 | 49.9908837E+00 | DC220MV |  | 9.9880677900499989e-05 | 3.7925942760439445e-08 | 1517777291.275 | 1517777306.804 | 9.9949e-05 | 9.9857e-05 | 9.9928e-05 | 9.991e-05 | 9.993e-05 | 9.9866e-05 | 9.9908e-05 | 9.9842e-05 | 9.9839e-05 | 9.9947e-05 | 9.984e-05 | 9.9915e-05 | 9.9856e-05 | 9.9866e-05 | 9.9851e-05 | 9.9881e-05 | 9.9864e-05 | 9.9815e-05 | 9.9864e-05 | 9.9885e-05
| 0.1 | 0.0001 | 0.2 | 0.0 | 0.1 | 0.0001 | 20.0 | 10.0 | 0.0 | 49.9947364E+00 | DC220MV |  | 9.9895164034999984e-05 | 4.1521815972941454e-08 | 1517777321.65 | 1517777337.195 | 9.9864e-05 | 9.9919e-05 | 9.989e-05 | 9.9859e-05 | 9.9917e-05 | 9.982e-05 | 9.9862e-05 | 9.9903e-05 | 9.9877e-05 | 9.9897e-05 | 9.9954e-05 | 9.9887e-05 | 9.986e-05 | 9.9917e-05 | 9.9924e-05 | 0.0001 | 9.9891e-05 | 9.9824e-05 | 9.9926e-05 | 9.9912e-05
| 0.1 | 0.0001 | 0.2 | 0.0001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9921427E+00 | DC220MV |  | -1.0457990797649998e-07 | 3.5403817539236998e-08 | 1517777352.049 | 1517777367.605 | -9.3722e-08 | -7.3712e-08 | -9.9166e-08 | -1.4433e-07 | -6.5031e-08 | -1.3374e-07 | -9.431e-08 | -6.8857e-08 | -1.3153e-07 | -1.3389e-07 | -1.6273e-07 | -1.0682e-07 | -1.5552e-07 | -1.6935e-07 | -1.1094e-07 | -6.7238e-08 | -8.2981e-08 | -7.0328e-08 | -4.9289e-08 | -7.8126e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0079211E+00 | DC220MV |  | -1.0092640449900001e-07 | 3.9504082234853246e-08 | 1517777382.452 | 1517777398.014 | -1.4724e-07 | -8.9137e-08 | -6.8544e-08 | -1.0944e-07 | -1.3003e-07 | -1.8945e-07 | -1.6062e-07 | -7.2075e-08 | -6.016e-08 | -1.0576e-07 | -7.737e-08 | -6.1925e-08 | -9.2814e-08 | -1.0576e-07 | -9.9433e-08 | -1.0002e-07 | -1.7136e-07 | -4.1185e-08 | -5.8395e-08 | -7.7811e-08
| 0.1 | 0.0001 | 0.2 | 0.0001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0167226E+00 | DC220MV |  | -1.115853687366e-07 | 4.3409448364541413e-08 | 1517777412.879 | 1517777428.42 | -1.0113e-07 | -1.2571e-07 | -1.1982e-07 | -1.2394e-07 | -2.6938e-08 | -1.3719e-07 | -9.6269e-08 | -1.1702e-07 | -1.2718e-07 | -1.1688e-07 | -1.3219e-07 | -2.2374e-08 | -9.9066e-08 | -1.4588e-07 | -1.6516e-07 | -1.1173e-07 | -1.4175e-07 | -1.4676e-07 | -1.6766e-07 | -7.0656e-09
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9971700E+00 | DC220MV |  | -1.11080996218e-07 | 3.2562010188762027e-08 | 1517777443.26 | 1517777458.803 | -8.081e-08 | -1.1908e-07 | -1.2099e-07 | -1.1658e-07 | -1.3557e-07 | -1.4528e-07 | -2.9733e-08 | -1.1746e-07 | -6.6532e-08 | -1.55e-07 | -5.564e-08 | -9.6707e-08 | -1.5794e-07 | -1.2835e-07 | -1.0775e-07 | -1.2099e-07 | -1.1908e-07 | -8.2135e-08 | -1.3527e-07 | -1.3071e-07
| 0.1 | 0.0001 | 0.2 | 0.0 | 0.1 | 0.0001 | 20.0 | 10.0 | 0.0 | 50.0126966E+00 | DC220MV |  | 9.9879676124500012e-05 | 3.758416508460765e-08 | 1517777473.636 | 1517777489.183 | 9.9917e-05 | 9.9845e-05 | 9.9956e-05 | 9.9892e-05 | 9.9814e-05 | 9.9835e-05 | 9.984e-05 | 9.9899e-05 | 9.991e-05 | 9.99e-05 | 9.9832e-05 | 9.9862e-05 | 9.9849e-05 | 9.9887e-05 | 9.9884e-05 | 9.9847e-05 | 9.9912e-05 | 9.9881e-05 | 9.9944e-05 | 9.9888e-05
| 0.1 | -0.0001 | 0.2 | 0.0 | 0.1 | -0.0001 | 20.0 | 10.0 | 0.0 | 49.9852166E+00 | DC220MV |  | -0.000100102002451 | 5.8893975048870975e-08 | 1517777504.04 | 1517777519.547 | -0.00010013 | -0.00010009 | -0.00010013 | -0.00010007 | -0.0001001 | -0.00010014 | -0.00010012 | -0.00010006 | -0.00010011 | -0.00010016 | -0.00010005 | -0.00010005 | -0.00010014 | -9.9979e-05 | -0.00010007 | -0.00010014 | -0.00010014 | -0.00010004 | -0.00010005 | -0.00010027
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0409817E+00 | DC220MV |  | -1.0924657843199999e-07 | 4.7942320756690017e-08 | 1517777534.388 | 1517777549.957 | -1.5298e-07 | -1.1017e-07 | -2.4859e-08 | -5.0159e-08 | -8.6932e-08 | -6.1338e-08 | -1.3312e-07 | -1.1473e-07 | -1.8343e-07 | -1.9887e-07 | -8.3108e-08 | -6.2809e-08 | -1.0797e-07 | -8.5462e-08 | -6.3692e-08 | -1.8504e-07 | -8.5609e-08 | -8.7962e-08 | -1.3944e-07 | -1.6725e-07
| 0.1 | -0.0001 | 0.2 | -0.0001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0256452E+00 | DC220MV |  | -8.7332700241849998e-08 | 4.8868244034103607e-08 | 1517777564.832 | 1517777580.38 | -1.3997e-07 | -5.284e-08 | -1.4056e-07 | -8.9047e-08 | -1.3791e-07 | -7.6537e-08 | -1.3865e-07 | -6.1524e-08 | -3.4883e-08 | -7.271e-08 | -1.0583e-07 | -8.7281e-08 | -3.5913e-08 | -5.9905e-08 | -2.355e-09 | -1.3041e-07 | -3.3706e-08 | -1.4836e-07 | -1.7765e-07 | -2.0606e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0049061E+00 | DC220MV |  | -1.24236533793e-07 | 3.6656918342398129e-08 | 1517777595.231 | 1517777610.773 | -1.0716e-07 | -1.2467e-07 | -9.8324e-08 | -7.7275e-08 | -3.7239e-08 | -1.1525e-07 | -1.4219e-07 | -1.7722e-07 | -1.7898e-07 | -1.335e-07 | -7.3743e-08 | -1.7707e-07 | -1.0186e-07 | -1.2555e-07 | -9.4791e-08 | -1.491e-07 | -1.207e-07 | -1.5794e-07 | -1.6044e-07 | -1.3174e-07
| 0.1 | -0.0001 | 0.2 | -0.0001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0313357E+00 | DC220MV |  | -1.0852081126200002e-07 | 2.8245173100622767e-08 | 1517777625.597 | 1517777641.167 | -1.0895e-07 | -1.2101e-07 | -1.3601e-07 | -1.0366e-07 | -1.2322e-07 | -1.2292e-07 | -1.2763e-07 | -1.1263e-07 | -1.3072e-07 | -9.8662e-08 | -3.6171e-08 | -9.1898e-08 | -1.1351e-07 | -1.3483e-07 | -1.3277e-07 | -1.1395e-07 | -1.3821e-07 | -9.4251e-08 | -9.5133e-08 | -3.426e-08
| 0.1 | -0.0001 | 0.2 | 0.0 | 0.1 | -0.0001 | 20.0 | 10.0 | 0.0 | 49.9978020E+00 | DC220MV |  | -0.0001001032105 | 3.7882339511149464e-08 | 1517777656.018 | 1517777671.528 | -0.00010011 | -0.00010006 | -0.00010009 | -0.00010005 | -0.00010013 | -0.00010008 | -0.00010012 | -0.00010011 | -0.00010017 | -0.00010015 | -0.00010009 | -0.00010008 | -0.00010012 | -0.00010011 | -0.00010014 | -0.00010005 | -0.00010007 | -0.00010014 | -0.00010004 | -0.00010015
| 0.1 | -0.0002 | 0.2 | -0.0001 | 0.1 | -0.0001 | 20.0 | 10.0 | 0.0 | 49.9641342E+00 | DC220MV |  | -0.00010007398044949999 | 6.9187635114952069e-08 | 1517777686.381 | 1517777701.901 | -0.00010007 | -0.00010011 | -0.00010005 | -0.00010013 | -0.00010007 | -0.00010006 | -0.00010013 | -0.0001 | -0.00010003 | -0.00010013 | -0.00010008 | -0.0001001 | -0.00010011 | -0.00010004 | -0.00010002 | -0.00010011 | -9.9843e-05 | -0.00010005 | -0.00010017 | -0.00010016
| 0.1 | -0.0002 | 0.2 | 0.0 | 0.1 | -0.0002 | 20.0 | 10.0 | 0.0 | 50.0528388E+00 | DC220MV |  | -0.00020009818263999997 | 3.419980611925036e-08 | 1517777716.729 | 1517777732.267 | -0.0002001 | -0.00020007 | -0.0002001 | -0.00020011 | -0.00020022 | -0.00020007 | -0.00020011 | -0.00020005 | -0.00020011 | -0.00020008 | -0.00020009 | -0.00020012 | -0.00020007 | -0.00020007 | -0.00020006 | -0.00020011 | -0.00020009 | -0.00020011 | -0.00020013 | -0.00020011
| 0.1 | -0.0003 | 0.2 | -0.0001 | 0.1 | -0.0002 | 20.0 | 10.0 | 0.0 | 49.9985991E+00 | DC220MV |  | -0.00020008351243500002 | 3.7368676268249033e-08 | 1517777747.111 | 1517777762.612 | -0.00020001 | -0.00020014 | -0.00020009 | -0.00020003 | -0.00020002 | -0.00020008 | -0.00020009 | -0.00020014 | -0.00020012 | -0.00020003 | -0.0002001 | -0.00020008 | -0.00020008 | -0.00020008 | -0.00020008 | -0.00020012 | -0.00020007 | -0.0002001 | -0.00020012 | -0.00020007
| 0.1 | -0.0003 | 0.2 | 0.0 | 0.1 | -0.0003 | 20.0 | 10.0 | 0.0 | 50.0186343E+00 | DC220MV |  | -0.00030009289342499996 | 5.0405877596957804e-08 | 1517777777.459 | 1517777792.996 | -0.00030006 | -0.00030012 | -0.00030012 | -0.00030027 | -0.00030005 | -0.00030006 | -0.0003001 | -0.00030003 | -0.00030011 | -0.00030009 | -0.00030005 | -0.00030015 | -0.00030007 | -0.00030011 | -0.00030008 | -0.0003001 | -0.00030006 | -0.00030009 | -0.00030007 | -0.00030005
| 0.1 | -0.0004 | 0.2 | -0.0001 | 0.1 | -0.0003 | 20.0 | 10.0 | 0.0 | 49.9737167E+00 | DC220MV |  | -0.00030009706850000009 | 4.6584244479384418e-08 | 1517777807.847 | 1517777823.353 | -0.00030006 | -0.00030009 | -0.00030007 | -0.00030005 | -0.00030008 | -0.00030008 | -0.00030006 | -0.00030011 | -0.00029999 | -0.00030012 | -0.00030005 | -0.00030011 | -0.00030019 | -0.00030011 | -0.00030019 | -0.00030016 | -0.00030009 | -0.00030008 | -0.00030012 | -0.00030013
| 0.1 | -0.0004 | 0.2 | 0.0 | 0.1 | -0.0004 | 20.0 | 10.0 | 0.0 | 50.0135036E+00 | DC220MV |  | -0.00040011419394999991 | 4.6127531030008896e-08 | 1517777838.212 | 1517777853.731 | -0.00040007 | -0.00040016 | -0.00040006 | -0.00040005 | -0.00040011 | -0.00040002 | -0.00040018 | -0.00040011 | -0.00040019 | -0.00040006 | -0.00040013 | -0.00040009 | -0.00040016 | -0.00040011 | -0.00040016 | -0.00040009 | -0.00040012 | -0.00040013 | -0.00040016 | -0.00040013
| 0.1 | -0.0005 | 0.2 | -0.0001 | 0.1 | -0.0004 | 20.0 | 10.0 | 0.0 | 49.9962739E+00 | DC220MV |  | -0.00040012293874000006 | 4.4668910679371655e-08 | 1517777868.58 | 1517777884.101 | -0.00040016 | -0.00040012 | -0.00040011 | -0.00040009 | -0.00040011 | -0.00040004 | -0.00040012 | -0.00040016 | -0.00040019 | -0.00040012 | -0.00040012 | -0.00040019 | -0.00040014 | -0.00040019 | -0.00040009 | -0.00040009 | -0.0004001 | -0.00040016 | -0.00040002 | -0.00040014
| 0.1 | -0.0005 | 0.2 | 0.0 | 0.1 | -0.0005 | 20.0 | 10.0 | 0.0 | 50.0137019E+00 | DC220MV |  | -0.00050010576267000011 | 3.1016581420955444e-08 | 1517777898.96 | 1517777914.481 | -0.00050009 | -0.00050009 | -0.00050011 | -0.00050014 | -0.0005001 | -0.00050013 | -0.00050006 | -0.00050013 | -0.00050011 | -0.0005001 | -0.00050011 | -0.00050016 | -0.00050008 | -0.00050003 | -0.00050008 | -0.00050008 | -0.00050014 | -0.00050013 | -0.00050014 | -0.00050012
| 0.1 | -0.0006 | 0.2 | -0.0001 | 0.1 | -0.0004999999999999999 | 20.0 | 10.0 | 0.0 | 49.9607808E+00 | DC220MV |  | -0.00050007717973499992 | 4.3498504291966048e-08 | 1517777929.335 | 1517777944.853 | -0.0005002 | -0.00050004 | -0.00050005 | -0.00050012 | -0.00050017 | -0.00050006 | -0.00050005 | -0.00050007 | -0.00050002 | -0.00050007 | -0.00050009 | -0.00050007 | -0.00050005 | -0.00050006 | -0.0005001 | -0.00050008 | -0.00050008 | -0.00050004 | -0.00050002 | -0.0005001
| 0.1 | -0.0006 | 0.2 | 0.0 | 0.1 | -0.0006 | 20.0 | 10.0 | 0.0 | 50.0237660E+00 | DC220MV |  | -0.00060011002273000001 | 3.2138452543012399e-08 | 1517777959.706 | 1517777975.247 | -0.00060015 | -0.00060011 | -0.00060007 | -0.00060013 | -0.00060011 | -0.00060013 | -0.00060005 | -0.0006001 | -0.00060014 | -0.00060008 | -0.00060006 | -0.00060013 | -0.00060012 | -0.00060017 | -0.00060009 | -0.00060012 | -0.00060012 | -0.00060012 | -0.00060014 | -0.00060006
| 0.1 | -0.0007 | 0.2 | -0.0001 | 0.1 | -0.0006 | 20.0 | 10.0 | 0.0 | 50.0072935E+00 | DC220MV |  | -0.00060009408991000004 | 4.5488115746675098e-08 | 1517777990.103 | 1517778005.614 | -0.00060011 | -0.00060011 | -0.00060011 | -0.0006001 | -0.00060007 | -0.0006001 | -0.00060008 | -0.0006002 | -0.00060009 | -0.00060012 | -0.00060006 | -0.00060004 | -0.00060015 | -0.0006001 | -0.00060001 | -0.00060013 | -0.0006001 | -0.00060013 | -0.0006 | -0.00060006
| 0.1 | -0.0007 | 0.2 | 0.0 | 0.1 | -0.0007 | 20.0 | 10.0 | 0.0 | 50.0137915E+00 | DC220MV |  | -0.00070011858391999993 | 4.5492077861263666e-08 | 1517778020.479 | 1517778036.009 | -0.00070007 | -0.00070013 | -0.00070013 | -0.00070005 | -0.00070006 | -0.00070013 | -0.00070015 | -0.00070013 | -0.00070009 | -0.00070017 | -0.00070008 | -0.00070011 | -0.0007001 | -0.00070015 | -0.00070009 | -0.0007001 | -0.00070016 | -0.00070013 | -0.00070008 | -0.00070025
| 0.1 | -0.0008 | 0.2 | -0.0001 | 0.1 | -0.0007 | 20.0 | 10.0 | 0.0 | 50.0064112E+00 | DC220MV |  | -0.00070007309651000003 | 2.8297835999681086e-08 | 1517778050.867 | 1517778066.378 | -0.00070003 | -0.00070006 | -0.0007001 | -0.0007001 | -0.00070012 | -0.00070006 | -0.00070009 | -0.00070004 | -0.00070008 | -0.00070005 | -0.00070006 | -0.0007001 | -0.00070007 | -0.00070006 | -0.00070006 | -0.00070004 | -0.00070006 | -0.00070008 | -0.00070014 | -0.00070005
| 0.1 | -0.0008 | 0.2 | 0.0 | 0.1 | -0.0008 | 20.0 | 10.0 | 0.0 | 49.9795932E+00 | DC220MV |  | -0.00080008410895500014 | 5.2350817133293123e-08 | 1517778081.233 | 1517778096.764 | -0.00080011 | -0.00080009 | -0.0008001 | -0.00079996 | -0.00080009 | -0.00080004 | -0.00080004 | -0.00080001 | -0.00080013 | -0.00080005 | -0.00080013 | -0.00080009 | -0.00080005 | -0.00080011 | -0.00080019 | -0.00080012 | -0.00080004 | -0.00080008 | -0.00080013 | -0.00080015
| 0.1 | -0.0009 | 0.2 | -0.0001 | 0.1 | -0.0007999999999999999 | 20.0 | 10.0 | 0.0 | 50.0121303E+00 | DC220MV |  | -0.00080005318158999999 | 3.9628726246347604e-08 | 1517778111.619 | 1517778127.147 | -0.00080003 | -0.00080005 | -0.00080008 | -0.0008001 | -0.00080003 | -0.00080011 | -0.00080001 | -0.0008 | -0.00080012 | -0.00080006 | -0.00080005 | -0.00080002 | -0.00080013 | -0.00080006 | -0.00079999 | -0.00080008 | -0.00080003 | -0.00080002 | -0.00080002 | -0.00080006
| 0.1 | -0.0009 | 0.2 | 0.0 | 0.1 | -0.0009 | 20.0 | 10.0 | 0.0 | 49.9795367E+00 | DC220MV |  | -0.00090009169463999992 | 5.2515766647840719e-08 | 1517778142.001 | 1517778157.511 | -0.00090006 | -0.00090011 | -0.0009001 | -0.00090011 | -0.00090004 | -0.00090023 | -0.00090004 | -0.00090003 | -0.00090008 | -0.00090013 | -0.00090005 | -0.00090008 | -0.00090001 | -0.00090015 | -0.00090012 | -0.00090011 | -0.00090005 | -0.00090015 | -0.00090003 | -0.00090015
| 0.1 | -0.001 | 0.2 | -0.0001 | 0.1 | -0.0009 | 20.0 | 10.0 | 0.0 | 49.9799326E+00 | DC220MV |  | -0.00090008109826000013 | 5.2326689370040385e-08 | 1517778172.375 | 1517778187.905 | -0.00090013 | -0.00090011 | -0.00090008 | -0.00090001 | -0.00090014 | -0.00090009 | -0.00090007 | -0.00090012 | -0.00090007 | -0.00090001 | -0.00090009 | -0.00089998 | -0.00090017 | -0.0009001 | -0.0009001 | -0.00090011 | -0.00089998 | -0.00090003 | -0.0009001 | -0.00090014
| 0.1 | -0.001 | 0.2 | 0.0 | 0.1 | -0.001 | 20.0 | 10.0 | 0.0 | 49.9768174E+00 | DC220MV |  | -0.0010000964505650003 | 6.6725806175240022e-08 | 1517778202.731 | 1517778218.249 | -0.0010001 | -0.0010001 | -0.0010001 | -0.0010001 | -0.001 | -0.0009999 | -0.0010001 | -0.0010001 | -0.0010001 | -0.0010002 | -0.0010001 | -0.0010001 | -0.0010003 | -0.0010001 | -0.0010001 | -0.0010001 | -0.0010002 | -0.0010001 | -0.0010001 | -0.0010001
| 0.1 | -0.001 | 0.2 | -0.0001 | 0.1 | -0.0009 | 20.0 | 10.0 | 0.0 | 49.9513540E+00 | DC220MV |  | -0.00090006829590500001 | 5.4097518881492528e-08 | 1517778233.099 | 1517778248.632 | -0.00090002 | -0.0008999 | -0.00090014 | -0.00090009 | -0.0009001 | -0.00090008 | -0.00090005 | -0.00090007 | -0.00090011 | -0.00090011 | -0.00090004 | -0.00090007 | -0.00090008 | -0.00090011 | -0.00090006 | -0.00090009 | -0.00090008 | -0.00090011 | -0.0009001 | -0.00089996
| 0.1 | -0.0009 | 0.2 | 0.0 | 0.1 | -0.0009 | 20.0 | 10.0 | 0.0 | 49.9796357E+00 | DC220MV |  | -0.00090011012380499986 | 3.1319104727854013e-08 | 1517778263.482 | 1517778279.031 | -0.00090016 | -0.00090016 | -0.00090007 | -0.0009001 | -0.00090013 | -0.00090009 | -0.00090006 | -0.00090008 | -0.00090007 | -0.00090015 | -0.00090014 | -0.00090012 | -0.00090011 | -0.00090008 | -0.00090017 | -0.0009001 | -0.00090013 | -0.00090011 | -0.0009001 | -0.00090009
| 0.1 | -0.0009 | 0.2 | -0.0001 | 0.1 | -0.0007999999999999999 | 20.0 | 10.0 | 0.0 | 49.9867111E+00 | DC220MV |  | -0.00080009775683499999 | 4.0390321171826939e-08 | 1517778293.887 | 1517778309.425 | -0.00080006 | -0.00080016 | -0.00080007 | -0.00080002 | -0.00080009 | -0.0008001 | -0.00080012 | -0.00080007 | -0.00080012 | -0.00080009 | -0.00080013 | -0.00080012 | -0.00080005 | -0.00080014 | -0.00080016 | -0.00080006 | -0.00080007 | -0.00080017 | -0.00080008 | -0.00080007
| 0.1 | -0.0008 | 0.2 | 0.0 | 0.1 | -0.0008 | 20.0 | 10.0 | 0.0 | 49.9555302E+00 | DC220MV |  | -0.00080012057766500001 | 4.0442781738297303e-08 | 1517778324.263 | 1517778339.805 | -0.00080014 | -0.00080011 | -0.00080013 | -0.00080007 | -0.00080013 | -0.00080014 | -0.00080015 | -0.00080007 | -0.0008001 | -0.00080013 | -0.00080014 | -0.00080018 | -0.00080009 | -0.00079999 | -0.00080017 | -0.00080012 | -0.0008001 | -0.00080015 | -0.00080015 | -0.00080012
| 0.1 | -0.0008 | 0.2 | -0.0001 | 0.1 | -0.0007 | 20.0 | 10.0 | 0.0 | 50.0010755E+00 | DC220MV |  | -0.00070010876146500006 | 4.1567173170447503e-08 | 1517778354.655 | 1517778370.168 | -0.00070008 | -0.00070009 | -0.0007001 | -0.00070014 | -0.00070011 | -0.00070012 | -0.00070004 | -0.00070011 | -0.00070014 | -0.00070013 | -0.00070012 | -0.00070011 | -0.0007001 | -0.00070004 | -0.0007002 | -0.00070013 | -0.00070004 | -0.00070018 | -0.00070009 | -0.00070012
| 0.1 | -0.0007 | 0.2 | 0.0 | 0.1 | -0.0007 | 20.0 | 10.0 | 0.0 | 50.0094027E+00 | DC220MV |  | -0.00070010668688999988 | 3.9263106080479108e-08 | 1517778385.026 | 1517778400.554 | -0.00070008 | -0.00070011 | -0.00070002 | -0.00070016 | -0.00070003 | -0.00070008 | -0.00070008 | -0.00070009 | -0.00070006 | -0.00070015 | -0.00070013 | -0.00070011 | -0.0007001 | -0.00070015 | -0.00070011 | -0.00070013 | -0.00070017 | -0.00070011 | -0.00070013 | -0.00070014
| 0.1 | -0.0007 | 0.2 | -0.0001 | 0.1 | -0.0006 | 20.0 | 10.0 | 0.0 | 49.9608232E+00 | DC220MV |  | -0.00060008343510999996 | 3.5294299683278414e-08 | 1517778415.412 | 1517778430.938 | -0.00060013 | -0.00060011 | -0.00060008 | -0.00060005 | -0.0006 | -0.00060009 | -0.00060007 | -0.00060009 | -0.00060007 | -0.00060006 | -0.00060006 | -0.00060013 | -0.00060007 | -0.00060004 | -0.00060013 | -0.00060006 | -0.00060014 | -0.00060011 | -0.00060005 | -0.00060012
| 0.1 | -0.0006 | 0.2 | 0.0 | 0.1 | -0.0006 | 20.0 | 10.0 | 0.0 | 50.0186296E+00 | DC220MV |  | -0.00060009094245500006 | 3.9989469015881241e-08 | 1517778445.783 | 1517778461.329 | -0.0006001 | -0.00060011 | -0.00060018 | -0.00060008 | -0.00060009 | -0.00060006 | -0.00060008 | -0.00060007 | -0.00060005 | -0.00059998 | -0.00060008 | -0.00060006 | -0.00060013 | -0.00060011 | -0.00060009 | -0.00060009 | -0.00060012 | -0.00060014 | -0.00060008 | -0.00060014
| 0.1 | -0.0006 | 0.2 | -0.0001 | 0.1 | -0.0004999999999999999 | 20.0 | 10.0 | 0.0 | 50.0244176E+00 | DC220MV |  | -0.00050006676541500001 | 3.1430265475534857e-08 | 1517778476.19 | 1517778491.68 | -0.00050007 | -0.00050013 | -0.00050004 | -0.00050001 | -0.00050006 | -0.00050008 | -0.00050006 | -0.00050009 | -0.00050008 | -0.00050007 | -0.00050009 | -0.00050001 | -0.00050008 | -0.00050009 | -0.00050004 | -0.00050008 | -0.00050001 | -0.00050005 | -0.00050012 | -0.00050009
| 0.1 | -0.0005 | 0.2 | 0.0 | 0.1 | -0.0005 | 20.0 | 10.0 | 0.0 | 50.0021605E+00 | DC220MV |  | -0.00050010893763500011 | 5.3832403844521013e-08 | 1517778506.524 | 1517778522.031 | -0.00050013 | -0.0005001 | -0.0005001 | -0.00050012 | -0.00050007 | -0.00050011 | -0.00050015 | -0.00050012 | -0.00050011 | -0.00050008 | -0.00050004 | -0.00050009 | -0.00050008 | -0.0005001 | -0.00050012 | -0.00050028 | -0.0005002 | -0.00050004 | -0.00050008 | -0.00050006
| 0.1 | -0.0005 | 0.2 | -0.0001 | 0.1 | -0.0004 | 20.0 | 10.0 | 0.0 | 49.9564813E+00 | DC220MV |  | -0.00040007767894000002 | 3.3175478118510766e-08 | 1517778536.892 | 1517778552.427 | -0.00040015 | -0.00040005 | -0.00040008 | -0.00040007 | -0.00040007 | -0.00040008 | -0.0004001 | -0.00040001 | -0.00040007 | -0.00040004 | -0.00040014 | -0.00040006 | -0.00040012 | -0.00040009 | -0.00040006 | -0.00040007 | -0.00040008 | -0.00040006 | -0.00040004 | -0.00040011
| 0.1 | -0.0004 | 0.2 | 0.0 | 0.1 | -0.0004 | 20.0 | 10.0 | 0.0 | 49.9981416E+00 | DC220MV |  | -0.00040009905138000003 | 3.8621702433566769e-08 | 1517778567.295 | 1517778582.852 | -0.00040011 | -0.00040009 | -0.00040011 | -0.00040009 | -0.00040018 | -0.00040013 | -0.00040006 | -0.00040006 | -0.00040014 | -0.00040014 | -0.00040012 | -0.00040008 | -0.00040006 | -0.00040017 | -0.00040007 | -0.00040002 | -0.00040009 | -0.0004001 | -0.0004001 | -0.00040007
| 0.1 | -0.0004 | 0.2 | -0.0001 | 0.1 | -0.0003 | 20.0 | 10.0 | 0.0 | 49.9813607E+00 | DC220MV |  | -0.00030008224619999997 | 3.4781062246283848e-08 | 1517778597.707 | 1517778613.227 | -0.0003001 | -0.0003001 | -0.00030011 | -0.00030006 | -0.00030008 | -0.00030003 | -0.00030002 | -0.00030011 | -0.00030004 | -0.00030012 | -0.00030013 | -0.00030009 | -0.00030013 | -0.00030013 | -0.00030005 | -0.00030009 | -0.0003001 | -0.00030008 | -0.00030004 | -0.00030004
| 0.1 | -0.0003 | 0.2 | 0.0 | 0.1 | -0.0003 | 20.0 | 10.0 | 0.0 | 50.0417235E+00 | DC220MV |  | -0.00030008744665000001 | 4.694849258681744e-08 | 1517778628.073 | 1517778643.611 | -0.00030018 | -0.00030008 | -0.00030006 | -0.00030006 | -0.00030004 | -0.00030007 | -0.00030013 | -0.00030008 | -0.00029996 | -0.00030012 | -0.00030008 | -0.00030007 | -0.00030012 | -0.00030018 | -0.00030006 | -0.0003001 | -0.0003001 | -0.00030009 | -0.00030004 | -0.00030013
| 0.1 | -0.0003 | 0.2 | -0.0001 | 0.1 | -0.0002 | 20.0 | 10.0 | 0.0 | 49.9758230E+00 | DC220MV |  | -0.00020008094772499995 | 4.8442630727127877e-08 | 1517778658.457 | 1517778673.969 | -0.00020011 | -0.00020005 | -0.00020016 | -0.00020008 | -0.00020007 | -0.0002001 | -0.00020009 | -0.00020013 | -0.00020002 | -0.00020004 | -0.00019999 | -0.00020011 | -0.00020012 | -0.00019997 | -0.0002001 | -0.00020011 | -0.00020015 | -0.00020008 | -0.00020007 | -0.00020008
| 0.1 | -0.0002 | 0.2 | 0.0 | 0.1 | -0.0002 | 20.0 | 10.0 | 0.0 | 49.9560387E+00 | DC220MV |  | -0.000200105103465 | 3.5061931352140357e-08 | 1517778688.829 | 1517778704.362 | -0.00020009 | -0.00020008 | -0.00020011 | -0.00020016 | -0.0002001 | -0.00020015 | -0.00020009 | -0.00020009 | -0.00020009 | -0.00020001 | -0.00020013 | -0.00020012 | -0.00020009 | -0.00020006 | -0.00020014 | -0.00020016 | -0.00020012 | -0.00020011 | -0.00020008 | -0.00020014
| 0.1 | -0.0002 | 0.2 | -0.0001 | 0.1 | -0.0001 | 20.0 | 10.0 | 0.0 | 50.0047174E+00 | DC220MV |  | -0.00010011040339000001 | 4.3292411845310589e-08 | 1517778719.209 | 1517778734.747 | -0.00010019 | -0.00010016 | -0.00010008 | -0.00010012 | -0.0001002 | -0.00010008 | -0.00010007 | -0.00010009 | -0.00010006 | -0.00010009 | -0.00010015 | -0.00010008 | -0.00010008 | -0.00010003 | -0.00010015 | -0.00010012 | -0.00010012 | -0.00010011 | -0.00010014 | -0.00010008
| 0.1 | -0.0001 | 0.2 | 0.0 | 0.1 | -0.0001 | 20.0 | 10.0 | 0.0 | 49.9944817E+00 | DC220MV |  | -0.00010008684305949998 | 5.9228899816546277e-08 | 1517778749.589 | 1517778765.108 | -0.00010011 | -0.00010008 | -0.00010018 | -0.00010009 | -9.9933e-05 | -0.00010009 | -0.00010002 | -0.00010009 | -0.00010006 | -0.00010012 | -0.00010014 | -0.00010009 | -0.00010009 | -0.00010009 | -0.00010007 | -0.00010003 | -0.00010014 | -0.00010003 | -0.00010022 | -0.00010008
| 0.1 | -0.0001 | 0.2 | -0.0001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0344530E+00 | DC220MV |  | -6.3582299113849995e-08 | 4.3437894656907278e-08 | 1517778779.959 | 1517778795.525 | -5.7802e-08 | -6.4126e-08 | -7.9423e-08 | -8.4718e-08 | -1.2766e-07 | 3.2357e-09 | -6.5597e-08 | -1.081e-07 | -5.7949e-08 | -7.1186e-08 | -4.0153e-08 | -3.4711e-08 | -1.2546e-07 | -3.4564e-08 | 5.589e-08 | -1.0222e-07 | -2.7357e-08 | -7.4569e-08 | -5.6331e-08 | -1.1884e-07
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9521590E+00 | DC220MV |  | -1.0763662423815e-07 | 4.1429627282465409e-08 | 1517778810.383 | 1517778825.927 | -1.0642e-07 | -1.9283e-07 | -5.7259e-08 | -8.1105e-08 | -1.2217e-07 | -1.232e-07 | -1.2394e-07 | -1.7678e-07 | -1.126e-07 | -8.5079e-08 | -7.1831e-08 | -1.5323e-07 | -8.9936e-08 | -9.656e-08 | -1.0775e-07 | -1.0745e-07 | 2.3551e-09 | -1.3792e-07 | -8.2577e-08 | -1.2644e-07
| 0.1 | -0.0001 | 0.2 | -0.0001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0148912E+00 | DC220MV |  | -8.2800814397500004e-08 | 4.8763695305063314e-08 | 1517778840.76 | 1517778856.342 | -9.8044e-08 | -7.8935e-08 | -6.4383e-08 | -6.747e-08 | -1.6654e-07 | -9.1429e-08 | -1.2465e-07 | -5.9679e-08 | -3.1897e-08 | -5.0565e-08 | -5.4975e-08 | -7.8788e-08 | -1.2568e-07 | 1.3817e-08 | -2.1784e-07 | -1.0275e-07 | -5.1888e-08 | -7.2467e-08 | -3.8512e-08 | -9.334e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9610068E+00 | DC220MV |  | -7.565287382005001e-08 | 4.1995323907893535e-08 | 1517778871.2 | 1517778886.758 | -1.3493e-07 | -1.1359e-07 | -1.5053e-07 | -7.3718e-08 | -2.4131e-08 | -7.2835e-08 | -4.7968e-08 | -1.0167e-07 | -4.326e-08 | -5.2235e-08 | -6.7391e-08 | -2.4278e-08 | -4.2818e-08 | -6.0181e-08 | -8.1369e-08 | -5.9445e-08 | -1.7113e-07 | -8.6372e-08 | -5.7385e-09 | -9.9468e-08
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 50.0047127E+00 | DC220MV |  | -0.0100000348969 | 4.83719306405779e-08 | 1517778901.607 | 1517778917.129 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.0099999 | -0.01 | -0.01 | -0.01
| 0.1 | 0.001 | 0.2 | 0.0 | 0.1 | 0.001 | 20.0 | 10.0 | 0.0 | 49.9414178E+00 | DC220MV |  | 0.00099992556788 | 3.5945870916886981e-08 | 1517778931.988 | 1517778947.514 | 0.00099992 | 0.00099991 | 0.00099993 | 0.00099995 | 0.00099998 | 0.0009999 | 0.00099994 | 0.00099996 | 0.00099998 | 0.00099993 | 0.0009999 | 0.00099995 | 0.00099987 | 0.00099995 | 0.00099983 | 0.00099996 | 0.00099994 | 0.00099988 | 0.00099991 | 0.00099993
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9970096E+00 | DC220MV |  | -8.9058714454499984e-08 | 2.77894539165886e-08 | 1517778962.352 | 1517778977.942 | -4.2465e-08 | -1.0668e-07 | -6.2301e-08 | -6.7738e-08 | -1.0139e-07 | -1.0682e-07 | -1.3841e-07 | -4.9518e-08 | -6.6563e-08 | -1.0197e-07 | -9.9917e-08 | -6.9061e-08 | -1.3077e-07 | -6.5828e-08 | -1.2798e-07 | -7.2734e-08 | -7.8024e-08 | -1.1946e-07 | -1.0932e-07 | -6.4212e-08
| 0.1 | 0.001 | 0.2 | 0.001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0167887E+00 | DC220MV |  | -8.6294353276500011e-08 | 2.8821031743273384e-08 | 1517778992.799 | 1517779008.367 | -9.7656e-08 | -1.3325e-07 | -7.883e-08 | -8.839e-08 | -5.8976e-08 | -1.7796e-08 | -8.5743e-08 | -8.5155e-08 | -1.2531e-07 | -8.839e-08 | -1.4016e-07 | -8.8684e-08 | -5.0004e-08 | -7.5595e-08 | -7.9125e-08 | -1.153e-07 | -6.7506e-08 | -9.1037e-08 | -1.0736e-07 | -5.1622e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0061517E+00 | DC220MV |  | -8.0416072987000013e-08 | 4.1106257379708357e-08 | 1517779023.189 | 1517779038.753 | -8.6205e-08 | -9.0177e-08 | -1.0768e-07 | -1.3343e-07 | -8.5175e-08 | -9.1353e-08 | -1.5226e-07 | 3.2658e-08 | -8.1644e-08 | -7.3259e-08 | -1.2092e-07 | -7.0023e-08 | -1.3843e-07 | -7.9585e-08 | -9.003e-08 | -8.1939e-08 | -3.9278e-08 | -3.0451e-08 | -4.3397e-08 | -4.575e-08
| 0.1 | 0.001 | 0.2 | 0.001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9638187E+00 | DC220MV |  | -7.2572000758e-08 | 6.9873130067311605e-08 | 1517779053.597 | 1517779069.162 | -1.3239e-07 | 2.3683e-08 | -2.5154e-08 | -1.0724e-07 | -3.4716e-08 | -5.1191e-08 | -1.1474e-08 | 1.1621e-08 | -2.1624e-07 | -5.3545e-08 | -1.0591e-08 | -6.2223e-08 | -3.4863e-08 | -2.6596e-07 | -4.1482e-08 | -1.0488e-07 | -5.9281e-08 | -9.6939e-08 | -6.1635e-08 | -1.1694e-07
| 0.1 | 0.001 | 0.2 | 0.0 | 0.1 | 0.001 | 20.0 | 10.0 | 0.0 | 50.0031370E+00 | DC220MV |  | 0.00099993074806000023 | 5.4876878595611946e-08 | 1517779084.004 | 1517779099.561 | 0.0009999 | 0.00099991 | 0.00099993 | 0.00099988 | 0.00099997 | 0.00099989 | 0.00099988 | 0.00099994 | 0.00099986 | 0.00099988 | 0.00099996 | 0.00099995 | 0.00099998 | 0.001 | 0.00099993 | 0.001 | 0.00099983 | 0.00099994 | 0.00099994 | 0.00099994
| 0.1 | 0.002 | 0.2 | 0.001 | 0.1 | 0.001 | 20.0 | 10.0 | 0.0 | 49.9842785E+00 | DC220MV |  | 0.00099991626185000002 | 5.6863103261573675e-08 | 1517779114.404 | 1517779129.93 | 0.00099991 | 0.0009999 | 0.00099994 | 0.00099989 | 0.00099992 | 0.00099988 | 0.00099993 | 0.00099986 | 0.001 | 0.00099995 | 0.00099973 | 0.00099992 | 0.00099991 | 0.00099992 | 0.0009999 | 0.00099991 | 0.001 | 0.00099997 | 0.00099989 | 0.00099998
| 0.1 | 0.002 | 0.2 | 0.0 | 0.1 | 0.002 | 20.0 | 10.0 | 0.0 | 49.9614637E+00 | DC220MV |  | 0.0019999364539999998 | 5.4152528849547355e-08 | 1517779144.786 | 1517779160.286 | 0.002 | 0.0019999 | 0.0019999 | 0.0019999 | 0.0019998 | 0.002 | 0.0019999 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.002 | 0.002 | 0.0019998 | 0.0019999 | 0.0019999 | 0.002 | 0.002 | 0.0019999
| 0.1 | 0.003 | 0.2 | 0.001 | 0.1 | 0.002 | 20.0 | 10.0 | 0.0 | 50.0304431E+00 | DC220MV |  | 0.00199992636415 | 3.1814224408677225e-08 | 1517779175.134 | 1517779190.659 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.0019999 | 0.0019999
| 0.1 | 0.003 | 0.2 | 0.0 | 0.1 | 0.003 | 20.0 | 10.0 | 0.0 | 49.9541364E+00 | DC220MV |  | 0.0029999418588500001 | 3.876063274288087e-08 | 1517779205.505 | 1517779221.023 | 0.003 | 0.0029999 | 0.0029999 | 0.0029999 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.003 | 0.003 | 0.0029999 | 0.003 | 0.0029999 | 0.0029999 | 0.0029999
| 0.1 | 0.004 | 0.2 | 0.001 | 0.1 | 0.003 | 20.0 | 10.0 | 0.0 | 49.9872957E+00 | DC220MV |  | 0.00299993038985 | 4.4732606904028518e-08 | 1517779235.877 | 1517779251.416 | 0.0029999 | 0.003 | 0.003 | 0.003 | 0.0029999 | 0.0029999 | 0.003 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.003 | 0.0029999 | 0.003 | 0.0029998 | 0.0029999 | 0.003 | 0.0029999 | 0.003 | 0.0029999
| 0.1 | 0.004 | 0.2 | 0.0 | 0.1 | 0.004 | 20.0 | 10.0 | 0.0 | 50.0583038E+00 | DC220MV |  | 0.0039999649997500006 | 4.7575552742816149e-08 | 1517779266.278 | 1517779281.805 | 0.0039999 | 0.004 | 0.004 | 0.0039999 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.004 | 0.0039999 | 0.0039999 | 0.004 | 0.004 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0040001
| 0.1 | 0.005 | 0.2 | 0.001 | 0.1 | 0.004 | 20.0 | 10.0 | 0.0 | 50.0546493E+00 | DC220MV |  | 0.0039999098518499997 | 4.1805295074047641e-08 | 1517779296.65 | 1517779312.154 | 0.004 | 0.004 | 0.0039999 | 0.004 | 0.0039999 | 0.0039999 | 0.0039999 | 0.004 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039998
| 0.1 | 0.005 | 0.2 | 0.0 | 0.1 | 0.005 | 20.0 | 10.0 | 0.0 | 50.0238699E+00 | DC220MV |  | 0.0049999253890000003 | 4.1872744816851176e-08 | 1517779327.006 | 1517779342.493 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049998 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999
| 0.1 | 0.006 | 0.2 | 0.001 | 0.1 | 0.005 | 20.0 | 10.0 | 0.0 | 49.9802814E+00 | DC220MV |  | 0.0049999016988999997 | 3.7641501392634847e-08 | 1517779357.341 | 1517779372.863 | 0.005 | 0.0049999 | 0.0049999 | 0.005 | 0.0049998 | 0.0049998 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999
| 0.1 | 0.006 | 0.2 | 0.0 | 0.1 | 0.006 | 20.0 | 10.0 | 0.0 | 49.9652317E+00 | DC220MV |  | 0.0059998964435500008 | 2.548208645991172e-08 | 1517779387.709 | 1517779403.233 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999
| 0.1 | 0.007 | 0.2 | 0.001 | 0.1 | 0.006 | 20.0 | 10.0 | 0.0 | 50.0334988E+00 | DC220MV |  | 0.0059998894498499997 | 4.1742820121774923e-08 | 1517779418.089 | 1517779433.639 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059998 | 0.0059998 | 0.0059999 | 0.0059999 | 0.0059998 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999
| 0.1 | 0.007 | 0.2 | 0.0 | 0.1 | 0.007 | 20.0 | 10.0 | 0.0 | 50.0164299E+00 | DC220MV |  | 0.0069999137534500003 | 3.4930624036116323e-08 | 1517779448.497 | 1517779464.003 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.007 | 0.007 | 0.0069999 | 0.0069999 | 0.0069998 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999
| 0.1 | 0.008 | 0.2 | 0.001 | 0.1 | 0.007 | 20.0 | 10.0 | 0.0 | 50.0435426E+00 | DC220MV |  | 0.0069999079378500004 | 4.1064231359175309e-08 | 1517779478.862 | 1517779494.4 | 0.0069999 | 0.0069999 | 0.0069999 | 0.007 | 0.0069999 | 0.007 | 0.0069999 | 0.0069998 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.007 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999
| 0.1 | 0.008 | 0.2 | 0.0 | 0.1 | 0.008 | 20.0 | 10.0 | 0.0 | 49.9982548E+00 | DC220MV |  | 0.0079999040672000012 | 5.5955932523200002e-08 | 1517779509.238 | 1517779524.734 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.008 | 0.0079999 | 0.0079999 | 0.0079998 | 0.008 | 0.0079999 | 0.0079998 | 0.0079999 | 0.0079999 | 0.008 | 0.0079998 | 0.008 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079998
| 0.1 | 0.009 | 0.2 | 0.001 | 0.1 | 0.008 | 20.0 | 10.0 | 0.0 | 49.9939536E+00 | DC220MV |  | 0.0079999227934000015 | 6.6873449773182269e-08 | 1517779539.584 | 1517779555.094 | 0.0079999 | 0.0079999 | 0.008 | 0.0079999 | 0.0079999 | 0.0079999 | 0.008 | 0.0079999 | 0.0079999 | 0.008 | 0.0079999 | 0.0079999 | 0.0080001 | 0.0079998 | 0.0079999 | 0.0080001 | 0.0079999 | 0.0079998 | 0.0079999 | 0.0079999
| 0.1 | 0.009 | 0.2 | 0.0 | 0.1 | 0.009 | 20.0 | 10.0 | 0.0 | 49.9382321E+00 | DC220MV |  | 0.0089999091758499981 | 4.2236849315838844e-08 | 1517779569.927 | 1517779585.457 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089998 | 0.009 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089998 | 0.009 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.009 | 0.0089999 | 0.009 | 0.0089999
| 0.1 | 0.01 | 0.2 | 0.001 | 0.1 | 0.009000000000000001 | 20.0 | 10.0 | 0.0 | 50.0492324E+00 | DC220MV |  | 0.0089999073636999986 | 4.9507132204438532e-08 | 1517779600.316 | 1517779615.861 | 0.0090001 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.009 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.009 | 0.0089999
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0224629E+00 | DC220MV |  | 0.0099998942095999976 | 4.1354717796599766e-08 | 1517779630.718 | 1517779646.226 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099998 | 0.0099998 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.0099999
| 0.1 | 0.01 | 0.2 | 0.001 | 0.1 | 0.009000000000000001 | 20.0 | 10.0 | 0.0 | 49.9928171E+00 | DC220MV |  | 0.0089999002751499999 | 5.6639572383927734e-08 | 1517779661.078 | 1517779676.589 | 0.0089999 | 0.0089999 | 0.009 | 0.009 | 0.0089999 | 0.009 | 0.0089999 | 0.0089999 | 0.0089998 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089998 | 0.0089998 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999
| 0.1 | 0.009 | 0.2 | 0.0 | 0.1 | 0.009 | 20.0 | 10.0 | 0.0 | 50.0192857E+00 | DC220MV |  | 0.0089999239469499993 | 4.2885511576262988e-08 | 1517779691.47 | 1517779707 | 0.0089999 | 0.009 | 0.0089999 | 0.0089999 | 0.009 | 0.009 | 0.0089999 | 0.009 | 0.0089998 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.0089999 | 0.009 | 0.0089999 | 0.0089999 | 0.0089999 | 0.009 | 0.0089999
| 0.1 | 0.009 | 0.2 | 0.001 | 0.1 | 0.008 | 20.0 | 10.0 | 0.0 | 49.9982784E+00 | DC220MV |  | 0.0079999008594499998 | 3.0173020512608711e-08 | 1517779721.846 | 1517779737.359 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079998 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.008 | 0.0079999 | 0.0079999
| 0.1 | 0.008 | 0.2 | 0.0 | 0.1 | 0.008 | 20.0 | 10.0 | 0.0 | 50.0093131E+00 | DC220MV |  | 0.0079998836402999998 | 5.8062252791315115e-08 | 1517779752.239 | 1517779767.765 | 0.0079998 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.008 | 0.0079999 | 0.0079999 | 0.0079999 | 0.0079999 | 0.008 | 0.0079997 | 0.0079998 | 0.0079999 | 0.0079999 | 0.0079998 | 0.0079999 | 0.0079998
| 0.1 | 0.008 | 0.2 | 0.001 | 0.1 | 0.007 | 20.0 | 10.0 | 0.0 | 49.9579740E+00 | DC220MV |  | 0.0069999189637499999 | 4.1779610482617531e-08 | 1517779782.616 | 1517779798.144 | 0.007 | 0.007 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069998 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.007 | 0.007 | 0.007 | 0.0069999 | 0.0069998 | 0.0069999 | 0.0069999 | 0.0069999
| 0.1 | 0.007 | 0.2 | 0.0 | 0.1 | 0.007 | 20.0 | 10.0 | 0.0 | 49.9494992E+00 | DC220MV |  | 0.0069999307810999986 | 3.8253031024756952e-08 | 1517779812.977 | 1517779828.547 | 0.0069999 | 0.0069999 | 0.007 | 0.0069999 | 0.0069999 | 0.0069999 | 0.0069999 | 0.007 | 0.0069999 | 0.007 | 0.0069999 | 0.007 | 0.0069999 | 0.0069999 | 0.007 | 0.0069999 | 0.0069999 | 0.0069999 | 0.007 | 0.0069999
| 0.1 | 0.007 | 0.2 | 0.001 | 0.1 | 0.006 | 20.0 | 10.0 | 0.0 | 50.0471198E+00 | DC220MV |  | 0.0059999081203499999 | 6.10312221475383e-08 | 1517779843.375 | 1517779858.937 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.006 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059998 | 0.006 | 0.006 | 0.006 | 0.006 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059997
| 0.1 | 0.006 | 0.2 | 0.0 | 0.1 | 0.006 | 20.0 | 10.0 | 0.0 | 50.0334752E+00 | DC220MV |  | 0.0059999361151499998 | 4.112447624873876e-08 | 1517779873.793 | 1517779889.294 | 0.0059999 | 0.0059999 | 0.0059999 | 0.006 | 0.006 | 0.0059999 | 0.0059999 | 0.0059999 | 0.006 | 0.006 | 0.0059999 | 0.006 | 0.0059999 | 0.0059999 | 0.006 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999 | 0.0059999
| 0.1 | 0.006 | 0.2 | 0.001 | 0.1 | 0.005 | 20.0 | 10.0 | 0.0 | 49.9667627E+00 | DC220MV |  | 0.0049999244473999992 | 4.3382292337266032e-08 | 1517779904.147 | 1517779919.658 | 0.0049999 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049998 | 0.005 | 0.0049999 | 0.0049999 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.005 | 0.005
| 0.1 | 0.005 | 0.2 | 0.0 | 0.1 | 0.005 | 20.0 | 10.0 | 0.0 | 49.9128644E+00 | DC220MV |  | 0.0049999329850000003 | 5.0824028378431856e-08 | 1517779934.511 | 1517779950.056 | 0.005 | 0.0049999 | 0.005 | 0.005 | 0.0049999 | 0.0049999 | 0.005 | 0.0049998 | 0.0049999 | 0.0049999 | 0.0049999 | 0.0049999 | 0.005 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999 | 0.005 | 0.0049999 | 0.0049999
| 0.1 | 0.005 | 0.2 | 0.001 | 0.1 | 0.004 | 20.0 | 10.0 | 0.0 | 49.9806254E+00 | DC220MV |  | 0.0039999331877500002 | 5.978420672043316e-08 | 1517779964.913 | 1517779980.476 | 0.004 | 0.004 | 0.0039999 | 0.0039998 | 0.0040001 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.0039999 | 0.004 | 0.0039999 | 0.004 | 0.0039999 | 0.0039999 | 0.004
| 0.1 | 0.004 | 0.2 | 0.0 | 0.1 | 0.004 | 20.0 | 10.0 | 0.0 | 50.0579302E+00 | DC220MV |  | 0.0039999352669499999 | 6.027101291547423e-08 | 1517779995.306 | 1517780010.846 | 0.0039999 | 0.0039999 | 0.004 | 0.0039999 | 0.004 | 0.0039999 | 0.0039999 | 0.004 | 0.0039999 | 0.0040001 | 0.0039999 | 0.004 | 0.004 | 0.0039999 | 0.0039999 | 0.0039999 | 0.004 | 0.0039999 | 0.0039999 | 0.0039999
| 0.1 | 0.004 | 0.2 | 0.001 | 0.1 | 0.003 | 20.0 | 10.0 | 0.0 | 50.0172324E+00 | DC220MV |  | 0.0029999241080500001 | 3.4679991788471972e-08 | 1517780025.701 | 1517780041.187 | 0.003 | 0.0029999 | 0.0029999 | 0.003 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.003 | 0.003 | 0.0029999 | 0.0029999
| 0.1 | 0.003 | 0.2 | 0.0 | 0.1 | 0.003 | 20.0 | 10.0 | 0.0 | 49.9692312E+00 | DC220MV |  | 0.0029999125102999997 | 4.1304496277157554e-08 | 1517780056.035 | 1517780071.547 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.003 | 0.0029999 | 0.003 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029998 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999 | 0.0029999
| 0.1 | 0.003 | 0.2 | 0.001 | 0.1 | 0.002 | 20.0 | 10.0 | 0.0 | 50.0504943E+00 | DC220MV |  | 0.00199992822115 | 4.0417564846572438e-08 | 1517780086.4 | 1517780101.947 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.0019999 | 0.0019999 | 0.0019999 | 0.002 | 0.002 | 0.0019999 | 0.002 | 0.002 | 0.0019999
| 0.1 | 0.002 | 0.2 | 0.0 | 0.1 | 0.002 | 20.0 | 10.0 | 0.0 | 50.0624316E+00 | DC220MV |  | 0.00199993265285 | 4.6097629887334998e-08 | 1517780116.795 | 1517780132.292 | 0.002 | 0.0019999 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999 | 0.0019999 | 0.0019998 | 0.0019999 | 0.002 | 0.002 | 0.0019999 | 0.0019999 | 0.0019998 | 0.002 | 0.0019999 | 0.0019999 | 0.0019999 | 0.002 | 0.0019999
| 0.1 | 0.002 | 0.2 | 0.001 | 0.1 | 0.001 | 20.0 | 10.0 | 0.0 | 49.9778636E+00 | DC220MV |  | 0.0009999475300799998 | 6.1494377179966879e-08 | 1517780147.143 | 1517780162.66 | 0.0009999 | 0.00099996 | 0.00099989 | 0.00099996 | 0.00099995 | 0.00099997 | 0.00099993 | 0.00099989 | 0.00099994 | 0.00099991 | 0.001 | 0.00099995 | 0.00099995 | 0.00099994 | 0.0010002 | 0.00099999 | 0.00099987 | 0.00099994 | 0.00099993 | 0.0009999
| 0.1 | 0.001 | 0.2 | 0.0 | 0.1 | 0.001 | 20.0 | 10.0 | 0.0 | 50.0093414E+00 | DC220MV |  | 0.00099994785416000027 | 4.8598102481933307e-08 | 1517780177.51 | 1517780193.048 | 0.00099987 | 0.0009999 | 0.00099996 | 0.00099992 | 0.00099987 | 0.00099994 | 0.00099996 | 0.00099993 | 0.00099998 | 0.00099998 | 0.00099999 | 0.00099995 | 0.00099993 | 0.0010001 | 0.00099993 | 0.00099997 | 0.00099995 | 0.00099995 | 0.0009999 | 0.00099998
| 0.1 | 0.001 | 0.2 | 0.001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9686423E+00 | DC220MV |  | -7.0550727621850002e-08 | 5.1989858483414096e-08 | 1517780207.897 | 1517780223.464 | 1.4856e-08 | -1.3017e-07 | -3.1624e-08 | -1.3547e-07 | -2.1916e-08 | -8.5164e-08 | -3.4566e-08 | -2.1504e-07 | -3.8096e-08 | 1.1767e-09 | -1.0443e-07 | -6.3983e-08 | -2.1328e-08 | -8.8988e-08 | -1.0031e-07 | -7.6044e-08 | -7.7662e-08 | -4.5597e-08 | -9.9137e-08 | -5.7511e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9996887E+00 | DC220MV |  | -4.8847658702250001e-08 | 3.074051724281309e-08 | 1517780238.315 | 1517780253.884 | -8.4134e-08 | -4.4126e-08 | -6.9866e-08 | 1.1032e-08 | -3.8243e-08 | -1.0252e-07 | -5.9276e-08 | -4.898e-08 | -4.2067e-08 | -5.8835e-09 | -6.31e-08 | -2.9417e-09 | -4.4715e-08 | -1.1473e-08 | -9.5313e-08 | -4.8539e-08 | -9.7078e-08 | -4.1332e-08 | -3.1918e-08 | -5.6481e-08
| 0.1 | 0.001 | 0.2 | 0.001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9908602E+00 | DC220MV |  | -6.721086195235001e-08 | 2.685731597980173e-08 | 1517780268.75 | 1517780284.315 | -1.0209e-07 | -4.2365e-08 | -1.0547e-07 | -8.9291e-08 | -5.5752e-08 | -7.0609e-08 | -4.972e-08 | -3.3686e-08 | -1.1062e-07 | -7.5905e-08 | -5.8841e-09 | -4.2512e-08 | -5.3545e-08 | -6.3842e-08 | -7.2815e-08 | -6.0312e-08 | -1.1209e-07 | -7.7376e-08 | -7.1197e-08 | -4.9132e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0045617E+00 | DC220MV |  | -6.4775164617800006e-08 | 3.4890816284139491e-08 | 1517780299.169 | 1517780314.739 | -4.0443e-08 | -1.4854e-08 | -1.0265e-07 | -7.3092e-08 | -1.1265e-07 | -6.265e-08 | -8.8828e-08 | -8.2063e-08 | -3.0884e-09 | -7.1768e-08 | -6.6327e-08 | -7.4562e-08 | -6.4856e-08 | 6.1768e-09 | -9.1034e-08 | -4.2355e-08 | -6.0003e-08 | -1.4251e-07 | -6.1326e-08 | -4.662e-08
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 49.9853863E+00 | DC220MV |  | 0.0099999107898500004 | 2.3625157328918294e-08 | 1517780329.586 | 1517780345.103 | 0.0099999 | 0.0099999 | 0.01 | 0.01 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 50.0520494E+00 | DC220MV |  | -0.0100000466925 | 4.3552284549056686e-08 | 1517780359.963 | 1517780375.486 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9944346E+00 | DC220MV |  | -6.5961948670000006e-08 | 6.6993342975903696e-08 | 1517780390.324 | 1517780405.855 | -3.5058e-08 | -7.8659e-08 | -5.3323e-08 | -7.9543e-08 | -7.5713e-08 | -1.8118e-08 | -6.7464e-08 | -9.1475e-08 | -1.1902e-07 | -4.0508e-08 | -6.2162e-08 | 6.4813e-08 | -5.0525e-08 | -7.9691e-08 | 7.645e-08 | -2.5984e-07 | -5.2145e-08 | -1.2005e-07 | -1.2211e-07 | -5.5091e-08
| 0.1 | -0.001 | 0.2 | -0.001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9816435E+00 | DC220MV |  | -4.7352577882699997e-08 | 3.7570931742949645e-08 | 1517780420.707 | 1517780436.271 | -5.8841e-08 | -4.8691e-08 | -2.9421e-09 | -1.0106e-07 | -4.0601e-08 | -2.7803e-08 | -5.5752e-08 | -3.0303e-08 | -1.471e-09 | 1.0297e-08 | -3.251e-08 | -2.1771e-08 | -4.4131e-09 | -5.9577e-08 | -9.9001e-08 | -5.4134e-08 | -4.1189e-08 | -1.4696e-07 | -8.5761e-08 | -4.4572e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9673138E+00 | DC220MV |  | -4.379227696145e-08 | 3.6271290712078266e-08 | 1517780451.125 | 1517780466.702 | -6.5291e-08 | 6.3233e-09 | -9.7937e-08 | -6.2791e-08 | -4.5733e-08 | -1.047e-07 | -7.5732e-08 | -7.882e-08 | -3.0146e-08 | -5.9997e-08 | -1.1029e-08 | -4.5586e-09 | -4.8086e-08 | 1.3823e-08 | -6.2497e-08 | -7.8379e-08 | 1.7793e-08 | -1.3088e-08 | -6.3233e-08 | -1.1764e-08
| 0.1 | -0.001 | 0.2 | -0.001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0163969E+00 | DC220MV |  | -1.6634599532239996e-08 | 4.8276290536489918e-08 | 1517780481.569 | 1517780497.153 | -3.8224e-09 | -4.2634e-08 | -5.6747e-08 | -7.2037e-08 | -5.7923e-08 | -7.7917e-09 | 3.5283e-08 | -1.4848e-08 | -2.955e-08 | -2.6904e-08 | -1.3158e-07 | 4.7926e-08 | 5.2778e-08 | 3.9253e-08 | -5.2925e-09 | 1.1761e-09 | -8.2475e-08 | -3.5136e-08 | -8.8208e-10 | 5.8511e-08
| 0.1 | -0.001 | 0.2 | 0.0 | 0.1 | -0.001 | 20.0 | 10.0 | 0.0 | 49.9924068E+00 | DC220MV |  | -0.0010000404738049998 | 4.8604960144816116e-08 | 1517780512.03 | 1517780527.543 | -0.001 | -0.0010001 | -0.001 | -0.00099998 | -0.001 | -0.001 | -0.001 | -0.0010001 | -0.00099999 | -0.001 | -0.0010001 | -0.001 | -0.001 | -0.0010001 | -0.0010001 | -0.00099999 | -0.001 | -0.00099999 | -0.0010001 | -0.0010002
| 0.1 | -0.002 | 0.2 | -0.001 | 0.1 | -0.001 | 20.0 | 10.0 | 0.0 | 49.9685999E+00 | DC220MV |  | -0.0010000325136999999 | 3.9444369777816499e-08 | 1517780542.391 | 1517780557.911 | -0.001 | -0.001 | -0.00099996 | -0.001 | -0.00099999 | -0.001 | -0.0010001 | -0.001 | -0.001 | -0.001 | -0.001 | -0.001 | -0.0010001 | -0.0010001 | -0.001 | -0.0010001 | -0.00099995 | -0.0010001 | -0.0010001 | -0.0010001
| 0.1 | -0.002 | 0.2 | 0.0 | 0.1 | -0.002 | 20.0 | 10.0 | 0.0 | 50.0039012E+00 | DC220MV |  | -0.0020000626063500001 | 4.9732009140272644e-08 | 1517780572.758 | 1517780588.298 | -0.0020001 | -0.0020001 | -0.002 | -0.0020001 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.0020001 | -0.0020001 | -0.0020001 | -0.002 | -0.0020001 | -0.0020001 | -0.0020001 | -0.0020001 | -0.0020001 | -0.002 | -0.002 | -0.002
| 0.1 | -0.003 | 0.2 | -0.001 | 0.1 | -0.002 | 20.0 | 10.0 | 0.0 | 49.9827984E+00 | DC220MV |  | -0.0020000445122000002 | 3.6709889785454377e-08 | 1517780603.13 | 1517780618.637 | -0.002 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.002 | -0.002 | -0.002 | -0.002 | -0.0020001 | -0.0020001 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.0020001 | -0.0020001 | -0.0020001 | -0.0020001 | -0.002
| 0.1 | -0.003 | 0.2 | 0.0 | 0.1 | -0.003 | 20.0 | 10.0 | 0.0 | 50.0050996E+00 | DC220MV |  | -0.0030000238020499999 | 4.3985953229922391e-08 | 1517780633.486 | 1517780649.005 | -0.003 | -0.003 | -0.003 | -0.0029999 | -0.003 | -0.0030001 | -0.003 | -0.003 | -0.003 | -0.0030001 | -0.003 | -0.003 | -0.003 | -0.003 | -0.0030001 | -0.0030001 | -0.0030001 | -0.003 | -0.003 | -0.003
| 0.1 | -0.004 | 0.2 | -0.001 | 0.1 | -0.003 | 20.0 | 10.0 | 0.0 | 49.9791031E+00 | DC220MV |  | -0.0030000152083999999 | 3.3128454809124745e-08 | 1517780663.854 | 1517780679.369 | -0.003 | -0.0030001 | -0.003 | -0.003 | -0.003 | -0.003 | -0.003 | -0.003 | -0.0030001 | -0.003 | -0.0030001 | -0.003 | -0.0030001 | -0.003 | -0.003 | -0.003 | -0.003 | -0.003 | -0.0030001 | -0.003
| 0.1 | -0.004 | 0.2 | 0.0 | 0.1 | -0.004 | 20.0 | 10.0 | 0.0 | 50.0129845E+00 | DC220MV |  | -0.0040000361797499998 | 3.7112524117756105e-08 | 1517780694.228 | 1517780709.749 | -0.004 | -0.0040001 | -0.0040001 | -0.004 | -0.004 | -0.0040001 | -0.0040001 | -0.004 | -0.0040001 | -0.004 | -0.004 | -0.0040001 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004
| 0.1 | -0.005 | 0.2 | -0.001 | 0.1 | -0.004 | 20.0 | 10.0 | 0.0 | 49.9978586E+00 | DC220MV |  | -0.0040000172573499996 | 4.7007215090026683e-08 | 1517780724.613 | 1517780740.122 | -0.004 | -0.0040001 | -0.004 | -0.0040001 | -0.004 | -0.0040001 | -0.004 | -0.0040001 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.0039999 | -0.0040001
| 0.1 | -0.005 | 0.2 | 0.0 | 0.1 | -0.005 | 20.0 | 10.0 | 0.0 | 49.9925671E+00 | DC220MV |  | -0.0050000431760499993 | 3.6494826725771037e-08 | 1517780754.973 | 1517780770.479 | -0.0049999 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.0050001 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.005 | -0.0050001 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.005
| 0.1 | -0.006 | 0.2 | -0.001 | 0.1 | -0.005 | 20.0 | 10.0 | 0.0 | 49.9985897E+00 | DC220MV |  | -0.0050000191510500003 | 4.201359122892739e-08 | 1517780785.337 | 1517780800.856 | -0.0050001 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.0050001 | -0.0049999 | -0.0050001 | -0.005 | -0.005 | -0.005 | -0.005 | -0.005 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.005 | -0.005 | -0.005
| 0.1 | -0.006 | 0.2 | 0.0 | 0.1 | -0.006 | 20.0 | 10.0 | 0.0 | 49.9991274E+00 | DC220MV |  | -0.0060000299052500004 | 5.0369778683109688e-08 | 1517780815.716 | 1517780831.238 | -0.0060001 | -0.006 | -0.006 | -0.006 | -0.0060001 | -0.006 | -0.006 | -0.0060001 | -0.006 | -0.006 | -0.0060001 | -0.006 | -0.006 | -0.006 | -0.006 | -0.0060001 | -0.0059999 | -0.0059999 | -0.006 | -0.0060001
| 0.1 | -0.007 | 0.2 | -0.001 | 0.1 | -0.006 | 20.0 | 10.0 | 0.0 | 50.0022926E+00 | DC220MV |  | -0.0060000151592000004 | 4.6450019710987728e-08 | 1517780846.097 | 1517780861.619 | -0.006 | -0.006 | -0.006 | -0.006 | -0.006 | -0.0060002 | -0.006 | -0.006 | -0.0060001 | -0.006 | -0.006 | -0.006 | -0.006 | -0.006 | -0.0060001 | -0.006 | -0.006 | -0.0060001 | -0.006 | -0.006
| 0.1 | -0.007 | 0.2 | 0.0 | 0.1 | -0.007 | 20.0 | 10.0 | 0.0 | 49.9499746E+00 | DC220MV |  | -0.0070000279297499999 | 4.0447781398779537e-08 | 1517780876.496 | 1517780892.004 | -0.007 | -0.007 | -0.0070001 | -0.0069999 | -0.007 | -0.007 | -0.0070001 | -0.0070001 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.0070001 | -0.007 | -0.007 | -0.007 | -0.0070001
| 0.1 | -0.008 | 0.2 | -0.001 | 0.1 | -0.007 | 20.0 | 10.0 | 0.0 | 50.0257019E+00 | DC220MV |  | -0.0070000283736000009 | 3.7497712357337833e-08 | 1517780906.844 | 1517780922.374 | -0.007 | -0.0070001 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.0070001 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.0070001 | -0.0070001
| 0.1 | -0.008 | 0.2 | 0.0 | 0.1 | -0.008 | 20.0 | 10.0 | 0.0 | 49.9920437E+00 | DC220MV |  | -0.0080000352846000006 | 3.6416010570222974e-08 | 1517780937.228 | 1517780952.718 | -0.008 | -0.008 | -0.0080001 | -0.008 | -0.008 | -0.008 | -0.008 | -0.008 | -0.008 | -0.008 | -0.008 | -0.0080001 | -0.008 | -0.008 | -0.0080001 | -0.0080001 | -0.008 | -0.008 | -0.0080001 | -0.008
| 0.1 | -0.009 | 0.2 | -0.001 | 0.1 | -0.008 | 20.0 | 10.0 | 0.0 | 50.0208625E+00 | DC220MV |  | -0.0080000277488499992 | 4.9569305856596101e-08 | 1517780967.58 | 1517780983.115 | -0.008 | -0.008 | -0.0080001 | -0.008 | -0.008 | -0.0080001 | -0.0080001 | -0.0080001 | -0.0079999 | -0.008 | -0.008 | -0.008 | -0.008 | -0.0080001 | -0.0080001 | -0.008 | -0.0080001 | -0.0079999 | -0.0080001 | -0.008
| 0.1 | -0.009 | 0.2 | 0.0 | 0.1 | -0.009 | 20.0 | 10.0 | 0.0 | 50.0354166E+00 | DC220MV |  | -0.0090000242967000002 | 5.4661894308604546e-08 | 1517780997.967 | 1517781013.469 | -0.009 | -0.009 | -0.009 | -0.0090001 | -0.0089999 | -0.009 | -0.0090001 | -0.0090001 | -0.009 | -0.009 | -0.009 | -0.0090001 | -0.0090001 | -0.009 | -0.0090001 | -0.009 | -0.009 | -0.0089999 | -0.009 | -0.0090001
| 0.1 | -0.01 | 0.2 | -0.001 | 0.1 | -0.009000000000000001 | 20.0 | 10.0 | 0.0 | 49.9588075E+00 | DC220MV |  | -0.0090000317858499983 | 6.419338398332715e-08 | 1517781028.333 | 1517781043.83 | -0.009 | -0.009 | -0.009 | -0.009 | -0.009 | -0.009 | -0.009 | -0.009 | -0.0090001 | -0.009 | -0.009 | -0.009 | -0.009 | -0.0089999 | -0.009 | -0.0090001 | -0.0090002 | -0.009 | -0.0090002 | -0.009
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 49.9940856E+00 | DC220MV |  | -0.01000003267215 | 3.5839945436983187e-08 | 1517781058.682 | 1517781074.202 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.01 | 0.2 | -0.001 | 0.1 | -0.009000000000000001 | 20.0 | 10.0 | 0.0 | 50.0313120E+00 | DC220MV |  | -0.0089999977504 | 4.9196718936655396e-08 | 1517781089.06 | 1517781104.585 | -0.0089999 | -0.009 | -0.0090001 | -0.009 | -0.009 | -0.0089999 | -0.009 | -0.0089999 | -0.0090001 | -0.009 | -0.009 | -0.0090001 | -0.009 | -0.009 | -0.009 | -0.009 | -0.0089999 | -0.009 | -0.0090001 | -0.009
| 0.1 | -0.009 | 0.2 | 0.0 | 0.1 | -0.009 | 20.0 | 10.0 | 0.0 | 50.0230956E+00 | DC220MV |  | -0.0090000387628000004 | 3.7056722834833882e-08 | 1517781119.44 | 1517781134.938 | -0.0090001 | -0.0090001 | -0.009 | -0.0090001 | -0.009 | -0.009 | -0.009 | -0.009 | -0.009 | -0.0090001 | -0.009 | -0.009 | -0.0090001 | -0.0090001 | -0.009 | -0.0090001 | -0.009 | -0.009 | -0.0090001 | -0.009
| 0.1 | -0.009 | 0.2 | -0.001 | 0.1 | -0.008 | 20.0 | 10.0 | 0.0 | 49.9979718E+00 | DC220MV |  | -0.0079999932673999984 | 6.2503627378498259e-08 | 1517781149.791 | 1517781165.298 | -0.0080001 | -0.0080001 | -0.008 | -0.008 | -0.008 | -0.0080001 | -0.0079999 | -0.008 | -0.0080001 | -0.008 | -0.008 | -0.008 | -0.008 | -0.0079999 | -0.008 | -0.008 | -0.008 | -0.008 | -0.0079999 | -0.0079998
| 0.1 | -0.008 | 0.2 | 0.0 | 0.1 | -0.008 | 20.0 | 10.0 | 0.0 | 49.9898605E+00 | DC220MV |  | -0.0080000385140500015 | 3.7476097650951555e-08 | 1517781180.149 | 1517781195.674 | -0.008 | -0.0080001 | -0.0080001 | -0.008 | -0.008 | -0.008 | -0.008 | -0.008 | -0.0080001 | -0.0080001 | -0.008 | -0.0080001 | -0.0080001 | -0.0080001 | -0.008 | -0.008 | -0.008 | -0.008 | -0.008 | -0.008
| 0.1 | -0.008 | 0.2 | -0.001 | 0.1 | -0.007 | 20.0 | 10.0 | 0.0 | 50.0042504E+00 | DC220MV |  | -0.0070000188980000001 | 3.688611677991139e-08 | 1517781210.517 | 1517781226.043 | -0.007 | -0.0070001 | -0.007 | -0.007 | -0.0070001 | -0.0070001 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.0070001 | -0.007 | -0.007 | -0.007 | -0.0070001 | -0.007 | -0.007 | -0.007 | -0.007
| 0.1 | -0.007 | 0.2 | 0.0 | 0.1 | -0.007 | 20.0 | 10.0 | 0.0 | 50.0151508E+00 | DC220MV |  | -0.007000019517599999 | 3.736436718635796e-08 | 1517781240.897 | 1517781256.417 | -0.0070001 | -0.007 | -0.007 | -0.0070001 | -0.007 | -0.0070001 | -0.0070001 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.007 | -0.0070001 | -0.007 | -0.007 | -0.0070001 | -0.007 | -0.007
| 0.1 | -0.007 | 0.2 | -0.001 | 0.1 | -0.006 | 20.0 | 10.0 | 0.0 | 49.9975709E+00 | DC220MV |  | -0.0060000143584000003 | 4.310961258985316e-08 | 1517781271.262 | 1517781286.773 | -0.006 | -0.006 | -0.006 | -0.0060001 | -0.0060001 | -0.006 | -0.006 | -0.006 | -0.006 | -0.006 | -0.006 | -0.006 | -0.006 | -0.006 | -0.0059999 | -0.006 | -0.006 | -0.006 | -0.006 | -0.0060001
| 0.1 | -0.006 | 0.2 | 0.0 | 0.1 | -0.006 | 20.0 | 10.0 | 0.0 | 49.9802955E+00 | DC220MV |  | -0.0060000136037500007 | 4.5341340475087371e-08 | 1517781301.625 | 1517781317.148 | -0.006 | -0.006 | -0.0060001 | -0.006 | -0.006 | -0.006 | -0.0059999 | -0.0060001 | -0.0060001 | -0.006 | -0.006 | -0.006 | -0.0059999 | -0.006 | -0.006 | -0.0060001 | -0.006 | -0.006 | -0.006 | -0.006
| 0.1 | -0.006 | 0.2 | -0.001 | 0.1 | -0.005 | 20.0 | 10.0 | 0.0 | 49.9608844E+00 | DC220MV |  | -0.0050000409687000003 | 3.1365585586822779e-08 | 1517781332.003 | 1517781347.551 | -0.005 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.005 | -0.0050001 | -0.005 | -0.0050001 | -0.005 | -0.0050001 | -0.005 | -0.0050001 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.005 | -0.005 | -0.005
| 0.1 | -0.005 | 0.2 | 0.0 | 0.1 | -0.005 | 20.0 | 10.0 | 0.0 | 50.0285588E+00 | DC220MV |  | -0.0050000415581000004 | 5.6622607878750302e-08 | 1517781362.379 | 1517781377.91 | -0.005 | -0.005 | -0.005 | -0.0050001 | -0.005 | -0.005 | -0.0050001 | -0.0050001 | -0.0050001 | -0.0050001 | -0.0050001 | -0.005 | -0.005 | -0.0050001 | -0.005 | -0.0050001 | -0.0050001 | -0.005 | -0.0050001 | -0.0049998
| 0.1 | -0.005 | 0.2 | -0.001 | 0.1 | -0.004 | 20.0 | 10.0 | 0.0 | 49.9963446E+00 | DC220MV |  | -0.0040000273800999998 | 4.5999346102848565e-08 | 1517781392.76 | 1517781408.259 | -0.004 | -0.004 | -0.004 | -0.0040001 | -0.0040001 | -0.0040001 | -0.0040001 | -0.004 | -0.004 | -0.0039999 | -0.0040001 | -0.0040001 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.004 | -0.0040001 | -0.004
| 0.1 | -0.004 | 0.2 | 0.0 | 0.1 | -0.004 | 20.0 | 10.0 | 0.0 | 50.0048542E+00 | DC220MV |  | -0.0040000493603999992 | 3.2934984037283285e-08 | 1517781423.113 | 1517781438.641 | -0.0040001 | -0.004 | -0.0040001 | -0.004 | -0.0040001 | -0.004 | -0.0040001 | -0.0040001 | -0.0040001 | -0.004 | -0.0040001 | -0.0040001 | -0.0040001 | -0.004 | -0.004 | -0.004 | -0.004 | -0.0040001 | -0.004 | -0.0040001
| 0.1 | -0.004 | 0.2 | -0.001 | 0.1 | -0.003 | 20.0 | 10.0 | 0.0 | 49.9808328E+00 | DC220MV |  | -0.0030000400831500002 | 4.3912638736789023e-08 | 1517781453.487 | 1517781468.997 | -0.0030001 | -0.003 | -0.003 | -0.0030001 | -0.003 | -0.003 | -0.0030001 | -0.0030001 | -0.0030001 | -0.003 | -0.003 | -0.003 | -0.0030001 | -0.003 | -0.0030001 | -0.003 | -0.0030001 | -0.003 | -0.0030001 | -0.0030001
| 0.1 | -0.003 | 0.2 | 0.0 | 0.1 | -0.003 | 20.0 | 10.0 | 0.0 | 50.0332296E+00 | DC220MV |  | -0.0030000503868499997 | 2.9379995046377003e-08 | 1517781483.851 | 1517781499.383 | -0.0030001 | -0.003 | -0.0030001 | -0.003 | -0.0030001 | -0.0030001 | -0.003 | -0.003 | -0.003 | -0.003 | -0.0030001 | -0.0030001 | -0.003 | -0.0030001 | -0.003 | -0.0030001 | -0.003 | -0.003 | -0.003 | -0.0030001
| 0.1 | -0.003 | 0.2 | -0.001 | 0.1 | -0.002 | 20.0 | 10.0 | 0.0 | 50.0219672E+00 | DC220MV |  | -0.0020000178086000001 | 4.7422598922235786e-08 | 1517781514.243 | 1517781529.749 | -0.002 | -0.0019999 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.002 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.0019999 | -0.002 | -0.002 | -0.0020001 | -0.0020001 | -0.002
| 0.1 | -0.002 | 0.2 | 0.0 | 0.1 | -0.002 | 20.0 | 10.0 | 0.0 | 49.9915439E+00 | DC220MV |  | -0.0020000469674499996 | 2.8841980199846318e-08 | 1517781544.595 | 1517781560.097 | -0.002 | -0.0020001 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.002 | -0.002 | -0.0020001 | -0.0020001 | -0.002 | -0.0020001 | -0.002 | -0.002 | -0.0020001 | -0.0020001 | -0.0020001
| 0.1 | -0.002 | 0.2 | -0.001 | 0.1 | -0.001 | 20.0 | 10.0 | 0.0 | 50.0639969E+00 | DC220MV |  | -0.0010000190468650002 | 3.0355344274132896e-08 | 1517781574.955 | 1517781590.479 | -0.001 | -0.00099999 | -0.00099999 | -0.00099995 | -0.001 | -0.001 | -0.001 | -0.00099997 | -0.001 | -0.001 | -0.001 | -0.001 | -0.001 | -0.001 | -0.001 | -0.001 | -0.0010001 | -0.001 | -0.0010001 | -0.0010001
| 0.1 | -0.001 | 0.2 | 0.0 | 0.1 | -0.001 | 20.0 | 10.0 | 0.0 | 50.0000047E+00 | DC220MV |  | -0.0010000301815850001 | 3.6173399789925867e-08 | 1517781605.324 | 1517781620.808 | -0.0010001 | -0.001 | -0.0010001 | -0.001 | -0.001 | -0.0010001 | -0.00099996 | -0.0010001 | -0.001 | -0.001 | -0.0010001 | -0.001 | -0.001 | -0.001 | -0.001 | -0.00099998 | -0.00099995 | -0.0010001 | -0.0010001 | -0.0010001
| 0.1 | -0.001 | 0.2 | -0.001 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9476445E+00 | DC220MV |  | -1.35270350737e-08 | 3.6684212541245154e-08 | 1517781635.667 | 1517781651.236 | 2.0596e-09 | -2.7363e-08 | 6.9437e-08 | -2.1037e-08 | -8.7679e-08 | -7.9441e-09 | -1.1181e-08 | -6.4435e-08 | 1.3093e-08 | -4.1633e-08 | 2.6039e-08 | -4.3251e-08 | -4.4134e-08 | -3.9573e-08 | 1.5888e-08 | -1.7212e-08 | 2.1037e-08 | -3.4277e-08 | -2.2067e-08 | 4.3693e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9309540E+00 | DC220MV |  | -6.0589034671899997e-08 | 5.8018037021789895e-08 | 1517781666.094 | 1517781681.698 | -4.0094e-08 | -1.2586e-07 | 1.1162e-08 | -1.1441e-07 | -1.031e-07 | -1.1558e-07 | -1.0471e-07 | -8.4007e-08 | -5.7277e-08 | 8.2979e-08 | -2.203e-08 | -3.0548e-08 | -9.6931e-09 | -9.2084e-08 | -1.1602e-08 | -8.2832e-08 | -4.171e-08 | 2.4967e-09 | -1.6757e-07 | -1.053e-07
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9345951E+00 | DC220MV |  | 1.050365811355e-08 | 2.9252425484590569e-08 | 1517781696.555 | 1517781712.159 | 4.7303e-08 | -4.4953e-08 | -1.1165e-08 | 1.4837e-08 | 2.3652e-08 | 2.6443e-08 | -1.5278e-08 | 1.0283e-08 | -2.0567e-09 | 1.6159e-09 | 2.9087e-08 | 4.3924e-08 | -2.1595e-08 | 6.0525e-08 | -1.6747e-08 | 3.3935e-08 | 5.8321e-08 | 3.9664e-09 | 4.554e-09 | -3.6579e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9508879E+00 | DC220MV |  | -4.6831680440949996e-08 | 3.5726544386862486e-08 | 1517781727.019 | 1517781742.6 | -3.5137e-08 | -3.2931e-08 | -6.5569e-08 | -8.6151e-08 | -2.4846e-08 | -4.7633e-08 | -1.0335e-07 | -3.0432e-08 | -9.3502e-08 | -1.426e-08 | -1.6466e-08 | -5.8365e-08 | -3.9547e-08 | -4.5722e-08 | -1.2643e-07 | -4.2193e-08 | 9.703e-09 | -8.1446e-08 | -7.6448e-09 | 5.2925e-09
| 0.1 | -0.001 | 0.2 | 0.0 | 0.1 | -0.001 | 20.0 | 10.0 | 0.0 | 50.0220050E+00 | DC220MV |  | -0.0010000689682499997 | 3.7558948148055816e-08 | 1517781757.467 | 1517781773.009 | -0.001 | -0.0010001 | -0.00099997 | -0.0010001 | -0.0010001 | -0.0010001 | -0.0010001 | -0.001 | -0.0010001 | -0.001 | -0.0010001 | -0.0010001 | -0.0010001 | -0.001 | -0.0010001 | -0.0010001 | -0.0010001 | -0.001 | -0.0010001 | -0.001
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 49.9488684E+00 | DC220MV |  | 0.0099999499416000021 | 3.978678831656072e-08 | 1517781787.854 | 1517781803.367 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.01 | 0.01 | 0.01 | 0.01 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.01 | 0.01
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9719120E+00 | DC220MV |  | -3.7745383073744996e-08 | 3.667287717898692e-08 | 1517781818.233 | 1517781833.819 | -1.9696e-08 | -3.4688e-08 | -1.7197e-08 | -2.1607e-08 | 6.6143e-09 | 1.8079e-08 | -4.8505e-08 | -5.7618e-08 | -8.525e-08 | -1.0759e-07 | 1.2494e-08 | -2.1313e-08 | -1.6315e-08 | -1.0039e-07 | -2.0137e-08 | -7.5403e-08 | -2.9397e-10 | -2.0578e-08 | -6.5849e-08 | -7.9665e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0530042E+00 | DC220MV |  | -8.3778451156000009e-08 | 5.2763367744036423e-08 | 1517781848.703 | 1517781864.279 | -6.5723e-08 | -8.9542e-08 | -6.1459e-08 | -1.1821e-07 | -6.2782e-08 | -8.5719e-08 | -9.263e-08 | -1.8055e-07 | -9.6306e-08 | -1.2821e-07 | -5.4255e-08 | -1.1027e-08 | -1.4483e-07 | -1.46e-07 | -7.1016e-08 | -6.587e-08 | -3.0583e-08 | 6.3518e-08 | -9.9687e-08 | -1.3468e-07
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0404478E+00 | DC220MV |  | -5.4956350446749997e-08 | 2.7697533186608304e-08 | 1517781879.142 | 1517781894.687 | -6.0515e-08 | -3.2392e-08 | -1.091e-07 | -5.8895e-08 | -1.0792e-07 | -6.3901e-08 | -7.3324e-08 | -3.3717e-08 | -4.1963e-08 | -1.9435e-08 | -1.9288e-08 | -5.7864e-08 | -3.1509e-08 | -8.3631e-08 | -6.0367e-08 | -4.8588e-09 | -4.3877e-08 | -7.8478e-08 | -7.9655e-08 | -3.8429e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9963022E+00 | DC220MV |  | -7.1618850826849991e-08 | 4.0947499866561834e-08 | 1517781909.528 | 1517781925.082 | -1.1596e-07 | -1.1066e-07 | -6.7543e-08 | -8.1522e-08 | 3.6788e-09 | -1.3862e-07 | -2.2809e-08 | -1.7364e-08 | -9.8739e-08 | -6.6366e-08 | 1.5745e-08 | -9.6826e-08 | -7.5783e-08 | -5.6359e-08 | -9.3e-08 | -5.1797e-08 | -5.8861e-08 | -7.4165e-08 | -9.3147e-08 | -1.3229e-07
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0934526E+00 | DC220MV |  | 0.009999923716350001 | 6.5580550175667095e-08 | 1517781939.943 | 1517781955.469 | 0.0099999 | 0.0099999 | 0.01 | 0.01 | 0.01 | 0.01 | 0.0099999 | 0.0099997 | 0.0099999 | 0.0099999 | 0.01 | 0.01 | 0.01 | 0.01 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099998 | 0.01
| 0.1 | 0.02 | 0.2 | 0.01 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0088082E+00 | DC220MV |  | 0.0099998777237499982 | 3.8145504981445916e-08 | 1517781970.306 | 1517781985.781 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099998 | 0.0099998 | 0.0099998 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099998
| 0.1 | 0.02 | 0.2 | 0.0 | 0.1 | 0.02 | 20.0 | 10.0 | 0.0 | 50.0182425E+00 | DC220MV |  | 0.019999889500500002 | 3.8950446656707441e-08 | 1517782000.636 | 1517782016.157 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| 0.1 | 0.03 | 0.2 | 0.01 | 0.1 | 0.02 | 20.0 | 10.0 | 0.0 | 50.0027832E+00 | DC220MV |  | 0.019999872130500005 | 6.3370593138535889e-08 | 1517782031.013 | 1517782046.522 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| 0.1 | 0.03 | 0.2 | 0.0 | 0.1 | 0.03 | 20.0 | 10.0 | 0.0 | 49.9973162E+00 | DC220MV |  | 0.029999926678500006 | 4.2470860159748308e-08 | 1517782061.369 | 1517782076.882 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
| 0.1 | 0.04 | 0.2 | 0.01 | 0.1 | 0.03 | 20.0 | 10.0 | 0.0 | 50.0001981E+00 | DC220MV |  | 0.029999874693 | 3.680233417880775e-08 | 1517782091.732 | 1517782107.259 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
| 0.1 | 0.04 | 0.2 | 0.0 | 0.1 | 0.04 | 20.0 | 10.0 | 0.0 | 50.0422905E+00 | DC220MV |  | 0.039999909847000012 | 4.1330167565048331e-08 | 1517782122.104 | 1517782137.637 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
| 0.1 | 0.05 | 0.2 | 0.01 | 0.1 | 0.04 | 20.0 | 10.0 | 0.0 | 50.0323842E+00 | DC220MV |  | 0.039999835892500002 | 5.3736553375584871e-08 | 1517782152.484 | 1517782167.993 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
| 0.1 | 0.05 | 0.2 | 0.0 | 0.1 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0030851E+00 | DC220MV |  | 0.049999888525000001 | 3.336367508234925e-08 | 1517782182.816 | 1517782198.317 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 0.1 | 0.05 | 0.2 | 0.01 | 0.1 | 0.04 | 20.0 | 10.0 | 0.0 | 49.9853722E+00 | DC220MV |  | 0.039999862602000009 | 3.6986419075085053e-08 | 1517782213.165 | 1517782228.679 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
| 0.1 | 0.04 | 0.2 | 0.0 | 0.1 | 0.04 | 20.0 | 10.0 | 0.0 | 49.9880500E+00 | DC220MV |  | 0.039999915484999995 | 4.2350529099420914e-08 | 1517782243.536 | 1517782259.053 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
| 0.1 | 0.04 | 0.2 | 0.01 | 0.1 | 0.03 | 20.0 | 10.0 | 0.0 | 50.0064678E+00 | DC220MV |  | 0.029999852732499999 | 5.1689846863224179e-08 | 1517782273.881 | 1517782289.385 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
| 0.1 | 0.03 | 0.2 | 0.0 | 0.1 | 0.03 | 20.0 | 10.0 | 0.0 | 50.0044013E+00 | DC220MV |  | 0.029999899869499996 | 4.0667044947399698e-08 | 1517782304.26 | 1517782319.776 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
| 0.1 | 0.03 | 0.2 | 0.01 | 0.1 | 0.02 | 20.0 | 10.0 | 0.0 | 50.0347269E+00 | DC220MV |  | 0.019999874556 | 5.148958471393142e-08 | 1517782334.626 | 1517782350.142 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| 0.1 | 0.02 | 0.2 | 0.0 | 0.1 | 0.02 | 20.0 | 10.0 | 0.0 | 49.9578657E+00 | DC220MV |  | 0.019999910422499999 | 4.9093337111230031e-08 | 1517782364.992 | 1517782380.487 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| 0.1 | 0.02 | 0.2 | 0.01 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0097802E+00 | DC220MV |  | 0.0099999087445499998 | 5.5121741045227916e-08 | 1517782395.346 | 1517782410.895 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099998 | 0.01 | 0.0099999 | 0.0099999 | 0.0099999
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 49.9831378E+00 | DC220MV |  | 0.0099999521751000017 | 3.5290940903931561e-08 | 1517782425.75 | 1517782441.258 | 0.01 | 0.0099999 | 0.0099999 | 0.01 | 0.01 | 0.01 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.01 | 0.01 | 0.01 | 0.01
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0362905E+00 | DC220MV |  | -8.2294205485499987e-08 | 4.6530658103218507e-08 | 1517782456.12 | 1517782471.696 | -5.4544e-08 | -4.0577e-08 | -8.7036e-08 | -7.792e-08 | -4.7928e-08 | -1.4687e-07 | -4.8516e-08 | -7.5568e-08 | -1.035e-07 | -1.5863e-07 | -8.7477e-08 | -2.3964e-08 | -1.9995e-08 | -2.367e-08 | -1.5305e-07 | -1.1585e-07 | -9.6004e-08 | -1.2644e-08 | -1.4687e-07 | -1.2526e-07
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9774206E+00 | DC220MV |  | -5.6050880255999997e-08 | 4.5062372811212344e-08 | 1517782486.543 | 1517782502.093 | -4.4894e-08 | -4.3422e-08 | -2.3256e-08 | -5.7699e-08 | -6.5353e-08 | -9.5969e-08 | -9.7589e-08 | -1.5455e-08 | -1.3247e-08 | 2.3845e-08 | -8.2133e-08 | -4.3422e-08 | -4.5335e-08 | -2.8703e-08 | -1.6942e-07 | -4.9751e-08 | -1.0451e-07 | -1.1334e-07 | -6.9916e-08 | 1.8546e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0520021E+00 | DC220MV |  | -1.0324281335487e-07 | 4.204079621243898e-08 | 1517782516.936 | 1517782532.515 | -9.1278e-08 | -5.071e-08 | -1.5654e-07 | -9.51e-08 | -1.5301e-07 | -1.1729e-07 | -1.5022e-07 | -1.2758e-07 | -1.6962e-07 | -8.7163e-08 | -6.144e-08 | -7.8637e-08 | -1.0039e-07 | -1.414e-07 | -5.4385e-08 | 8.8191e-10 | -9.9068e-08 | -9.3924e-08 | -1.486e-07 | -8.9367e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0302636E+00 | DC220MV |  | -4.382397353115e-08 | 2.6366209568630072e-08 | 1517782547.36 | 1517782562.93 | -5.6186e-08 | -4.0448e-08 | -2.9711e-08 | -2.3387e-08 | -6.81e-08 | -3.8683e-08 | -2.471e-08 | -4.3537e-08 | -6.457e-08 | -7.516e-08 | -2.3092e-08 | -6.8836e-08 | -7.6484e-09 | -4.3831e-08 | 2.6034e-08 | -7.869e-08 | -2.7211e-08 | -4.9421e-08 | -8.7368e-08 | -5.1921e-08
| 0.1 | 0.1 | 0.2 | 0.0 | 0.1 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9869987E+00 | DC220MV |  | 0.09999979760500001 | 3.6706689238945601e-08 | 1517782577.787 | 1517782593.292 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 0.1 | -0.1 | 0.2 | 0.0 | 0.1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9572959E+00 | DC220MV |  | -0.09999993339249999 | 5.1395566431811718e-08 | 1517782608.141 | 1517782623.66 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0181339E+00 | DC220MV |  | -4.7784735053000009e-08 | 4.0765898504190422e-08 | 1517782638.506 | 1517782654.1 | -5.76e-08 | -4.0849e-08 | -1.3974e-07 | -4.511e-08 | -7.7878e-08 | -2.498e-08 | -6.7739e-08 | -8.3021e-08 | -5.1429e-08 | -3.747e-08 | -9.0074e-08 | -2.3804e-08 | -1.969e-08 | 7.3029e-08 | -6.6711e-08 | -3.3355e-08 | -2.7625e-08 | -8.4637e-08 | -4.6139e-08 | -1.0874e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0715456E+00 | DC220MV |  | 2.3404516588534997e-08 | 4.6014285036365029e-08 | 1517782668.957 | 1517782684.516 | 4.6206e-08 | 1.0595e-08 | 5.1503e-09 | -5.8861e-08 | -7.4753e-08 | 7.9609e-08 | 4.2674e-09 | 9.712e-08 | -2.4869e-08 | -2.2073e-09 | 3.9142e-08 | -2.0896e-08 | 3.6199e-08 | 2.943e-10 | 6.7837e-08 | 4.6059e-08 | 7.5195e-08 | 2.2073e-09 | 7.7991e-08 | 6.1804e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9518059E+00 | DC220MV |  | -1.5657915837600001e-08 | 3.6594794513831777e-08 | 1517782699.356 | 1517782714.89 | -8.5434e-09 | -2.2684e-08 | -5.0082e-09 | -8.5434e-09 | -1.6203e-08 | -5.7594e-08 | -5.7299e-08 | -1.5025e-08 | 3.2406e-09 | -1.6792e-08 | -7.6448e-08 | -6.0245e-08 | -4.3748e-08 | -3.1522e-08 | 3.7267e-08 | 6.0982e-08 | -6.0393e-08 | 4.743e-08 | 2.946e-09 | 1.5025e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0252769E+00 | DC220MV |  | -7.0178347586500012e-09 | 6.7301447980943809e-08 | 1517782729.752 | 1517782745.343 | -1.7489e-08 | -3.3803e-08 | -4.997e-09 | -3.7477e-08 | -2.1825e-07 | 5.5555e-08 | -5.7465e-08 | 4.997e-09 | -1.3374e-08 | 6.9811e-08 | -1.7783e-08 | -7.6425e-09 | -4.9529e-08 | -4.4091e-08 | 4.6002e-08 | 4.997e-08 | 9.9646e-08 | -2.6161e-08 | -3.1158e-08 | 9.2885e-08
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 50.0142069E+00 | DC220MV |  | -0.0100000389927 | 3.769985284058962e-08 | 1517782760.191 | 1517782775.676 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.02 | 0.2 | -0.01 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 49.9766713E+00 | DC220MV |  | -0.0099999899949500017 | 3.7172981608434396e-08 | 1517782790.529 | 1517782806.035 | -0.01 | -0.01 | -0.0099999 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.0099999 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.02 | 0.2 | 0.0 | 0.1 | -0.02 | 20.0 | 10.0 | 0.0 | 49.9607808E+00 | DC220MV |  | -0.0200000472605 | 3.827193716996355e-08 | 1517782820.883 | 1517782836.392 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02
| 0.1 | -0.03 | 0.2 | -0.01 | 0.1 | -0.02 | 20.0 | 10.0 | 0.0 | 50.0284076E+00 | DC220MV |  | -0.019999940548999999 | 4.0798877177781718e-08 | 1517782851.25 | 1517782866.776 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02
| 0.1 | -0.03 | 0.2 | 0.0 | 0.1 | -0.03 | 20.0 | 10.0 | 0.0 | 49.9673044E+00 | DC220MV |  | -0.030000020577999997 | 5.3196286205076332e-08 | 1517782881.62 | 1517782897.127 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03
| 0.1 | -0.04 | 0.2 | -0.01 | 0.1 | -0.03 | 20.0 | 10.0 | 0.0 | 49.9829775E+00 | DC220MV |  | -0.029999951089499998 | 4.5186286025446713e-08 | 1517782911.97 | 1517782927.485 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03
| 0.1 | -0.04 | 0.2 | 0.0 | 0.1 | -0.04 | 20.0 | 10.0 | 0.0 | 49.9708849E+00 | DC220MV |  | -0.039999992488499996 | 2.9054290781489951e-08 | 1517782942.349 | 1517782957.88 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04
| 0.1 | -0.05 | 0.2 | -0.01 | 0.1 | -0.04 | 20.0 | 10.0 | 0.0 | 50.0249181E+00 | DC220MV |  | -0.039999912724999996 | 4.8932089522104468e-08 | 1517782972.719 | 1517782988.243 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04
| 0.1 | -0.05 | 0.2 | 0.0 | 0.1 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9575220E+00 | DC220MV |  | -0.049999971559000003 | 3.9790297548784646e-08 | 1517783003.09 | 1517783018.591 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 0.1 | -0.05 | 0.2 | -0.01 | 0.1 | -0.04 | 20.0 | 10.0 | 0.0 | 50.0314396E+00 | DC220MV |  | -0.039999931103999992 | 4.4653769090333126e-08 | 1517783033.444 | 1517783048.982 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04
| 0.1 | -0.04 | 0.2 | 0.0 | 0.1 | -0.04 | 20.0 | 10.0 | 0.0 | 49.9975426E+00 | DC220MV |  | -0.039999990818999999 | 3.206408612385594e-08 | 1517783063.833 | 1517783079.337 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04
| 0.1 | -0.04 | 0.2 | -0.01 | 0.1 | -0.03 | 20.0 | 10.0 | 0.0 | 50.0591595E+00 | DC220MV |  | -0.029999941282000002 | 4.3253979192812169e-08 | 1517783094.18 | 1517783109.692 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03
| 0.1 | -0.03 | 0.2 | 0.0 | 0.1 | -0.03 | 20.0 | 10.0 | 0.0 | 50.0168689E+00 | DC220MV |  | -0.029999985045500006 | 3.4320874475820553e-08 | 1517783124.536 | 1517783140.021 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03
| 0.1 | -0.03 | 0.2 | -0.01 | 0.1 | -0.02 | 20.0 | 10.0 | 0.0 | 50.0448232E+00 | DC220MV |  | -0.0199999641505 | 4.7903977024202404e-08 | 1517783154.859 | 1517783170.366 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02
| 0.1 | -0.02 | 0.2 | 0.0 | 0.1 | -0.02 | 20.0 | 10.0 | 0.0 | 50.0247198E+00 | DC220MV |  | -0.020000020662999999 | 3.8254553206139821e-08 | 1517783185.199 | 1517783200.68 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02
| 0.1 | -0.02 | 0.2 | -0.01 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 50.0021180E+00 | DC220MV |  | -0.0099999812267500005 | 4.8027935975765743e-08 | 1517783215.525 | 1517783231.035 | -0.01 | -0.01 | -0.01 | -0.0099999 | -0.01 | -0.01 | -0.0099999 | -0.01 | -0.0099999 | -0.01 | -0.0099999 | -0.01 | -0.01 | -0.01 | -0.0099999 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 50.0022076E+00 | DC220MV |  | -0.010000037211949999 | 2.7245598658153784e-08 | 1517783245.865 | 1517783261.385 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9870647E+00 | DC220MV |  | 2.6684403213349999e-08 | 3.9284545351122812e-08 | 1517783276.238 | 1517783291.799 | 3.9287e-08 | 6.4743e-09 | 1.1477e-08 | 8.4607e-08 | 5.0912e-08 | -7.416e-08 | 3.7963e-08 | 6.1506e-08 | 4.5762e-08 | 2.3543e-08 | 4.2966e-08 | 1.0344e-07 | 5.3119e-08 | 3.09e-09 | 2.1483e-08 | 4.2819e-08 | -4.0759e-08 | 1.133e-08 | -1.5156e-08 | 2.3984e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0245215E+00 | DC220MV |  | -4.8327673816500003e-08 | 4.3826712834087822e-08 | 1517783306.661 | 1517783322.229 | -4.5158e-08 | -5.9867e-08 | 5.5896e-08 | -7.1782e-08 | -4.7364e-08 | -7.2517e-08 | -5.2218e-08 | -5.1483e-08 | -1.521e-07 | -8.4579e-08 | 1.2062e-08 | -9.414e-09 | -2.6771e-08 | -1.1209e-07 | -8.9139e-08 | -9.414e-09 | -6.1485e-08 | -2.3829e-08 | -2.8683e-08 | -3.6626e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0551929E+00 | DC220MV |  | 2.1768097021579999e-08 | 5.2146812874375148e-08 | 1517783337.076 | 1517783352.657 | -2.5142e-08 | -2.3966e-08 | -4.1022e-08 | 8.7484e-08 | -9.8659e-08 | -1.9261e-08 | 5.5284e-08 | 5.6166e-08 | 2.5142e-08 | 6.1606e-08 | 5.4549e-08 | -5.2932e-09 | 3.1465e-08 | 5.5872e-09 | 8.322e-08 | 1.3336e-07 | -2.5584e-08 | 4.0287e-08 | 5.8813e-10 | 3.9552e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0589184E+00 | DC220MV |  | -1.2275307111250001e-08 | 4.6917656106702229e-08 | 1517783367.507 | 1517783383.045 | 1.62e-08 | -5.1546e-08 | -1.4875e-08 | -4.7275e-08 | 4.8453e-08 | 3.7555e-08 | -2.7982e-09 | -9.0426e-08 | -2.3564e-08 | -5.4344e-08 | 1.2901e-07 | -3.8291e-08 | -6.7746e-09 | 9.7201e-09 | 1.3402e-08 | -2.7098e-08 | -4.8453e-08 | -2.6509e-08 | -4.86e-09 | -6.3033e-08
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 49.9907941E+00 | DC220MV |  | -0.010000017318799998 | 4.009897935555373e-08 | 1517783397.884 | 1517783413.351 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0223496E+00 | DC220MV |  | 0.0099999808176500001 | 4.3397543438803829e-08 | 1517783428.176 | 1517783443.711 | 0.01 | 0.01 | 0.01 | 0.0099999 | 0.01 | 0.01 | 0.01 | 0.01 | 0.0099999 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.0099999 | 0.01 | 0.0099999 | 0.01 | 0.01
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9877624E+00 | DC220MV |  | -1.1920201786989999e-08 | 3.2553370380900584e-08 | 1517783458.554 | 1517783474.113 | -3.7379e-08 | 4.5768e-08 | -1.8543e-08 | 1.9131e-09 | -1.3245e-09 | -5.2684e-08 | -3.4289e-08 | 1.4863e-08 | 3.4289e-08 | 2.3693e-08 | 1.766e-09 | 7.3581e-10 | -3.4878e-08 | 1.2362e-08 | -6.2103e-08 | -1.4716e-10 | 1.9426e-08 | -4.7386e-08 | -7.4317e-08 | -3.0168e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0055336E+00 | DC220MV |  | -7.3043653650999997e-08 | 3.3479805095852895e-08 | 1517783488.963 | 1517783504.536 | -7.2051e-08 | -7.4698e-08 | -3.3085e-08 | -8.058e-08 | -2.9262e-08 | -4.3231e-08 | -1.3969e-07 | -1.7057e-08 | -1.0837e-07 | -1.244e-07 | -7.7933e-08 | -9.4843e-08 | -7.0434e-08 | -1.1469e-08 | -7.2345e-08 | -7.6315e-08 | -5.6906e-08 | -1.0028e-07 | -1.0734e-07 | -7.0581e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0281810E+00 | DC220MV |  | -2.9694519499799994e-08 | 4.187630164106836e-08 | 1517783519.377 | 1517783534.937 | -2.0306e-08 | -5.003e-08 | -3.4433e-08 | 1.4421e-08 | -7.6664e-08 | -6.7394e-08 | -7.3574e-08 | -3.2373e-09 | 3.6493e-08 | -9.6235e-08 | -2.1336e-08 | -1.4715e-08 | -2.9724e-08 | -4.5616e-09 | -5.8271e-08 | -6.0331e-09 | -2.3249e-08 | -1.136e-07 | -9.8589e-09 | 5.8418e-08
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9805123E+00 | DC220MV |  | -7.6517695113500009e-08 | 3.787044919050865e-08 | 1517783549.785 | 1517783565.355 | -1.5176e-07 | -1.0764e-07 | -5.1027e-08 | -7.1908e-08 | -6.1173e-08 | -1.544e-08 | -9.7053e-08 | -8.3525e-08 | -4.9115e-08 | -1.2146e-07 | -1.0499e-07 | -2.0587e-08 | -3.838e-08 | -8.573e-08 | -6.3526e-08 | -9.4995e-08 | -3.2351e-08 | -1.1573e-07 | -3.3822e-08 | -1.3014e-07
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0634389E+00 | DC220MV |  | 0.0099999334759499989 | 3.7055278898442374e-08 | 1517783580.203 | 1517783595.73 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099998 | 0.01 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.01 | 0.0099999
| 0.1 | 0.02 | 0.2 | 0.01 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0096198E+00 | DC220MV |  | 0.0099998898379999989 | 4.3762236736646463e-08 | 1517783610.584 | 1517783626.089 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.0099999 | 0.01 | 0.01 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999
| 0.1 | 0.02 | 0.2 | 0.0 | 0.1 | 0.02 | 20.0 | 10.0 | 0.0 | 50.0127768E+00 | DC220MV |  | 0.019999937520999999 | 2.7506639907465737e-08 | 1517783640.944 | 1517783656.439 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| 0.1 | 0.03 | 0.2 | 0.01 | 0.1 | 0.02 | 20.0 | 10.0 | 0.0 | 50.0332343E+00 | DC220MV |  | 0.019999882059999998 | 4.6990617148769115e-08 | 1517783671.277 | 1517783686.781 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| 0.1 | 0.03 | 0.2 | 0.0 | 0.1 | 0.03 | 20.0 | 10.0 | 0.0 | 50.0516571E+00 | DC220MV |  | 0.029999904507500001 | 6.166321187455789e-08 | 1517783701.625 | 1517783717.13 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
| 0.1 | 0.04 | 0.2 | 0.01 | 0.1 | 0.03 | 20.0 | 10.0 | 0.0 | 49.9827230E+00 | DC220MV |  | 0.029999858117500007 | 5.2032338778294822e-08 | 1517783731.985 | 1517783747.471 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
| 0.1 | 0.04 | 0.2 | 0.0 | 0.1 | 0.04 | 20.0 | 10.0 | 0.0 | 50.0174259E+00 | DC220MV |  | 0.039999928888999989 | 3.6995828670582955e-08 | 1517783762.316 | 1517783777.838 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
| 0.1 | 0.05 | 0.2 | 0.01 | 0.1 | 0.04 | 20.0 | 10.0 | 0.0 | 50.0042645E+00 | DC220MV |  | 0.039999883421000006 | 2.7359981158077526e-08 | 1517783792.667 | 1517783808.188 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
| 0.1 | 0.05 | 0.2 | 0.0 | 0.1 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0154434E+00 | DC220MV |  | 0.049999924009999995 | 2.5283203911179226e-08 | 1517783823.046 | 1517783838.563 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 0.1 | 0.06 | 0.2 | 0.01 | 0.1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9939064E+00 | DC220MV |  | 0.049999867311999999 | 4.1796984413089599e-08 | 1517783853.419 | 1517783868.926 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 0.1 | 0.06 | 0.2 | 0.0 | 0.1 | 0.06 | 20.0 | 10.0 | 0.0 | 49.9670265E+00 | DC220MV |  | 0.05999987227100001 | 4.3207758203455561e-08 | 1517783883.768 | 1517783899.274 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06
| 0.1 | 0.07 | 0.2 | 0.01 | 0.1 | 0.06 | 20.0 | 10.0 | 0.0 | 49.9697023E+00 | DC220MV |  | 0.059999841982000014 | 4.8126307420645428e-08 | 1517783914.111 | 1517783929.634 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06
| 0.1 | 0.07 | 0.2 | 0.0 | 0.1 | 0.07 | 20.0 | 10.0 | 0.0 | 50.0241107E+00 | DC220MV |  | 0.069999848248000013 | 4.8753871497485309e-08 | 1517783944.471 | 1517783959.991 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07
| 0.1 | 0.08 | 0.2 | 0.01 | 0.1 | 0.07 | 20.0 | 10.0 | 0.0 | 50.0122955E+00 | DC220MV |  | 0.069999824439500008 | 2.825464766594587e-08 | 1517783974.84 | 1517783990.343 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07
| 0.1 | 0.08 | 0.2 | 0.0 | 0.1 | 0.08 | 20.0 | 10.0 | 0.0 | 50.0227745E+00 | DC220MV |  | 0.079999823349999985 | 8.0230453132685024e-08 | 1517784005.181 | 1517784020.692 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08
| 0.1 | 0.09 | 0.2 | 0.01 | 0.1 | 0.08 | 20.0 | 10.0 | 0.0 | 50.0381045E+00 | DC220MV |  | 0.079999843705499993 | 3.0850233464412574e-08 | 1517784035.559 | 1517784051.067 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08
| 0.1 | 0.09 | 0.2 | 0.0 | 0.1 | 0.09 | 20.0 | 10.0 | 0.0 | 50.0126919E+00 | DC220MV |  | 0.089999833183999997 | 4.2630286227748363e-08 | 1517784065.908 | 1517784081.409 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09
| 0.1 | 0.1 | 0.2 | 0.01 | 0.1 | 0.09000000000000001 | 20.0 | 10.0 | 0.0 | 49.9741549E+00 | DC220MV |  | 0.089999787826500005 | 2.3154953956284659e-08 | 1517784096.261 | 1517784111.77 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09
| 0.1 | 0.1 | 0.2 | 0.0 | 0.1 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9933264E+00 | DC220MV |  | 0.099999833405499991 | 3.3566467564100488e-08 | 1517784126.619 | 1517784142.145 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 0.1 | 0.1 | 0.2 | 0.01 | 0.1 | 0.09000000000000001 | 20.0 | 10.0 | 0.0 | 49.9790418E+00 | DC220MV |  | 0.089999783375000003 | 4.865110312348356e-08 | 1517784156.995 | 1517784172.519 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09
| 0.1 | 0.09 | 0.2 | 0.0 | 0.1 | 0.09 | 20.0 | 10.0 | 0.0 | 50.0178838E+00 | DC220MV |  | 0.089999814488000013 | 4.5451833801384779e-08 | 1517784187.355 | 1517784202.891 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09
| 0.1 | 0.09 | 0.2 | 0.01 | 0.1 | 0.08 | 20.0 | 10.0 | 0.0 | 49.9685669E+00 | DC220MV |  | 0.079999820533500007 | 5.6605736571397169e-08 | 1517784217.714 | 1517784233.229 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08
| 0.1 | 0.08 | 0.2 | 0.0 | 0.1 | 0.08 | 20.0 | 10.0 | 0.0 | 50.0100775E+00 | DC220MV |  | 0.079999819101499989 | 4.4536304546977753e-08 | 1517784248.076 | 1517784263.613 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08
| 0.1 | 0.08 | 0.2 | 0.01 | 0.1 | 0.07 | 20.0 | 10.0 | 0.0 | 49.9495086E+00 | DC220MV |  | 0.069999830039499986 | 3.6381191908441888e-08 | 1517784278.465 | 1517784293.97 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07
| 0.1 | 0.07 | 0.2 | 0.0 | 0.1 | 0.07 | 20.0 | 10.0 | 0.0 | 50.0296213E+00 | DC220MV |  | 0.069999870858499999 | 4.0583328015249709e-08 | 1517784308.797 | 1517784324.339 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07
| 0.1 | 0.07 | 0.2 | 0.01 | 0.1 | 0.06 | 20.0 | 10.0 | 0.0 | 49.9951844E+00 | DC220MV |  | 0.059999843093499997 | 3.1426115139879014e-08 | 1517784339.197 | 1517784354.712 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06
| 0.1 | 0.06 | 0.2 | 0.0 | 0.1 | 0.06 | 20.0 | 10.0 | 0.0 | 49.9554454E+00 | DC220MV |  | 0.059999887289500009 | 3.2217229471326044e-08 | 1517784369.567 | 1517784385.087 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06
| 0.1 | 0.06 | 0.2 | 0.01 | 0.1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9790937E+00 | DC220MV |  | 0.049999878780999998 | 3.7401799942796678e-08 | 1517784399.94 | 1517784415.473 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 0.1 | 0.05 | 0.2 | 0.0 | 0.1 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0021888E+00 | DC220MV |  | 0.049999878771499993 | 5.7596878845674264e-08 | 1517784430.307 | 1517784445.84 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 0.1 | 0.05 | 0.2 | 0.01 | 0.1 | 0.04 | 20.0 | 10.0 | 0.0 | 49.9744329E+00 | DC220MV |  | 0.039999864141499995 | 2.8328923960345273e-08 | 1517784460.693 | 1517784476.212 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
| 0.1 | 0.04 | 0.2 | 0.0 | 0.1 | 0.04 | 20.0 | 10.0 | 0.0 | 50.0595425E+00 | DC220MV |  | 0.039999910178000003 | 5.0282481700980781e-08 | 1517784491.048 | 1517784506.581 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
| 0.1 | 0.04 | 0.2 | 0.01 | 0.1 | 0.03 | 20.0 | 10.0 | 0.0 | 50.0592541E+00 | DC220MV |  | 0.029999888032499998 | 4.5173813418044126e-08 | 1517784521.439 | 1517784536.914 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
| 0.1 | 0.03 | 0.2 | 0.0 | 0.1 | 0.03 | 20.0 | 10.0 | 0.0 | 49.9946374E+00 | DC220MV |  | 0.029999930620500003 | 3.3729761558047737e-08 | 1517784551.758 | 1517784567.241 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
| 0.1 | 0.03 | 0.2 | 0.01 | 0.1 | 0.02 | 20.0 | 10.0 | 0.0 | 49.9849621E+00 | DC220MV |  | 0.019999879795500001 | 3.5931378414527066e-08 | 1517784582.092 | 1517784597.62 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| 0.1 | 0.02 | 0.2 | 0.0 | 0.1 | 0.02 | 20.0 | 10.0 | 0.0 | 50.0270854E+00 | DC220MV |  | 0.019999920804500002 | 5.2401802304195159e-08 | 1517784612.484 | 1517784628.007 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| 0.1 | 0.02 | 0.2 | 0.01 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0234355E+00 | DC220MV |  | 0.0099998843405999984 | 2.844490128193478e-08 | 1517784642.861 | 1517784658.378 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0170483E+00 | DC220MV |  | 0.0099999482692499989 | 3.6985988204684533e-08 | 1517784673.234 | 1517784688.746 | 0.0099999 | 0.0099999 | 0.0099999 | 0.01 | 0.01 | 0.0099999 | 0.01 | 0.01 | 0.01 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.01 | 0.0099999 | 0.01 | 0.01
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0447287E+00 | DC220MV |  | -1.333518199925e-07 | 5.2522147524388505e-08 | 1517784703.567 | 1517784719.12 | -1.1873e-07 | -1.3374e-07 | -1.308e-07 | -9.8136e-08 | -7.283e-08 | -4.6052e-08 | -1.4434e-07 | -1.2079e-07 | -1.3919e-07 | -8.9897e-08 | -1.5655e-07 | -1.5978e-07 | -8.9455e-08 | -1.5846e-07 | -1.6052e-07 | -1.4463e-07 | -1.1241e-07 | -3.2074e-07 | -1.3315e-07 | -1.3683e-07
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9532325E+00 | DC220MV |  | -7.4284902021100009e-08 | 4.6585241808544667e-08 | 1517784733.975 | 1517784749.53 | -1.0728e-07 | -1.3097e-08 | -5.0769e-08 | -3.1786e-08 | -4.0763e-08 | -1.1037e-08 | -8.4762e-08 | -1.3671e-07 | -1.2994e-07 | -7.5933e-08 | -6.8281e-08 | -1.6629e-07 | -5.8569e-08 | -2.1926e-08 | -9.1237e-08 | -8.2114e-08 | -8.1378e-08 | -1.4201e-07 | 8.388e-09 | -1.0021e-07
| 0.1 | 0.01 | 0.2 | 0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0196634E+00 | DC220MV |  | -1.3368823041600002e-07 | 5.0706161735230997e-08 | 1517784764.394 | 1517784779.955 | -1.5617e-07 | -1.8058e-07 | -1.0279e-07 | -1.1088e-07 | -1.2823e-07 | -7.7645e-08 | -1.8603e-07 | -2.372e-07 | -1.0867e-07 | -1.3911e-07 | -1.4073e-07 | -5.4558e-08 | -1.5985e-07 | -1.0309e-07 | -1.6279e-07 | -1.4985e-07 | -2.2882e-07 | -3.1029e-08 | -1.0823e-07 | -1.075e-07
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9612894E+00 | DC220MV |  | -8.1101281937500003e-08 | 3.7363668051110097e-08 | 1517784794.797 | 1517784810.356 | -8.6802e-08 | -5.8996e-08 | -3.4721e-08 | -9.4894e-08 | -1.252e-07 | -7.1502e-08 | -1.3123e-07 | -1.0519e-07 | -1.3947e-07 | -6.5322e-08 | -1.5301e-08 | -6.9442e-08 | -2.0744e-08 | -1.0505e-07 | -6.3851e-08 | -1.5536e-07 | -7.9005e-08 | -4.0753e-08 | -7.1207e-08 | -8.7979e-08
| 0.1 | 0.01 | 0.2 | 0.0 | 0.1 | 0.01 | 20.0 | 10.0 | 0.0 | 50.0159249E+00 | DC220MV |  | 0.0099998802700000007 | 3.471770156280289e-08 | 1517784825.205 | 1517784840.763 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099998 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099998 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999 | 0.0099999
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 49.9499746E+00 | DC220MV |  | -0.010000070741000001 | 2.9131495825109908e-08 | 1517784855.613 | 1517784871.11 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9963116E+00 | DC220MV |  | -1.0969792990800001e-07 | 4.3942566821717977e-08 | 1517784885.975 | 1517784901.565 | -1.3518e-07 | -1.4149e-07 | -6.4355e-08 | -8.2281e-08 | -7.3612e-08 | -1.7602e-07 | -5.9213e-08 | -7.2877e-08 | -1.4678e-07 | -1.5207e-07 | -1.5854e-07 | -9.4182e-08 | -7.5228e-08 | -5.2748e-08 | -5.3483e-08 | -7.9195e-08 | -1.246e-07 | -1.5604e-07 | -9.8443e-08 | -1.9762e-07
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0657373E+00 | DC220MV |  | -2.6913862386049994e-08 | 3.8497034528285577e-08 | 1517784916.424 | 1517784931.997 | 2.6326e-08 | -3.4267e-08 | 1.4707e-09 | 1.7648e-08 | -4.6621e-08 | -2.5737e-08 | -2.9855e-08 | 3.6032e-08 | -9.5596e-09 | 7.7947e-09 | -5.7357e-09 | -3.9709e-08 | -9.1037e-08 | -1.6913e-08 | -5.8828e-10 | -3.3238e-08 | -1.0266e-07 | -3.4709e-08 | -5.1769e-08 | -1.0516e-07
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0720849E+00 | DC220MV |  | -9.3181605004999999e-08 | 3.0993720770445945e-08 | 1517784946.835 | 1517784962.371 | -8.0842e-08 | -6.1257e-08 | -7.1123e-08 | -1.287e-07 | -1.315e-07 | -1.0205e-07 | -5.0802e-08 | -1.5079e-07 | -2.3266e-08 | -9.9101e-08 | -1.4416e-07 | -1.1206e-07 | -7.775e-08 | -9.6451e-08 | -9.6451e-08 | -8.1136e-08 | -9.7776e-08 | -6.8473e-08 | -7.613e-08 | -1.1383e-07
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0217170E+00 | DC220MV |  | -2.4726353025270006e-08 | 4.7459945500714813e-08 | 1517784977.209 | 1517784992.747 | -4.418e-10 | 2.2679e-08 | -6.038e-08 | -5.9055e-08 | 8.0998e-09 | -1.9734e-08 | 2.4152e-08 | 4.3886e-08 | -2.7392e-08 | -1.6641e-08 | -6.3178e-08 | -1.3593e-07 | -7.4223e-08 | 4.0057e-08 | -4.7126e-09 | -5.1986e-08 | -3.1368e-08 | 2.2974e-08 | -1.0868e-07 | -2.6508e-09
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 50.0263393E+00 | DC220MV |  | -0.0100000576017 | 3.5280853887220127e-08 | 1517785007.603 | 1517785023.095 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.02 | 0.2 | -0.01 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 49.9735093E+00 | DC220MV |  | -0.010000009913100001 | 3.677667391135971e-08 | 1517785037.944 | 1517785053.447 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.02 | 0.2 | 0.0 | 0.1 | -0.02 | 20.0 | 10.0 | 0.0 | 49.9562882E+00 | DC220MV |  | -0.020000049321499999 | 4.8944299491861806e-08 | 1517785068.291 | 1517785083.812 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02
| 0.1 | -0.03 | 0.2 | -0.01 | 0.1 | -0.02 | 20.0 | 10.0 | 0.0 | 50.0244837E+00 | DC220MV |  | -0.019999991301000001 | 4.9717249109326262e-08 | 1517785098.655 | 1517785114.164 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02
| 0.1 | -0.03 | 0.2 | 0.0 | 0.1 | -0.03 | 20.0 | 10.0 | 0.0 | 49.9978539E+00 | DC220MV |  | -0.030000049344999995 | 4.4699932941614315e-08 | 1517785129.01 | 1517785144.515 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03
| 0.1 | -0.04 | 0.2 | -0.01 | 0.1 | -0.03 | 20.0 | 10.0 | 0.0 | 49.9966512E+00 | DC220MV |  | -0.029999970743999998 | 5.7655441148067177e-08 | 1517785159.363 | 1517785174.876 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03
| 0.1 | -0.04 | 0.2 | 0.0 | 0.1 | -0.04 | 20.0 | 10.0 | 0.0 | 49.9853392E+00 | DC220MV |  | -0.040000037489500005 | 3.2875035281714135e-08 | 1517785189.721 | 1517785205.215 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04
| 0.1 | -0.05 | 0.2 | -0.01 | 0.1 | -0.04 | 20.0 | 10.0 | 0.0 | 49.9895682E+00 | DC220MV |  | -0.039999967182999994 | 5.4283388905466033e-08 | 1517785220.064 | 1517785235.582 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04
| 0.1 | -0.05 | 0.2 | 0.0 | 0.1 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9271676E+00 | DC220MV |  | -0.049999992198500001 | 2.8986019608893186e-08 | 1517785250.421 | 1517785265.93 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 0.1 | -0.06 | 0.2 | -0.01 | 0.1 | -0.05 | 20.0 | 10.0 | 0.0 | 50.0143107E+00 | DC220MV |  | -0.049999936761499997 | 5.1691592669964852e-08 | 1517785280.798 | 1517785296.344 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 0.1 | -0.06 | 0.2 | 0.0 | 0.1 | -0.06 | 20.0 | 10.0 | 0.0 | 50.0002406E+00 | DC220MV |  | -0.059999972322500018 | 5.246569706765773e-08 | 1517785311.206 | 1517785326.726 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06
| 0.1 | -0.07 | 0.2 | -0.01 | 0.1 | -0.06 | 20.0 | 10.0 | 0.0 | 49.9825297E+00 | DC220MV |  | -0.059999934137500002 | 4.1423174054675284e-08 | 1517785341.588 | 1517785357.098 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06
| 0.1 | -0.07 | 0.2 | 0.0 | 0.1 | -0.07 | 20.0 | 10.0 | 0.0 | 49.9744282E+00 | DC220MV |  | -0.069999975510499995 | 5.2539146403351437e-08 | 1517785371.954 | 1517785387.474 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07
| 0.1 | -0.08 | 0.2 | -0.01 | 0.1 | -0.07 | 20.0 | 10.0 | 0.0 | 49.9758419E+00 | DC220MV |  | -0.069999934235500008 | 4.1061136307001629e-08 | 1517785402.337 | 1517785417.865 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07
| 0.1 | -0.08 | 0.2 | 0.0 | 0.1 | -0.08 | 20.0 | 10.0 | 0.0 | 50.0229445E+00 | DC220MV |  | -0.07999996155950001 | 7.2431017766665208e-08 | 1517785432.728 | 1517785448.253 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08
| 0.1 | -0.09 | 0.2 | -0.01 | 0.1 | -0.08 | 20.0 | 10.0 | 0.0 | 50.0438356E+00 | DC220MV |  | -0.079999907803499998 | 4.1569773907312687e-08 | 1517785463.103 | 1517785478.614 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08
| 0.1 | -0.09 | 0.2 | 0.0 | 0.1 | -0.09 | 20.0 | 10.0 | 0.0 | 50.0183369E+00 | DC220MV |  | -0.089999953451499998 | 4.133032376877386e-08 | 1517785493.455 | 1517785508.949 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09
| 0.1 | -0.1 | 0.2 | -0.01 | 0.1 | -0.09000000000000001 | 20.0 | 10.0 | 0.0 | 49.9970426E+00 | DC220MV |  | -0.089999908345999996 | 3.8168868779756846e-08 | 1517785523.822 | 1517785539.325 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09
| 0.1 | -0.1 | 0.2 | 0.0 | 0.1 | -0.1 | 20.0 | 10.0 | 0.0 | 50.0183605E+00 | DC220MV |  | -0.099999963894499999 | 4.1987216088320033e-08 | 1517785554.171 | 1517785569.67 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 0.1 | -0.1 | 0.2 | -0.01 | 0.1 | -0.09000000000000001 | 20.0 | 10.0 | 0.0 | 49.9800928E+00 | DC220MV |  | -0.089999923064499984 | 2.7479156915658227e-08 | 1517785584.513 | 1517785600.014 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09
| 0.1 | -0.09 | 0.2 | 0.0 | 0.1 | -0.09 | 20.0 | 10.0 | 0.0 | 49.9980284E+00 | DC220MV |  | -0.089999968416500001 | 3.1111039563446352e-08 | 1517785614.842 | 1517785630.355 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09 | -0.09
| 0.1 | -0.09 | 0.2 | -0.01 | 0.1 | -0.08 | 20.0 | 10.0 | 0.0 | 50.0463590E+00 | DC220MV |  | -0.079999924677499987 | 4.5751548375787689e-08 | 1517785645.192 | 1517785660.714 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08
| 0.1 | -0.08 | 0.2 | 0.0 | 0.1 | -0.08 | 20.0 | 10.0 | 0.0 | 49.9575549E+00 | DC220MV |  | -0.07999997868799999 | 2.9815684228094742e-08 | 1517785675.56 | 1517785691.053 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08 | -0.08
| 0.1 | -0.08 | 0.2 | -0.01 | 0.1 | -0.07 | 20.0 | 10.0 | 0.0 | 50.0291160E+00 | DC220MV |  | -0.069999916838000012 | 3.6149207543565578e-08 | 1517785705.894 | 1517785721.428 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07
| 0.1 | -0.07 | 0.2 | 0.0 | 0.1 | -0.07 | 20.0 | 10.0 | 0.0 | 49.9565190E+00 | DC220MV |  | -0.069999977891500018 | 4.8706368399273356e-08 | 1517785736.275 | 1517785751.768 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07 | -0.07
| 0.1 | -0.07 | 0.2 | -0.01 | 0.1 | -0.06 | 20.0 | 10.0 | 0.0 | 50.0198617E+00 | DC220MV |  | -0.059999917461999987 | 3.9224372856671382e-08 | 1517785766.623 | 1517785782.169 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06
| 0.1 | -0.06 | 0.2 | 0.0 | 0.1 | -0.06 | 20.0 | 10.0 | 0.0 | 49.9523615E+00 | DC220MV |  | -0.059999970673000003 | 5.6510933818042501e-08 | 1517785797.017 | 1517785812.523 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06 | -0.06
| 0.1 | -0.06 | 0.2 | -0.01 | 0.1 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9918787E+00 | DC220MV |  | -0.049999933105999997 | 5.1143140048698017e-08 | 1517785827.374 | 1517785842.915 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 0.1 | -0.05 | 0.2 | 0.0 | 0.1 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9590006E+00 | DC220MV |  | -0.0499999903065 | 5.840291407786301e-08 | 1517785857.771 | 1517785873.276 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 0.1 | -0.05 | 0.2 | -0.01 | 0.1 | -0.04 | 20.0 | 10.0 | 0.0 | 50.0344105E+00 | DC220MV |  | -0.039999944189999997 | 4.1687457226012224e-08 | 1517785888.138 | 1517785903.661 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04
| 0.1 | -0.04 | 0.2 | 0.0 | 0.1 | -0.04 | 20.0 | 10.0 | 0.0 | 50.0128004E+00 | DC220MV |  | -0.040000023293500005 | 3.2753934919830581e-08 | 1517785918.516 | 1517785934.009 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04 | -0.04
| 0.1 | -0.04 | 0.2 | -0.01 | 0.1 | -0.03 | 20.0 | 10.0 | 0.0 | 49.9991510E+00 | DC220MV |  | -0.029999968499500002 | 3.972328353441309e-08 | 1517785948.862 | 1517785964.365 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03
| 0.1 | -0.03 | 0.2 | 0.0 | 0.1 | -0.03 | 20.0 | 10.0 | 0.0 | 49.9882292E+00 | DC220MV |  | -0.030000014092 | 3.3773822200261172e-08 | 1517785979.211 | 1517785994.71 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03 | -0.03
| 0.1 | -0.03 | 0.2 | -0.01 | 0.1 | -0.02 | 20.0 | 10.0 | 0.0 | 49.9833546E+00 | DC220MV |  | -0.019999981138999999 | 4.469187587251419e-08 | 1517786009.574 | 1517786025.102 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02
| 0.1 | -0.02 | 0.2 | 0.0 | 0.1 | -0.02 | 20.0 | 10.0 | 0.0 | 50.0355205E+00 | DC220MV |  | -0.020000044605499999 | 3.7509690411292237e-08 | 1517786039.937 | 1517786055.442 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02 | -0.02
| 0.1 | -0.02 | 0.2 | -0.01 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 50.0291018E+00 | DC220MV |  | -0.00999999686575 | 4.0725450108012629e-08 | 1517786070.29 | 1517786085.788 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.0099999 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 49.9773028E+00 | DC220MV |  | -0.010000049099099999 | 4.5055801112554369e-08 | 1517786100.637 | 1517786116.135 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.0099999 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0072935E+00 | DC220MV |  | -2.0523214307800001e-08 | 4.6370117516341986e-08 | 1517786130.996 | 1517786146.579 | -4.5869e-08 | -2.5433e-08 | 1.5584e-08 | -3.9841e-08 | -5.072e-08 | -1.7936e-08 | -3.249e-08 | -3.2343e-09 | 7.2037e-09 | -8.7915e-08 | -1.2364e-07 | 6.1011e-08 | -1.676e-08 | 3.4842e-08 | 5.7336e-09 | -5.8512e-08 | 7.88e-08 | -2.2346e-08 | -2.8962e-08 | -5.9982e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0060951E+00 | DC220MV |  | -7.9675725903400012e-08 | 4.3119153710930346e-08 | 1517786161.44 | 1517786177.009 | -6.5594e-08 | -1.7899e-07 | -1.7546e-07 | -1.0457e-07 | -4.8975e-08 | -5.5152e-08 | -4.3533e-08 | -7.9124e-08 | -9.0007e-08 | -4.7063e-08 | -6.427e-08 | -3.8827e-08 | -5.9711e-08 | 1.1766e-09 | -1.2104e-07 | -5.9123e-08 | -9.5743e-08 | -6.5888e-08 | -1.231e-07 | -7.8536e-08
| 0.1 | -0.01 | 0.2 | -0.01 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9702253E+00 | DC220MV |  | -2.8481300013324998e-08 | 4.4596959416491562e-08 | 1517786191.863 | 1517786207.436 | -2.6621e-08 | -1.912e-09 | 8.8246e-10 | -9.3982e-08 | -1.4708e-08 | -8.6922e-08 | -8.6775e-09 | -1.1163e-07 | 8.9275e-08 | -4.5594e-09 | -3.1474e-08 | -2.7209e-08 | -8.0598e-08 | -4.2211e-08 | -3.1327e-08 | 1.7943e-08 | -5.0741e-08 | 1.0295e-08 | -6.339e-08 | -1.206e-08
| 0.1 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9962267E+00 | DC220MV |  | -9.2560134091500006e-08 | 4.3912666052063486e-08 | 1517786222.304 | 1517786237.886 | -1.1391e-07 | -1.0979e-07 | -4.1301e-08 | -1.0068e-07 | -8.0251e-08 | -1.1112e-07 | -8.3337e-08 | -1.07e-07 | -7.2755e-08 | -1.0494e-07 | -4.7033e-08 | -1.8887e-07 | -4.0566e-08 | -1.6859e-07 | -5.497e-08 | -8.4513e-08 | -3.8949e-08 | -4.7915e-08 | -1.7829e-07 | -7.6429e-08
| 0.1 | -0.01 | 0.2 | 0.0 | 0.1 | -0.01 | 20.0 | 10.0 | 0.0 | 49.9968917E+00 | DC220MV |  | -0.0100000728115 | 3.4968301828048265e-08 | 1517786252.735 | 1517786268.239 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
| 2 | 0.05 | 0.2 | 0.0 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9788721E+00 | DC220MV |  | 0.049999896004999995 | 5.1849129741976074e-08 | 1517786283.47 | 1517786298.926 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0570178E+00 | DC220MV |  | -2.3452725017000002e-07 | 5.4525701995723932e-08 | 1517786313.749 | 1517786329.25 | -2.1622e-07 | -2.3239e-07 | -2.7211e-07 | -2.7652e-07 | -2.2945e-07 | -1.7944e-07 | -1.6621e-07 | -1.3238e-07 | -1.7503e-07 | -3.2065e-07 | -2.9711e-07 | -2.6475e-07 | -2.6475e-07 | -2.6181e-07 | -2.5446e-07 | -3.1917e-07 | -1.7944e-07 | -1.4267e-07 | -2.7064e-07 | -2.3534e-07
| 2 | 0.05 | 0.2 | 0.05 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9946138E+00 | DC220MV |  | -3.1271979179999995e-07 | 4.2425682332449437e-08 | 1517786344.071 | 1517786359.549 | -2.0171e-07 | -3.0919e-07 | -3.4158e-07 | -3.3274e-07 | -2.9299e-07 | -2.9152e-07 | -3.3127e-07 | -2.8563e-07 | -3.1066e-07 | -3.1949e-07 | -3.7102e-07 | -2.5913e-07 | -3.4747e-07 | -3.2833e-07 | -2.9888e-07 | -3.6072e-07 | -3.7986e-07 | -2.7532e-07 | -2.6207e-07 | -3.5483e-07
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0138529E+00 | DC220MV |  | -2.5268659550499994e-07 | 4.201135920315177e-08 | 1517786374.389 | 1517786389.873 | -2.4577e-07 | -1.9573e-07 | -2.2517e-07 | -2.9581e-07 | -2.649e-07 | -3.2524e-07 | -2.5901e-07 | -2.5901e-07 | -2.7079e-07 | -1.972e-07 | -1.7513e-07 | -2.5166e-07 | -3.0611e-07 | -2.9875e-07 | -1.9132e-07 | -2.8256e-07 | -2.3547e-07 | -2.0898e-07 | -3.0317e-07 | -2.6196e-07
| 2 | 0.05 | 0.2 | 0.05 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9953495E+00 | DC220MV |  | -2.9206615528e-07 | 7.3293377630074826e-08 | 1517786404.719 | 1517786420.211 | -3.1487e-07 | -4.0168e-07 | -3.6931e-07 | -3.9727e-07 | -3.443e-07 | -1.1035e-07 | -2.8397e-07 | -2.7515e-07 | -3.7961e-07 | -2.5455e-07 | -2.3689e-07 | -2.2365e-07 | -2.0599e-07 | -3.1487e-07 | -2.3248e-07 | -2.413e-07 | -2.4719e-07 | -2.9133e-07 | -3.5019e-07 | -3.6637e-07
| 2 | 0.05 | 0.2 | 0.0 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0096811E+00 | DC220MV |  | 0.049999848339500004 | 6.7333659003643492e-08 | 1517786435.048 | 1517786450.509 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 2 | 0.1 | 0.2 | 0.05 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9836139E+00 | DC220MV |  | 0.049999773034500004 | 4.6634055846974343e-08 | 1517786465.351 | 1517786480.8 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 2 | 0.1 | 0.2 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0166990E+00 | DC220MV |  | 0.099999848322500001 | 5.7828290900982302e-08 | 1517786495.63 | 1517786511.088 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.15 | 0.2 | 0.05 | 1 | 0.09999999999999999 | 20.0 | 10.0 | 0.0 | 50.0254658E+00 | DC220MV |  | 0.099999878165 | 3.3220960025156565e-08 | 1517786525.913 | 1517786541.352 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.15 | 0.2 | 0.0 | 1 | 0.15 | 20.0 | 10.0 | 0.0 | 49.9922606E+00 | DC220MV |  | 0.14999992200000001 | 6.1827170401397779e-08 | 1517786556.203 | 1517786571.631 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
| 2 | 0.2 | 0.2 | 0.05 | 1 | 0.15 | 20.0 | 10.0 | 0.0 | 49.9882056E+00 | DC220MV |  | 0.14999988348499999 | 6.5880560680499862e-08 | 1517786586.474 | 1517786601.917 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
| 2 | 0.2 | 0.2 | 0.0 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9777552E+00 | DC220MV |  | 0.20000001754999999 | 6.4248031095243562e-08 | 1517786616.756 | 1517786632.183 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.25 | 0.2 | 0.05 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0044485E+00 | DC220MV |  | 0.19999996183999996 | 6.6549818932844993e-08 | 1517786647.025 | 1517786662.475 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.25 | 0.2 | 0.0 | 1 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0106815E+00 | DC220MV |  | 0.25000011595500005 | 5.6741179712559862e-08 | 1517786677.326 | 1517786692.761 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
| 2 | 0.3 | 0.2 | 0.05 | 1 | 0.25 | 20.0 | 10.0 | 0.0 | 49.9904499E+00 | DC220MV |  | 0.25000002289000001 | 6.2419082817269656e-08 | 1517786707.609 | 1517786723.042 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
| 2 | 0.3 | 0.2 | 0.0 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0300369E+00 | DC220MV |  | 0.30000019824000002 | 7.3979952686065164e-08 | 1517786737.889 | 1517786753.324 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.35 | 0.2 | 0.05 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9751350E+00 | DC220MV |  | 0.30000013667499997 | 4.4940069808788546e-08 | 1517786768.175 | 1517786783.588 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.35 | 0.2 | 0.0 | 1 | 0.35 | 20.0 | 10.0 | 0.0 | 50.0485471E+00 | DC220MV |  | 0.35000027075500001 | 5.910938567527118e-08 | 1517786798.447 | 1517786813.883 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35
| 2 | 0.4 | 0.2 | 0.05 | 1 | 0.35 | 20.0 | 10.0 | 0.0 | 49.9945572E+00 | DC220MV |  | 0.35000020514500008 | 5.8274792787939784e-08 | 1517786828.737 | 1517786844.137 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35
| 2 | 0.4 | 0.2 | 0.0 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0382368E+00 | DC220MV |  | 0.40000038993500003 | 5.8383904247906768e-08 | 1517786858.98 | 1517786874.393 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 2 | 0.45 | 0.2 | 0.05 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9979906E+00 | DC220MV |  | 0.40000032001500002 | 5.8505899489262382e-08 | 1517786889.218 | 1517786904.625 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 2 | 0.45 | 0.2 | 0.0 | 1 | 0.45 | 20.0 | 10.0 | 0.0 | 49.9767043E+00 | DC220MV |  | 0.45000042833999998 | 4.8645209426217437e-08 | 1517786919.467 | 1517786934.899 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45
| 2 | 0.5 | 0.2 | 0.05 | 1 | 0.45 | 20.0 | 10.0 | 0.0 | 50.0076285E+00 | DC220MV |  | 0.45000037144499999 | 8.6521947935726156e-08 | 1517786949.744 | 1517786965.179 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45
| 2 | 0.5 | 0.2 | 0.0 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 49.9847735E+00 | DC220MV |  | 0.50000051378499999 | 6.0840046641380695e-08 | 1517786980.028 | 1517786995.459 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 2 | 0.5 | 0.2 | 0.05 | 1 | 0.45 | 20.0 | 10.0 | 0.0 | 49.9888845E+00 | DC220MV |  | 0.45000031942499996 | 5.5316334610197524e-08 | 1517787010.288 | 1517787025.742 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45
| 2 | 0.45 | 0.2 | 0.0 | 1 | 0.45 | 20.0 | 10.0 | 0.0 | 50.0445207E+00 | DC220MV |  | 0.45000041175999994 | 4.0553019612566335e-08 | 1517787040.582 | 1517787056.004 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45
| 2 | 0.45 | 0.2 | 0.05 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0044060E+00 | DC220MV |  | 0.40000030812499998 | 5.4903058884853991e-08 | 1517787070.841 | 1517787086.244 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 2 | 0.4 | 0.2 | 0.0 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0000566E+00 | DC220MV |  | 0.4000003766350001 | 5.7004826770342085e-08 | 1517787101.096 | 1517787116.518 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 2 | 0.4 | 0.2 | 0.05 | 1 | 0.35 | 20.0 | 10.0 | 0.0 | 50.1017538E+00 | DC220MV |  | 0.35000027012000001 | 4.7085248213350097e-08 | 1517787131.369 | 1517787146.801 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35
| 2 | 0.35 | 0.2 | 0.0 | 1 | 0.35 | 20.0 | 10.0 | 0.0 | 49.9654013E+00 | DC220MV |  | 0.35000031611500004 | 8.1119654056782432e-08 | 1517787161.647 | 1517787177.026 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35
| 2 | 0.35 | 0.2 | 0.05 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0125126E+00 | DC220MV |  | 0.30000020586499992 | 4.7474122155877433e-08 | 1517787191.863 | 1517787207.298 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.3 | 0.2 | 0.0 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0202724E+00 | DC220MV |  | 0.30000022878999999 | 6.8698201583141247e-08 | 1517787222.172 | 1517787237.569 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.3 | 0.2 | 0.05 | 1 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0167934E+00 | DC220MV |  | 0.25000008574500004 | 4.5808006666742955e-08 | 1517787252.413 | 1517787267.821 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
| 2 | 0.25 | 0.2 | 0.0 | 1 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0734521E+00 | DC220MV |  | 0.25000011939000005 | 5.6102343088897307e-08 | 1517787282.677 | 1517787298.103 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
| 2 | 0.25 | 0.2 | 0.05 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0017453E+00 | DC220MV |  | 0.20000005344999999 | 5.9807361586988168e-08 | 1517787312.945 | 1517787328.34 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.2 | 0.2 | 0.0 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9657311E+00 | DC220MV |  | 0.20000006452500002 | 5.1964140278773961e-08 | 1517787343.195 | 1517787358.615 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.2 | 0.2 | 0.05 | 1 | 0.15 | 20.0 | 10.0 | 0.0 | 49.9844577E+00 | DC220MV |  | 0.149999878335 | 5.8630574572478428e-08 | 1517787373.466 | 1517787388.919 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
| 2 | 0.15 | 0.2 | 0.0 | 1 | 0.15 | 20.0 | 10.0 | 0.0 | 49.9521308E+00 | DC220MV |  | 0.14999998270500001 | 4.9431705158559576e-08 | 1517787403.748 | 1517787419.202 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
| 2 | 0.15 | 0.2 | 0.05 | 1 | 0.09999999999999999 | 20.0 | 10.0 | 0.0 | 49.9927558E+00 | DC220MV |  | 0.099999853702500002 | 5.074573557233937e-08 | 1517787434.038 | 1517787449.514 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.1 | 0.2 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0094924E+00 | DC220MV |  | 0.099999879147499998 | 5.6086226282723189e-08 | 1517787464.353 | 1517787479.809 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.1 | 0.2 | 0.05 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0134612E+00 | DC220MV |  | 0.049999722278000001 | 6.4150335665073845e-08 | 1517787494.654 | 1517787510.094 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 2 | 0.05 | 0.2 | 0.0 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0802134E+00 | DC220MV |  | 0.049999795553999998 | 5.0544742198451472e-08 | 1517787524.952 | 1517787540.392 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| 2 | 0.05 | 0.2 | 0.05 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0121161E+00 | DC220MV |  | -3.2194729661000002e-07 | 5.5688543971621177e-08 | 1517787555.213 | 1517787570.669 | -3.2268e-07 | -3.6836e-07 | -2.608e-07 | -2.6375e-07 | -4.0814e-07 | -2.8437e-07 | -3.0206e-07 | -4.1846e-07 | -3.4773e-07 | -1.9449e-07 | -3.0795e-07 | -2.7259e-07 | -3.4921e-07 | -3.0942e-07 | -3.8457e-07 | -3.2268e-07 | -2.9322e-07 | -2.7848e-07 | -4.0225e-07 | -3.4773e-07
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9936423E+00 | DC220MV |  | -2.7894237900999995e-07 | 5.7848514711534411e-08 | 1517787585.5 | 1517787600.98 | -3.0912e-07 | -3.4297e-07 | -2.7526e-07 | -3.1059e-07 | -1.8105e-07 | -2.0608e-07 | -3.1942e-07 | -2.8851e-07 | -2.0313e-07 | -2.8557e-07 | -2.1049e-07 | -3.4886e-07 | -3.8125e-07 | -2.3257e-07 | -2.944e-07 | -2.2963e-07 | -3.6653e-07 | -2.1491e-07 | -3.1795e-07 | -2.6054e-07
| 2 | 0.05 | 0.2 | 0.05 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9771237E+00 | DC220MV |  | -3.1492112798499998e-07 | 8.39969593302784e-08 | 1517787615.814 | 1517787631.301 | -4.5472e-07 | -2.2515e-07 | -3.37e-07 | -3.0168e-07 | -2.9432e-07 | -1.7218e-07 | -3.782e-07 | -5.4596e-07 | -2.4428e-07 | -2.6047e-07 | -3.6054e-07 | -2.487e-07 | -2.3398e-07 | -4.0175e-07 | -3.2964e-07 | -3.2817e-07 | -2.8696e-07 | -2.8549e-07 | -2.6047e-07 | -3.4877e-07
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0190828E+00 | DC220MV |  | -2.7811170003749997e-07 | 8.0305794759503327e-08 | 1517787646.124 | 1517787661.621 | -2.6627e-07 | -3.663e-07 | -3.5748e-07 | -3.2658e-07 | -2.4567e-07 | -3.663e-07 | -2.2361e-07 | -3.0599e-07 | -2.2655e-07 | -3.4571e-07 | -1.6623e-07 | -2.9716e-07 | -9.7093e-08 | -2.751e-07 | -3.0893e-07 | -2.1037e-07 | -2.0595e-07 | -2.5891e-07 | -2.4715e-07 | -4.6487e-07
| 2 | 0.05 | 0.2 | 0.0 | 1 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9834489E+00 | DC220MV |  | 0.0499998039405 | 6.2624874808353306e-08 | 1517787676.454 | 1517787691.891 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | 0.0 | 1 | -0.005 | 20.0 | 10.0 | 0.0 | 50.0281101E+00 | DC220MV |  | -0.050000259041999992 | 7.554337499730809e-08 | 1517787706.753 | 1517787722.203 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9596599E+00 | DC220MV |  | -1.1858966726874999e-07 | 5.4639041014612688e-08 | 1517787737.035 | 1517787752.516 | -6.4779e-08 | -2.2525e-07 | -1.1189e-07 | -8.539e-08 | -1.8256e-07 | -1.0306e-07 | -1.6195e-07 | -2.3703e-07 | -1.1042e-07 | -6.9196e-08 | -1.5164e-07 | -7.9501e-08 | -9.7168e-08 | -1.3692e-07 | -1.5017e-07 | -1.1189e-07 | -7.0668e-08 | -8.0974e-08 | -1.3692e-07 | -4.4167e-09
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | -0.005 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0383076E+00 | DC220MV |  | -0.045000238983500002 | 8.455633260036301e-08 | 1517787767.355 | 1517787782.823 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0030332E+00 | DC220MV |  | -1.4584092614069999e-07 | 7.9306429018141504e-08 | 1517787797.669 | 1517787813.136 | -2.1655e-07 | 5.156e-08 | -1.9887e-07 | -2.0624e-08 | -1.7678e-07 | -4.4194e-09 | -1.37e-07 | -1.753e-07 | -1.3553e-07 | -2.5927e-07 | -1.37e-07 | -2.3276e-07 | -1.0901e-07 | -2.1655e-07 | -2.0771e-07 | -7.2184e-08 | -1.5321e-07 | -2.1655e-07 | -1.5763e-07 | -1.4142e-07
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | -0.005 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0032832E+00 | DC220MV |  | -0.045000267003999998 | 6.3471826774457052e-08 | 1517787827.965 | 1517787843.415 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045
| 2 | -0.05 | 0.2 | 0.0 | 1 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9823694E+00 | DC220MV |  | -0.050000285157000002 | 6.6391545779336718e-08 | 1517787858.262 | 1517787873.701 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 2 | -0.1 | 0.2 | -0.005 | 1 | -0.095 | 20.0 | 10.0 | 0.0 | 50.0187948E+00 | DC220MV |  | -0.095000336541500002 | 3.9743303498545006e-08 | 1517787888.544 | 1517787903.994 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095
| 2 | -0.1 | 0.2 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 50.0141125E+00 | DC220MV |  | -0.100000365625 | 4.5943137409988389e-08 | 1517787918.815 | 1517787934.229 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 2 | -0.15 | 0.2 | -0.005 | 1 | -0.145 | 20.0 | 10.0 | 0.0 | 49.9566838E+00 | DC220MV |  | -0.14500043466499998 | 5.2573960046483539e-08 | 1517787949.066 | 1517787964.495 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145
| 2 | -0.15 | 0.2 | 0.0 | 1 | -0.15 | 20.0 | 10.0 | 0.0 | 50.0366306E+00 | DC220MV |  | -0.15000051998500002 | 4.7495497419563264e-08 | 1517787979.326 | 1517787994.768 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15
| 2 | -0.2 | 0.2 | -0.005 | 1 | -0.195 | 20.0 | 10.0 | 0.0 | 49.9870788E+00 | DC220MV |  | -0.195000541975 | 5.4977630679017304e-08 | 1517788009.614 | 1517788025.038 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195
| 2 | -0.2 | 0.2 | 0.0 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9723172E+00 | DC220MV |  | -0.20000060830500002 | 4.2952921609742933e-08 | 1517788039.884 | 1517788055.303 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 2 | -0.25 | 0.2 | -0.005 | 1 | -0.245 | 20.0 | 10.0 | 0.0 | 49.9881019E+00 | DC220MV |  | -0.24500069026499999 | 5.4505772858334734e-08 | 1517788070.139 | 1517788085.591 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245
| 2 | -0.25 | 0.2 | 0.0 | 1 | -0.25 | 20.0 | 10.0 | 0.0 | 49.9933782E+00 | DC220MV |  | -0.25000067648499996 | 4.5305896690650563e-08 | 1517788100.431 | 1517788115.866 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25
| 2 | -0.3 | 0.2 | -0.005 | 1 | -0.295 | 20.0 | 10.0 | 0.0 | 49.9714032E+00 | DC220MV |  | -0.29500073450500003 | 5.8251665002368258e-08 | 1517788130.689 | 1517788146.121 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295
| 2 | -0.3 | 0.2 | 0.0 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0073784E+00 | DC220MV |  | -0.30000077547000004 | 6.5090284225917061e-08 | 1517788160.966 | 1517788176.392 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 2 | -0.35 | 0.2 | -0.005 | 1 | -0.345 | 20.0 | 10.0 | 0.0 | 50.0298244E+00 | DC220MV |  | -0.34500089069000001 | 6.1147910028417072e-08 | 1517788191.244 | 1517788206.67 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345
| 2 | -0.35 | 0.2 | 0.0 | 1 | -0.35 | 20.0 | 10.0 | 0.0 | 50.0136028E+00 | DC220MV |  | -0.35000090103999998 | 5.2412025338583359e-08 | 1517788221.489 | 1517788236.899 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35
| 2 | -0.4 | 0.2 | -0.005 | 1 | -0.395 | 20.0 | 10.0 | 0.0 | 49.9668757E+00 | DC220MV |  | -0.39500093720500001 | 5.6845549297105412e-08 | 1517788251.729 | 1517788267.149 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395
| 2 | -0.4 | 0.2 | 0.0 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0184596E+00 | DC220MV |  | -0.40000095039500005 | 6.7974859141555183e-08 | 1517788281.988 | 1517788297.433 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 2 | -0.45 | 0.2 | -0.005 | 1 | -0.445 | 20.0 | 10.0 | 0.0 | 49.9878190E+00 | DC220MV |  | -0.44500106138000006 | 6.9362919492408728e-08 | 1517788312.261 | 1517788327.691 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445
| 2 | -0.45 | 0.2 | 0.0 | 1 | -0.45 | 20.0 | 10.0 | 0.0 | 49.9974105E+00 | DC220MV |  | -0.45000108500499997 | 4.6314160632904889e-08 | 1517788342.523 | 1517788357.959 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45
| 2 | -0.5 | 0.2 | -0.005 | 1 | -0.495 | 20.0 | 10.0 | 0.0 | 49.9917561E+00 | DC220MV |  | -0.49500112484000003 | 5.1525065747384395e-08 | 1517788372.797 | 1517788388.232 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495
| 2 | -0.5 | 0.2 | 0.0 | 1 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9941705E+00 | DC220MV |  | -0.50000118479 | 5.6252812371942332e-08 | 1517788403.082 | 1517788418.518 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 2 | -0.5 | 0.2 | -0.005 | 1 | -0.495 | 20.0 | 10.0 | 0.0 | 49.9944299E+00 | DC220MV |  | -0.49500121161499999 | 5.6926252943478187e-08 | 1517788433.34 | 1517788448.788 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495 | -0.495
| 2 | -0.45 | 0.2 | 0.0 | 1 | -0.45 | 20.0 | 10.0 | 0.0 | 50.0098038E+00 | DC220MV |  | -0.45000116108000005 | 3.9190210514372672e-08 | 1517788463.625 | 1517788479.06 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45
| 2 | -0.45 | 0.2 | -0.005 | 1 | -0.445 | 20.0 | 10.0 | 0.0 | 49.9072338E+00 | DC220MV |  | -0.44500110728999998 | 7.6121198753635509e-08 | 1517788493.912 | 1517788509.338 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445 | -0.445
| 2 | -0.4 | 0.2 | 0.0 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0175769E+00 | DC220MV |  | -0.40000101520999998 | 5.5535420231380955e-08 | 1517788524.172 | 1517788539.63 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 2 | -0.4 | 0.2 | -0.005 | 1 | -0.395 | 20.0 | 10.0 | 0.0 | 49.9802531E+00 | DC220MV |  | -0.39500099284500001 | 7.9537478434823165e-08 | 1517788554.477 | 1517788569.907 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395 | -0.395
| 2 | -0.35 | 0.2 | 0.0 | 1 | -0.35 | 20.0 | 10.0 | 0.0 | 49.9852072E+00 | DC220MV |  | -0.35000093934499998 | 1.1284970303161499e-07 | 1517788584.759 | 1517788600.202 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35
| 2 | -0.35 | 0.2 | -0.005 | 1 | -0.345 | 20.0 | 10.0 | 0.0 | 50.0083221E+00 | DC220MV |  | -0.34500143756000001 | 7.6297274748027263e-07 | 1517788615.047 | 1517788630.498 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345 | -0.345
| 2 | -0.3 | 0.2 | 0.0 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0031889E+00 | DC220MV |  | -0.30000081231999998 | 7.108321602144361e-08 | 1517788645.348 | 1517788660.774 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 2 | -0.3 | 0.2 | -0.005 | 1 | -0.295 | 20.0 | 10.0 | 0.0 | 49.9866074E+00 | DC220MV |  | -0.29500075913000001 | 5.3840673295233615e-08 | 1517788675.611 | 1517788691.044 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295 | -0.295
| 2 | -0.25 | 0.2 | 0.0 | 1 | -0.25 | 20.0 | 10.0 | 0.0 | 49.9779673E+00 | DC220MV |  | -0.25000071863000001 | 5.7704177489813814e-08 | 1517788705.899 | 1517788721.328 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25
| 2 | -0.25 | 0.2 | -0.005 | 1 | -0.245 | 20.0 | 10.0 | 0.0 | 50.0223874E+00 | DC220MV |  | -0.24500064290000001 | 4.1475788114797309e-08 | 1517788736.182 | 1517788751.624 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245 | -0.245
| 2 | -0.2 | 0.2 | 0.0 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9886158E+00 | DC220MV |  | -0.20000061821999998 | 4.9514024273664279e-08 | 1517788766.48 | 1517788781.894 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 2 | -0.2 | 0.2 | -0.005 | 1 | -0.195 | 20.0 | 10.0 | 0.0 | 50.0708928E+00 | DC220MV |  | -0.19500056047499997 | 5.3957046573747587e-08 | 1517788796.723 | 1517788812.153 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195 | -0.195
| 2 | -0.15 | 0.2 | 0.0 | 1 | -0.15 | 20.0 | 10.0 | 0.0 | 50.0031370E+00 | DC220MV |  | -0.15000052017499999 | 8.1404587554340028e-08 | 1517788826.97 | 1517788842.352 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15
| 2 | -0.15 | 0.2 | -0.005 | 1 | -0.145 | 20.0 | 10.0 | 0.0 | 49.9789475E+00 | DC220MV |  | -0.14500049051 | 6.0253961696462354e-08 | 1517788857.174 | 1517788872.587 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145 | -0.145
| 2 | -0.1 | 0.2 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 50.0093650E+00 | DC220MV |  | -0.100000415675 | 3.6109401476446293e-08 | 1517788887.439 | 1517788902.87 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 2 | -0.1 | 0.2 | -0.005 | 1 | -0.095 | 20.0 | 10.0 | 0.0 | 50.0268446E+00 | DC220MV |  | -0.095000398032 | 5.5814711735323083e-08 | 1517788917.716 | 1517788933.152 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095001 | -0.095 | -0.095 | -0.095 | -0.095 | -0.095
| 2 | -0.05 | 0.2 | 0.0 | 1 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9628955E+00 | DC220MV |  | -0.0500003241875 | 5.2075743094040778e-08 | 1517788947.99 | 1517788963.399 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 2 | -0.05 | 0.2 | -0.005 | 1 | -0.04500000000000001 | 20.0 | 10.0 | 0.0 | 50.0037975E+00 | DC220MV |  | -0.045000273627499997 | 5.6368016274777214e-08 | 1517788978.228 | 1517788993.674 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0456738E+00 | DC220MV |  | -2.00399609583e-07 | 5.5260651676430591e-08 | 1517789008.507 | 1517789023.989 | -1.6485e-07 | -2.1048e-07 | -1.7221e-07 | -1.8104e-07 | -1.6927e-07 | -3.0321e-07 | -2.1637e-07 | -2.0312e-07 | -3.2676e-07 | -2.4581e-07 | -2.149e-07 | -1.9871e-07 | -7.6539e-08 | -1.8546e-07 | -2.3256e-07 | -1.678e-07 | -2.6053e-07 | -1.5161e-07 | -1.31e-07 | -1.9576e-07
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | -0.005 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9594951E+00 | DC220MV |  | -0.045000232076500009 | 6.0020413716656782e-08 | 1517789038.84 | 1517789054.261 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045 | -0.045
| 2 | 0.0 | 0.2 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0097425E+00 | DC220MV |  | -1.5957939912444999e-07 | 7.2478912817492113e-08 | 1517789069.109 | 1517789084.614 | -2.6915e-07 | -2.1768e-07 | -3.1033e-07 | -1.4267e-07 | -1.2649e-07 | -1.1325e-07 | -1.5149e-07 | -1.6326e-07 | -2.3091e-07 | -1.059e-07 | -5.4419e-08 | -1.7208e-07 | -5.1477e-08 | -5.8831e-09 | -1.7208e-07 | -1.6179e-07 | -1.2796e-07 | -2.1032e-07 | -1.7502e-07 | -2.2944e-07
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 2 | -0.05 | 0.2 | 0.0 | 1 | -0.005 | 20.0 | 10.0 | 0.0 | 49.9737873E+00 | DC220MV |  | -0.050000300897500002 | 5.6828135273816448e-08 | 1517789099.435 | 1517789114.86 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05 | -0.05
| 2 | 0.1 | 2.0 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9775761E+00 | DC2_2V |  | 0.099999975300000016 | 5.7720500083854148e-08 | 1517789130.022 | 1517789145.48 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9463736E+00 | DC2_2V |  | -1.4756620991250002e-07 | 5.9066727879536981e-08 | 1517789160.309 | 1517789175.805 | -6.3264e-08 | -1.9126e-07 | -1.7214e-07 | -2.3246e-07 | -1.2653e-07 | -2.1039e-07 | -2.2804e-07 | -6.0321e-08 | -2.0597e-07 | -1.7066e-07 | -1.5742e-07 | -1.942e-07 | -1.8243e-07 | -6.9149e-08 | -1.2358e-07 | -1.5154e-07 | -1.383e-07 | -9.8574e-08 | -1.6037e-07 | -1.4712e-08
| 2 | 0.1 | 2.0 | 0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0351946E+00 | DC2_2V |  | -5.5065823766000002e-07 | 5.564926310730967e-08 | 1517789190.657 | 1517789206.175 | -6.0556e-07 | -4.7033e-07 | -6.2172e-07 | -5.159e-07 | -5.012e-07 | -5.2031e-07 | -5.6146e-07 | -4.8944e-07 | -4.8356e-07 | -4.9385e-07 | -6.1731e-07 | -5.2913e-07 | -6.2025e-07 | -5.6293e-07 | -6.0556e-07 | -6.4377e-07 | -5.6146e-07 | -6.0409e-07 | -4.7033e-07 | -5.35e-07
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0032880E+00 | DC2_2V |  | -1.5483100244350002e-07 | 5.6661258930952316e-08 | 1517789221.03 | 1517789236.505 | -1.6639e-07 | -1.7081e-07 | -7.3624e-08 | -2.3412e-07 | -1.1927e-07 | -1.9289e-07 | -2.0615e-07 | -1.1338e-07 | -8.3931e-08 | -2.459e-07 | -1.178e-07 | -1.8553e-07 | -9.2766e-08 | -2.2234e-07 | -1.178e-07 | -1.1485e-07 | -1.2516e-07 | -1.5167e-07 | -9.5711e-08 | -2.6652e-07
| 2 | 0.1 | 2.0 | 0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0206925E+00 | DC2_2V |  | -5.9960426974499996e-07 | 7.1547517791802119e-08 | 1517789251.341 | 1517789266.836 | -5.885e-07 | -4.8551e-07 | -6.5912e-07 | -6.1939e-07 | -5.3995e-07 | -6.547e-07 | -6.2233e-07 | -5.6937e-07 | -6.9148e-07 | -6.753e-07 | -5.3995e-07 | -5.8997e-07 | -6.9737e-07 | -6.0321e-07 | -7.1208e-07 | -4.3402e-07 | -6.1939e-07 | -6.2381e-07 | -5.3847e-07 | -5.2818e-07
| 2 | 0.1 | 2.0 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9721146E+00 | DC2_2V |  | 0.099999922465000002 | 5.0370027843377476e-08 | 1517789281.665 | 1517789297.111 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.2 | 2.0 | 0.1 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9605312E+00 | DC2_2V |  | 0.099999550206999993 | 6.7396343083412235e-08 | 1517789311.938 | 1517789327.395 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.099999 | 0.1 | 0.1 | 0.099999 | 0.099999 | 0.099999 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.2 | 2.0 | 0.0 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9860416E+00 | DC2_2V |  | 0.20000013556000001 | 4.4766386943340775e-08 | 1517789342.252 | 1517789357.692 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.3 | 2.0 | 0.1 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9907517E+00 | DC2_2V |  | 0.19999968636499998 | 6.5203767335841742e-08 | 1517789372.521 | 1517789387.973 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.3 | 2.0 | 0.0 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0868203E+00 | DC2_2V |  | 0.30000030820500001 | 8.0325248056581602e-08 | 1517789402.815 | 1517789418.238 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.4 | 2.0 | 0.1 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0245262E+00 | DC2_2V |  | 0.29999985849499999 | 7.9523666136629443e-08 | 1517789433.106 | 1517789448.495 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.4 | 2.0 | 0.0 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0052222E+00 | DC2_2V |  | 0.40000041174999995 | 4.4327502754546649e-08 | 1517789463.32 | 1517789478.747 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 2 | 0.5 | 2.0 | 0.1 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9800127E+00 | DC2_2V |  | 0.40000001459500006 | 7.4134037226287693e-08 | 1517789493.603 | 1517789509.02 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 2 | 0.5 | 2.0 | 0.0 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 49.9986746E+00 | DC2_2V |  | 0.50000060268500002 | 6.2632182425481682e-08 | 1517789523.882 | 1517789539.313 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 2 | 0.6 | 2.0 | 0.1 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0298197E+00 | DC2_2V |  | 0.50000024165000001 | 5.5864367001685448e-08 | 1517789554.165 | 1517789569.6 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 2 | 0.6 | 2.0 | 0.0 | 1 | 0.6 | 20.0 | 10.0 | 0.0 | 49.8950415E+00 | DC2_2V |  | 0.60000077062000001 | 6.8456516136830384e-08 | 1517789584.443 | 1517789599.852 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
| 2 | 0.7 | 2.0 | 0.1 | 1 | 0.6 | 20.0 | 10.0 | 0.0 | 49.9626977E+00 | DC2_2V |  | 0.60000034195000007 | 7.4966802668019504e-08 | 1517789614.685 | 1517789630.159 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
| 2 | 0.7 | 2.0 | 0.0 | 1 | 0.7 | 20.0 | 10.0 | 0.0 | 50.0024152E+00 | DC2_2V |  | 0.70000096308000004 | 6.4724775794750047e-08 | 1517789644.984 | 1517789660.431 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7
| 2 | 0.8 | 2.0 | 0.1 | 1 | 0.7000000000000001 | 20.0 | 10.0 | 0.0 | 50.0515153E+00 | DC2_2V |  | 0.70000057937499993 | 7.09107740353939e-08 | 1517789675.29 | 1517789690.718 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7
| 2 | 0.8 | 2.0 | 0.0 | 1 | 0.8 | 20.0 | 10.0 | 0.0 | 50.1229431E+00 | DC2_2V |  | 0.80000106674500004 | 6.722108654188963e-08 | 1517789705.585 | 1517789720.991 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| 2 | 0.9 | 2.0 | 0.1 | 1 | 0.8 | 20.0 | 10.0 | 0.0 | 50.0197861E+00 | DC2_2V |  | 0.8000006033849999 | 7.9967351312859805e-08 | 1517789735.862 | 1517789751.224 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| 2 | 0.9 | 2.0 | 0.0 | 1 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0187806E+00 | DC2_2V |  | 0.90000112629999995 | 6.2197266821031176e-08 | 1517789766.061 | 1517789781.491 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9
| 2 | 1.0 | 2.0 | 0.1 | 1 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0595756E+00 | DC2_2V |  | 0.90000074898500004 | 6.398196835262503e-08 | 1517789796.34 | 1517789811.752 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9
| 2 | 1.0 | 2.0 | 0.0 | 1 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9931472E+00 | DC2_2V |  | 1.0000012543499999 | 6.5055572378620412e-08 | 1517789826.593 | 1517789841.987 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 2 | 1.0 | 2.0 | 0.1 | 1 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0236952E+00 | DC2_2V |  | 0.90000081202000004 | 5.7622726414096651e-08 | 1517789856.866 | 1517789872.285 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9
| 2 | 0.9 | 2.0 | 0.0 | 1 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0078739E+00 | DC2_2V |  | 0.90000122412500017 | 7.911715283948435e-08 | 1517789887.117 | 1517789902.53 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9
| 2 | 0.9 | 2.0 | 0.1 | 1 | 0.8 | 20.0 | 10.0 | 0.0 | 50.0498894E+00 | DC2_2V |  | 0.80000065931999997 | 7.8352317123434827e-08 | 1517789917.406 | 1517789932.833 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| 2 | 0.8 | 2.0 | 0.0 | 1 | 0.8 | 20.0 | 10.0 | 0.0 | 49.9555019E+00 | DC2_2V |  | 0.8000011455099999 | 8.3709395539466181e-08 | 1517789947.672 | 1517789963.084 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| 2 | 0.8 | 2.0 | 0.1 | 1 | 0.7000000000000001 | 20.0 | 10.0 | 0.0 | 49.9816482E+00 | DC2_2V |  | 0.70000048396000003 | 6.5578215882830418e-08 | 1517789977.926 | 1517789993.372 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7
| 2 | 0.7 | 2.0 | 0.0 | 1 | 0.7 | 20.0 | 10.0 | 0.0 | 49.9497063E+00 | DC2_2V |  | 0.70000087410999989 | 8.1092705585667192e-08 | 1517790008.232 | 1517790023.675 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7
| 2 | 0.7 | 2.0 | 0.1 | 1 | 0.6 | 20.0 | 10.0 | 0.0 | 50.0042079E+00 | DC2_2V |  | 0.60000039216000001 | 6.9426136299735201e-08 | 1517790038.539 | 1517790053.994 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
| 2 | 0.6 | 2.0 | 0.0 | 1 | 0.6 | 20.0 | 10.0 | 0.0 | 50.0332391E+00 | DC2_2V |  | 0.60000079011499996 | 7.8156607377146125e-08 | 1517790068.814 | 1517790084.239 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
| 2 | 0.6 | 2.0 | 0.1 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0130553E+00 | DC2_2V |  | 0.50000018880000019 | 5.3934747595925294e-08 | 1517790099.082 | 1517790114.494 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 2 | 0.5 | 2.0 | 0.0 | 1 | 0.5 | 20.0 | 10.0 | 0.0 | 49.9399543E+00 | DC2_2V |  | 0.50000068131500008 | 7.1791714520381117e-08 | 1517790129.312 | 1517790144.718 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 2 | 0.5 | 2.0 | 0.1 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0719714E+00 | DC2_2V |  | 0.40000002891000008 | 8.7913570626488905e-08 | 1517790159.555 | 1517790175.008 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 2 | 0.4 | 2.0 | 0.0 | 1 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9952127E+00 | DC2_2V |  | 0.40000049958000006 | 5.2759630402416802e-08 | 1517790189.841 | 1517790205.239 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 2 | 0.4 | 2.0 | 0.1 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9792162E+00 | DC2_2V |  | 0.29999985424500009 | 7.5185440573953366e-08 | 1517790220.064 | 1517790235.497 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.3 | 2.0 | 0.0 | 1 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0371266E+00 | DC2_2V |  | 0.30000029612500001 | 6.9887844970683929e-08 | 1517790250.349 | 1517790265.777 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 2 | 0.3 | 2.0 | 0.1 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9035496E+00 | DC2_2V |  | 0.19999968510500002 | 4.5475152284226694e-08 | 1517790280.613 | 1517790296.042 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.2 | 2.0 | 0.0 | 1 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0732345E+00 | DC2_2V |  | 0.20000012355500002 | 5.6873873395754618e-08 | 1517790310.876 | 1517790326.352 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 2 | 0.2 | 2.0 | 0.1 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9557891E+00 | DC2_2V |  | 0.099999605431499994 | 8.0795706277908224e-08 | 1517790341.182 | 1517790356.612 | 0.1 | 0.1 | 0.1 | 0.099999 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.099999 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.1 | 2.0 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0040145E+00 | DC2_2V |  | 0.10000000614849999 | 6.8905222900714804e-08 | 1517790371.46 | 1517790386.901 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | 0.1 | 2.0 | 0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0205320E+00 | DC2_2V |  | -6.9787619232999996e-07 | 6.2299572066524635e-08 | 1517790401.722 | 1517790417.216 | -7.1796e-07 | -6.8118e-07 | -6.3704e-07 | -7.0177e-07 | -7.4444e-07 | -6.7088e-07 | -6.2968e-07 | -6.6205e-07 | -5.7525e-07 | -7.415e-07 | -7.5915e-07 | -6.6794e-07 | -6.6058e-07 | -6.3557e-07 | -8.7244e-07 | -7.1354e-07 | -6.6941e-07 | -7.5621e-07 | -7.4591e-07 | -7.1502e-07
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0281196E+00 | DC2_2V |  | -1.7583517863599998e-07 | 7.3737700930272217e-08 | 1517790432.052 | 1517790447.513 | -2.151e-07 | -2.3867e-07 | -2.0479e-07 | -5.3038e-08 | -1.2228e-07 | -2.4751e-07 | -1.5617e-07 | -2.5488e-07 | -9.7236e-08 | -2.0331e-07 | -4.8618e-08 | -1.3554e-07 | -1.9005e-07 | -2.9907e-07 | -1.3259e-07 | -2.6372e-07 | -1.2965e-07 | -1.9742e-07 | -6.3351e-08 | -2.6372e-07
| 2 | 0.1 | 2.0 | 0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9795037E+00 | DC2_2V |  | -6.1144100098999988e-07 | 5.8843509545910504e-08 | 1517790462.343 | 1517790477.83 | -5.4896e-07 | -7.1675e-07 | -6.1519e-07 | -7.0056e-07 | -5.8576e-07 | -6.2697e-07 | -6.5787e-07 | -5.2247e-07 | -5.9606e-07 | -5.1806e-07 | -5.2395e-07 | -6.0048e-07 | -5.9312e-07 | -7.0056e-07 | -5.9606e-07 | -7.035e-07 | -5.8282e-07 | -6.0489e-07 | -6.358e-07 | -5.99e-07
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0193660E+00 | DC2_2V |  | -2.2317610497999999e-07 | 5.7209501792829003e-08 | 1517790492.656 | 1517790508.152 | -2.1479e-07 | -1.7801e-07 | -2.3244e-07 | -2.8099e-07 | -3.6044e-07 | -2.1038e-07 | -2.0596e-07 | -1.9419e-07 | -1.3682e-07 | -2.2215e-07 | -2.0743e-07 | -1.4859e-07 | -1.839e-07 | -2.2509e-07 | -2.604e-07 | -1.7213e-07 | -3.3248e-07 | -1.6624e-07 | -2.9865e-07 | -2.3244e-07
| 2 | 0.1 | 2.0 | 0.0 | 1 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0574480E+00 | DC2_2V |  | 0.099999949851999989 | 7.3224443841624563e-08 | 1517790523.01 | 1517790538.463 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 2 | -0.1 | 2.0 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9907187E+00 | DC2_2V |  | -0.10000034973499998 | 6.5158293986904343e-08 | 1517790553.29 | 1517790568.687 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0349206E+00 | DC2_2V |  | -1.0286031058650001e-07 | 4.8280933166767952e-08 | 1517790583.546 | 1517790599.045 | -1.1184e-07 | -5.1504e-08 | -3.5317e-08 | -6.7691e-08 | -9.7121e-08 | -7.652e-08 | -1.3097e-07 | -1.6481e-07 | -1.9277e-07 | -4.4146e-08 | -8.3878e-08 | -5.4447e-08 | -7.0634e-08 | -2.1632e-07 | -1.6334e-07 | -1.1331e-07 | -1.0006e-07 | -9.8593e-08 | -7.5048e-08 | -1.0889e-07
| 2 | -0.1 | 2.0 | -0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9802154E+00 | DC2_2V |  | 4.0887838719500003e-07 | 7.0000599179079395e-08 | 1517790613.898 | 1517790629.358 | 4.2435e-07 | 2.9321e-07 | 3.4036e-07 | 3.33e-07 | 4.6855e-07 | 3.993e-07 | 2.9027e-07 | 4.435e-07 | 5.2602e-07 | 4.9949e-07 | 4.5971e-07 | 2.8879e-07 | 4.273e-07 | 3.6394e-07 | 4.7739e-07 | 4.0372e-07 | 3.9046e-07 | 4.1551e-07 | 4.3466e-07 | 4.9802e-07
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9922088E+00 | DC2_2V |  | -7.0764216264849998e-08 | 7.5869601949250209e-08 | 1517790644.199 | 1517790659.702 | -3.9722e-08 | -1.4271e-07 | -6.179e-08 | 8.9743e-08 | -8.0915e-08 | 2.7953e-08 | -1.7654e-08 | -1.7507e-07 | -3.2366e-08 | 2.9424e-08 | -7.3559e-09 | -7.9444e-08 | -1.7949e-07 | -1.4712e-07 | -8.5329e-08 | -1.4712e-08 | -1.4565e-07 | -1.074e-07 | -1.9567e-07 | -5.002e-08
| 2 | -0.1 | 2.0 | -0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9979057E+00 | DC2_2V |  | 3.9878983934000002e-07 | 7.1440308194584024e-08 | 1517790674.529 | 1517790690.024 | 4.3697e-07 | 4.0313e-07 | 4.7228e-07 | 3.8253e-07 | 4.7964e-07 | 4.355e-07 | 3.5605e-07 | 3.3545e-07 | 5.0612e-07 | 3.8695e-07 | 2.9867e-07 | 2.6042e-07 | 3.075e-07 | 3.5016e-07 | 3.8842e-07 | 4.1637e-07 | 4.0902e-07 | 4.4727e-07 | 5.5467e-07 | 3.4869e-07
| 2 | -0.1 | 2.0 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9786978E+00 | DC2_2V |  | -0.10000034398999999 | 7.3746504324080595e-08 | 1517790704.863 | 1517790720.291 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 2 | -0.2 | 2.0 | -0.1 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9800834E+00 | DC2_2V |  | -0.099999800197000013 | 6.8469666283722152e-08 | 1517790735.114 | 1517790750.567 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 2 | -0.2 | 2.0 | 0.0 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9730287E+00 | DC2_2V |  | -0.20000047859499998 | 5.9602076095766979e-08 | 1517790765.397 | 1517790780.821 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 2 | -0.3 | 2.0 | -0.1 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9484541E+00 | DC2_2V |  | -0.200000009305 | 5.969487813127813e-08 | 1517790795.656 | 1517790811.093 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 2 | -0.3 | 2.0 | 0.0 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0314632E+00 | DC2_2V |  | -0.30000067432999999 | 4.9626364968520541e-08 | 1517790825.931 | 1517790841.379 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 2 | -0.4 | 2.0 | -0.1 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9814880E+00 | DC2_2V |  | -0.30000018186499999 | 6.8862175939696774e-08 | 1517790856.219 | 1517790871.629 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 2 | -0.4 | 2.0 | 0.0 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0149337E+00 | DC2_2V |  | -0.40000088715499993 | 4.6213542124585238e-08 | 1517790886.48 | 1517790901.909 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 2 | -0.5 | 2.0 | -0.1 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 49.8896122E+00 | DC2_2V |  | -0.40000034570000009 | 7.8660587340397716e-08 | 1517790916.751 | 1517790932.164 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 2 | -0.5 | 2.0 | 0.0 | 1 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9863104E+00 | DC2_2V |  | -0.50000107470999999 | 7.1104851448211245e-08 | 1517790947.008 | 1517790962.472 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 2 | -0.6 | 2.0 | -0.1 | 1 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9952127E+00 | DC2_2V |  | -0.50000056996499997 | 7.8278485375308686e-08 | 1517790977.34 | 1517790992.77 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 2 | -0.6 | 2.0 | 0.0 | 1 | -0.6 | 20.0 | 10.0 | 0.0 | 50.0205556E+00 | DC2_2V |  | -0.60000125569499996 | 7.7491234821364107e-08 | 1517791007.622 | 1517791023.039 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 2 | -0.7 | 2.0 | -0.1 | 1 | -0.6 | 20.0 | 10.0 | 0.0 | 50.0555144E+00 | DC2_2V |  | -0.60000078737000007 | 5.1595882573262485e-08 | 1517791037.875 | 1517791053.316 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 2 | -0.7 | 2.0 | 0.0 | 1 | -0.7 | 20.0 | 10.0 | 0.0 | 49.9652035E+00 | DC2_2V |  | -0.70000145959999993 | 8.4291280685458812e-08 | 1517791068.169 | 1517791083.581 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 2 | -0.8 | 2.0 | -0.1 | 1 | -0.7000000000000001 | 20.0 | 10.0 | 0.0 | 49.9581012E+00 | DC2_2V |  | -0.70000093974000011 | 9.0790370624451641e-08 | 1517791098.421 | 1517791113.858 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 2 | -0.8 | 2.0 | 0.0 | 1 | -0.8 | 20.0 | 10.0 | 0.0 | 50.0271940E+00 | DC2_2V |  | -0.800001674025 | 6.0010356383375305e-08 | 1517791128.719 | 1517791144.163 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 2 | -0.9 | 2.0 | -0.1 | 1 | -0.8 | 20.0 | 10.0 | 0.0 | 50.0215706E+00 | DC2_2V |  | -0.80000120349999992 | 7.117031684467616e-08 | 1517791158.986 | 1517791174.411 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 2 | -0.9 | 2.0 | 0.0 | 1 | -0.9 | 20.0 | 10.0 | 0.0 | 50.0379344E+00 | DC2_2V |  | -0.90000194413000012 | 7.7254016708904522e-08 | 1517791189.237 | 1517791204.664 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 2 | -1.0 | 2.0 | -0.1 | 1 | -0.9 | 20.0 | 10.0 | 0.0 | 50.0041041E+00 | DC2_2V |  | -0.90000138756000025 | 9.8467387493049943e-08 | 1517791219.49 | 1517791234.904 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 2 | -1.0 | 2.0 | 0.0 | 1 | -1.0 | 20.0 | 10.0 | 0.0 | 50.0221277E+00 | DC2_2V |  | -1.0000021298499999 | 7.5485942399409128e-08 | 1517791249.725 | 1517791265.142 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 2 | -1.0 | 2.0 | -0.1 | 1 | -0.9 | 20.0 | 10.0 | 0.0 | 49.9665837E+00 | DC2_2V |  | -0.90000136142999998 | 6.811755354331787e-08 | 1517791279.993 | 1517791295.411 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 2 | -0.9 | 2.0 | 0.0 | 1 | -0.9 | 20.0 | 10.0 | 0.0 | 50.0294655E+00 | DC2_2V |  | -0.90000192780000021 | 8.1791405425181602e-08 | 1517791310.246 | 1517791325.677 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 2 | -0.9 | 2.0 | -0.1 | 1 | -0.8 | 20.0 | 10.0 | 0.0 | 49.9985755E+00 | DC2_2V |  | -0.80000127036499991 | 6.9880149362324377e-08 | 1517791340.497 | 1517791355.908 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 2 | -0.8 | 2.0 | 0.0 | 1 | -0.8 | 20.0 | 10.0 | 0.0 | 50.0544035E+00 | DC2_2V |  | -0.80000168724500009 | 5.8505063662427288e-08 | 1517791370.768 | 1517791386.195 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 2 | -0.8 | 2.0 | -0.1 | 1 | -0.7000000000000001 | 20.0 | 10.0 | 0.0 | 49.9768739E+00 | DC2_2V |  | -0.7000009988 | 6.5157854475874829e-08 | 1517791401.039 | 1517791416.441 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 2 | -0.7 | 2.0 | 0.0 | 1 | -0.7 | 20.0 | 10.0 | 0.0 | 49.9592314E+00 | DC2_2V |  | -0.70000142720999992 | 6.5872474518463952e-08 | 1517791431.286 | 1517791446.731 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 2 | -0.7 | 2.0 | -0.1 | 1 | -0.6 | 20.0 | 10.0 | 0.0 | 49.9744800E+00 | DC2_2V |  | -0.60000078225999998 | 5.2470357338801516e-08 | 1517791461.575 | 1517791476.995 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 2 | -0.6 | 2.0 | 0.0 | 1 | -0.6 | 20.0 | 10.0 | 0.0 | 50.0233977E+00 | DC2_2V |  | -0.60000129836499994 | 6.0091141400396869e-08 | 1517791491.838 | 1517791507.275 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 2 | -0.6 | 2.0 | -0.1 | 1 | -0.5 | 20.0 | 10.0 | 0.0 | 50.0408021E+00 | DC2_2V |  | -0.50000056430999995 | 6.5113461751242031e-08 | 1517791522.125 | 1517791537.533 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 2 | -0.5 | 2.0 | 0.0 | 1 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9846274E+00 | DC2_2V |  | -0.50000104533500012 | 1.024234508097424e-07 | 1517791552.403 | 1517791567.798 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 2 | -0.5 | 2.0 | -0.1 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0022359E+00 | DC2_2V |  | -0.40000033415999992 | 7.2708647356312412e-08 | 1517791582.682 | 1517791598.11 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 2 | -0.4 | 2.0 | 0.0 | 1 | -0.4 | 20.0 | 10.0 | 0.0 | 49.9962031E+00 | DC2_2V |  | -0.40000083829999999 | 4.9632529660736975e-08 | 1517791612.967 | 1517791628.397 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 2 | -0.4 | 2.0 | -0.1 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0078739E+00 | DC2_2V |  | -0.30000018786499999 | 7.2809039781863962e-08 | 1517791643.239 | 1517791658.656 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 2 | -0.3 | 2.0 | 0.0 | 1 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0120359E+00 | DC2_2V |  | -0.30000067595499996 | 7.524538174420712e-08 | 1517791673.505 | 1517791688.925 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 2 | -0.3 | 2.0 | -0.1 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9834394E+00 | DC2_2V |  | -0.19999997833000002 | 6.541306520869194e-08 | 1517791703.778 | 1517791719.208 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 2 | -0.2 | 2.0 | 0.0 | 1 | -0.2 | 20.0 | 10.0 | 0.0 | 50.0184360E+00 | DC2_2V |  | -0.20000047545000005 | 6.7827719262067286e-08 | 1517791734.039 | 1517791749.469 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 2 | -0.2 | 2.0 | -0.1 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 50.0048307E+00 | DC2_2V |  | -0.09999975315399999 | 6.2800897717201012e-08 | 1517791764.308 | 1517791779.753 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 2 | -0.1 | 2.0 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9691982E+00 | DC2_2V |  | -0.10000024288000002 | 6.2099714975421769e-08 | 1517791794.582 | 1517791810.005 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 2 | -0.1 | 2.0 | -0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0320630E+00 | DC2_2V |  | 4.5148505636499998e-07 | 6.6891355964823537e-08 | 1517791824.846 | 1517791840.335 | 4.3265e-07 | 5.283e-07 | 5.5038e-07 | 5.0181e-07 | 3.8261e-07 | 3.7967e-07 | 5.3713e-07 | 4.2088e-07 | 4.5472e-07 | 4.9446e-07 | 4.4736e-07 | 3.9439e-07 | 4.8857e-07 | 4.091e-07 | 5.2977e-07 | 3.4582e-07 | 4.9298e-07 | 3.576e-07 | 5.3125e-07 | 3.5024e-07
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9130571E+00 | DC2_2V |  | -8.04829109463e-07 | 7.6201627578914526e-07 | 1517791855.153 | 1517791870.618 | 4.2724e-08 | -1.9152e-08 | -1.1933e-07 | -2.3572e-08 | -6.6296e-08 | -1.7679e-08 | -1.7679e-08 | -7.3662e-08 | -6.4822e-08 | -1.1477e-06 | -2.1377e-06 | -1.6603e-06 | -2.1244e-06 | -1.9167e-06 | -9.8118e-07 | -1.1418e-06 | -1.1138e-06 | -1.3539e-06 | -9.4876e-07 | -1.211e-06
| 2 | -0.1 | 2.0 | -0.1 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9906433E+00 | DC2_2V |  | 1.0658346545000007e-08 | 4.4177678433995107e-07 | 1517791885.451 | 1517791900.963 | -7.2624e-07 | -3.3225e-07 | -3.5283e-07 | -5.7335e-07 | -4.278e-07 | -5.557e-07 | -1.1761e-07 | -3.8958e-07 | -3.0578e-07 | -3.6606e-07 | 4.2633e-07 | 3.9987e-07 | 3.9105e-07 | 4.5868e-07 | 4.6603e-07 | 5.7041e-07 | 3.984e-07 | 3.7782e-07 | 4.5427e-07 | 4.1751e-07
| 2 | 0.0 | 2.0 | 0.0 | 1 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9977690E+00 | DC2_2V |  | -3.1661590946499991e-08 | 8.0617026741158621e-08 | 1517791915.782 | 1517791931.264 | -6.7741e-08 | -3.5343e-08 | -1.7672e-08 | 3.9761e-08 | -5.7433e-08 | -1.9144e-08 | -9.8666e-08 | -1.3254e-08 | -1.5021e-07 | -2.3562e-08 | 2.9453e-08 | -1.7377e-07 | -1.4432e-07 | -8.6885e-08 | -1.4726e-08 | 1.8261e-07 | -6.3323e-08 | 7.805e-08 | -3.3871e-08 | 3.6816e-08
| 2 | -0.1 | 2.0 | 0.0 | 1 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9526769E+00 | DC2_2V |  | -0.10000029802500002 | 5.828753618862215e-08 | 1517791946.114 | 1517791961.549 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | -0.1
| 20 | 0.05 | 0.2 | 0.0 | 10 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9909167E+00 | DC220MV |  | 0.050001674766000002 | 3.4168190609691153e-07 | 1517791977.1 | 1517791992.566 | 0.050002 | 0.050002 | 0.050001 | 0.050001 | 0.050002 | 0.050002 | 0.050001 | 0.050001 | 0.050002 | 0.050001 | 0.050001 | 0.050002 | 0.050002 | 0.050001 | 0.050002 | 0.050002 | 0.050002 | 0.050002 | 0.050001 | 0.050002
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9834960E+00 | DC220MV |  | 1.1293226218350001e-06 | 3.4195716394826731e-07 | 1517792007.428 | 1517792022.92 | 1.7804e-06 | 1.2066e-06 | 1.2949e-06 | 1.4714e-06 | 1.4567e-06 | 5.2971e-07 | 8.24e-07 | 1.4567e-06 | 1.0153e-06 | 6.6214e-07 | 1.2654e-06 | 1.236e-06 | 5.8857e-07 | 8.6814e-07 | 1.236e-06 | 1.2801e-06 | 1.2949e-06 | 1.133e-06 | 5.5914e-07 | 1.4273e-06
| 20 | 0.05 | 0.2 | 0.05 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0341507E+00 | DC220MV |  | 1.246775282395e-06 | 2.5951325712449154e-07 | 1517792037.748 | 1517792053.247 | 1.2357e-06 | 1.6477e-06 | 1.53e-06 | 1.2946e-06 | 1.53e-06 | 6.62e-07 | 1.1769e-06 | 9.8565e-07 | 1.0592e-06 | 1.1475e-06 | 1.4858e-06 | 9.121e-07 | 9.2681e-07 | 1.1622e-06 | 1.2946e-06 | 1.427e-06 | 1.324e-06 | 1.7359e-06 | 1.1916e-06 | 1.2063e-06
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9894881E+00 | DC220MV |  | 9.8709170392500034e-07 | 3.0275633549964709e-07 | 1517792068.095 | 1517792083.56 | 7.3664e-07 | 7.5137e-07 | 9.2816e-07 | 6.777e-07 | 1.1344e-06 | 8.6923e-07 | 1.3996e-06 | 1.3554e-06 | 1.3112e-06 | 1.3701e-06 | 1.3112e-06 | 1.3112e-06 | 9.7236e-07 | 1.2965e-06 | 7.0717e-07 | 7.3664e-07 | 3.3885e-07 | 7.661e-07 | 6.777e-07 | 1.0902e-06
| 20 | 0.05 | 0.2 | 0.05 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0340987E+00 | DC220MV |  | 1.11979191034e-06 | 2.0547492875317196e-07 | 1517792098.383 | 1517792113.874 | 1.2508e-06 | 1.2213e-06 | 1.236e-06 | 1.0153e-06 | 5.1502e-07 | 9.2703e-07 | 1.5745e-06 | 1.2213e-06 | 8.0931e-07 | 1.2508e-06 | 1.1625e-06 | 1.2655e-06 | 1.2655e-06 | 1.133e-06 | 1.1183e-06 | 1.03e-06 | 1.0742e-06 | 1.1477e-06 | 1.1183e-06 | 1.0595e-06
| 20 | 0.05 | 0.2 | 0.0 | 10 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0118000E+00 | DC220MV |  | 0.050001187340999995 | 2.8233400802837618e-07 | 1517792128.73 | 1517792144.148 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050002 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050002 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001
| 20 | 0.1 | 0.2 | 0.05 | 10 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0201166E+00 | DC220MV |  | 0.0500009675455 | 3.0592020403195339e-07 | 1517792158.973 | 1517792174.4 | 0.050002 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001
| 20 | 0.1 | 0.2 | 0.0 | 10 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0139898E+00 | DC220MV |  | 0.10000096160000001 | 2.8734773010973324e-07 | 1517792189.236 | 1517792204.684 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.15 | 0.2 | 0.05 | 10 | 0.09999999999999999 | 20.0 | 10.0 | 0.0 | 49.9599849E+00 | DC220MV |  | 0.10000120202 | 2.9602061009407102e-07 | 1517792219.538 | 1517792234.969 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.15 | 0.2 | 0.0 | 10 | 0.15 | 20.0 | 10.0 | 0.0 | 49.9870034E+00 | DC220MV |  | 0.15000140165499998 | 2.4815796677755105e-07 | 1517792249.824 | 1517792265.293 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
| 20 | 0.2 | 0.2 | 0.05 | 10 | 0.15 | 20.0 | 10.0 | 0.0 | 50.0846620E+00 | DC220MV |  | 0.15000140691000002 | 3.6186104087146662e-07 | 1517792280.128 | 1517792295.573 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
| 20 | 0.2 | 0.2 | 0.0 | 10 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9376298E+00 | DC220MV |  | 0.20000179689 | 2.8705557632548757e-07 | 1517792310.406 | 1517792325.809 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 20 | 0.25 | 0.2 | 0.05 | 10 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0046655E+00 | DC220MV |  | 0.20000143432 | 2.500985597730585e-07 | 1517792340.638 | 1517792356.116 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 20 | 0.25 | 0.2 | 0.0 | 10 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0091432E+00 | DC220MV |  | 0.25000157984500004 | 2.9233150270879348e-07 | 1517792370.947 | 1517792386.381 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
| 20 | 0.3 | 0.2 | 0.05 | 10 | 0.25 | 20.0 | 10.0 | 0.0 | 49.9377239E+00 | DC220MV |  | 0.25000155600000001 | 2.3339884318709346e-07 | 1517792401.204 | 1517792416.622 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
| 20 | 0.3 | 0.2 | 0.0 | 10 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0345380E+00 | DC220MV |  | 0.30000184250499995 | 3.5394588354082367e-07 | 1517792431.453 | 1517792446.914 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 20 | 0.35 | 0.2 | 0.05 | 10 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9973681E+00 | DC220MV |  | 0.30000164364500004 | 2.7328627385459956e-07 | 1517792461.758 | 1517792477.185 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 20 | 0.35 | 0.2 | 0.0 | 10 | 0.35 | 20.0 | 10.0 | 0.0 | 50.0403864E+00 | DC220MV |  | 0.35000168747999999 | 2.9891598752452466e-07 | 1517792492.016 | 1517792507.449 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35
| 20 | 0.4 | 0.2 | 0.05 | 10 | 0.35 | 20.0 | 10.0 | 0.0 | 49.9651941E+00 | DC220MV |  | 0.35000200849999996 | 3.1872173286437856e-07 | 1517792522.299 | 1517792537.705 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35
| 20 | 0.4 | 0.2 | 0.0 | 10 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0371266E+00 | DC220MV |  | 0.40000206227500001 | 2.9085220624444259e-07 | 1517792552.544 | 1517792567.99 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 20 | 0.45 | 0.2 | 0.05 | 10 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9621843E+00 | DC220MV |  | 0.40000172444500004 | 3.0963226491881021e-07 | 1517792582.813 | 1517792598.231 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 20 | 0.45 | 0.2 | 0.0 | 10 | 0.45 | 20.0 | 10.0 | 0.0 | 50.0108892E+00 | DC220MV |  | 0.45000208276500003 | 3.9420164291052749e-07 | 1517792613.079 | 1517792628.549 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45
| 20 | 0.5 | 0.2 | 0.05 | 10 | 0.45 | 20.0 | 10.0 | 0.0 | 50.0058167E+00 | DC220MV |  | 0.45000205449999992 | 2.8027177523714661e-07 | 1517792643.379 | 1517792658.816 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45
| 20 | 0.5 | 0.2 | 0.0 | 10 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0300936E+00 | DC220MV |  | 0.50000223356999995 | 2.4478026492819311e-07 | 1517792673.674 | 1517792689.109 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 20 | 0.5 | 0.2 | 0.05 | 10 | 0.45 | 20.0 | 10.0 | 0.0 | 49.9600979E+00 | DC220MV |  | 0.45000207602000009 | 2.9964026365037016e-07 | 1517792703.938 | 1517792719.375 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45
| 20 | 0.45 | 0.2 | 0.0 | 10 | 0.45 | 20.0 | 10.0 | 0.0 | 49.9910912E+00 | DC220MV |  | 0.45000195278999999 | 3.1833888059271738e-07 | 1517792734.21 | 1517792749.674 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45 | 0.45
| 20 | 0.45 | 0.2 | 0.05 | 10 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0047740E+00 | DC220MV |  | 0.40000169090000004 | 2.2902380880570046e-07 | 1517792764.522 | 1517792779.965 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 20 | 0.4 | 0.2 | 0.0 | 10 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0650799E+00 | DC220MV |  | 0.40000205967000008 | 2.771338036067212e-07 | 1517792794.798 | 1517792810.235 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 20 | 0.4 | 0.2 | 0.05 | 10 | 0.35 | 20.0 | 10.0 | 0.0 | 50.0001415E+00 | DC220MV |  | 0.35000185377500004 | 3.4506906971513098e-07 | 1517792825.065 | 1517792840.472 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35
| 20 | 0.35 | 0.2 | 0.0 | 10 | 0.35 | 20.0 | 10.0 | 0.0 | 49.9822139E+00 | DC220MV |  | 0.35000167252500003 | 3.0905679878614157e-07 | 1517792855.325 | 1517792870.771 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35 | 0.35
| 20 | 0.35 | 0.2 | 0.05 | 10 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9985236E+00 | DC220MV |  | 0.30000144806500001 | 2.3622480029691268e-07 | 1517792885.618 | 1517792901.051 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 20 | 0.3 | 0.2 | 0.0 | 10 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0234355E+00 | DC220MV |  | 0.30000179134499999 | 2.8945516314357291e-07 | 1517792915.908 | 1517792931.351 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 20 | 0.3 | 0.2 | 0.05 | 10 | 0.25 | 20.0 | 10.0 | 0.0 | 49.9619205E+00 | DC220MV |  | 0.25000142784500001 | 2.6791147694228347e-07 | 1517792946.177 | 1517792961.601 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
| 20 | 0.25 | 0.2 | 0.0 | 10 | 0.25 | 20.0 | 10.0 | 0.0 | 50.0037880E+00 | DC220MV |  | 0.25000148019499996 | 2.839469677124048e-07 | 1517792976.438 | 1517792991.894 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
| 20 | 0.25 | 0.2 | 0.05 | 10 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9933216E+00 | DC220MV |  | 0.20000121794 | 3.1439408932296668e-07 | 1517793006.743 | 1517793022.166 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 20 | 0.2 | 0.2 | 0.0 | 10 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0256594E+00 | DC220MV |  | 0.20000159185499999 | 2.4217646763328239e-07 | 1517793037.002 | 1517793052.427 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 20 | 0.2 | 0.2 | 0.05 | 10 | 0.15 | 20.0 | 10.0 | 0.0 | 50.0337350E+00 | DC220MV |  | 0.15000128726500001 | 3.5678275361311069e-07 | 1517793067.277 | 1517793082.726 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
| 20 | 0.15 | 0.2 | 0.0 | 10 | 0.15 | 20.0 | 10.0 | 0.0 | 49.9760775E+00 | DC220MV |  | 0.15000122496500001 | 2.4446694515639489e-07 | 1517793097.564 | 1517793112.987 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
| 20 | 0.15 | 0.2 | 0.05 | 10 | 0.09999999999999999 | 20.0 | 10.0 | 0.0 | 50.0032502E+00 | DC220MV |  | 0.10000098183500002 | 3.6882025334177547e-07 | 1517793127.821 | 1517793143.275 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.1 | 0.2 | 0.0 | 10 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9611528E+00 | DC220MV |  | 0.10000099852 | 3.1882364655088575e-07 | 1517793158.124 | 1517793173.545 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.1 | 0.2 | 0.05 | 10 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0527868E+00 | DC220MV |  | 0.050000773592500006 | 2.7486774057489929e-07 | 1517793188.4 | 1517793203.876 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.05 | 0.050001 | 0.050001 | 0.05 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.05 | 0.05 | 0.05 | 0.050001
| 20 | 0.05 | 0.2 | 0.0 | 10 | 0.05 | 20.0 | 10.0 | 0.0 | 49.9940337E+00 | DC220MV |  | 0.050000831728000003 | 2.3841705588328776e-07 | 1517793218.726 | 1517793234.143 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001
| 20 | 0.05 | 0.2 | 0.05 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0241721E+00 | DC220MV |  | 7.3848442807999999e-07 | 2.9577945935248367e-07 | 1517793249.001 | 1517793264.505 | 4.119e-07 | 2.7951e-07 | 1.2946e-06 | 8.2381e-07 | 8.9736e-07 | 1.0445e-06 | 5.2959e-07 | 1.0886e-06 | 7.6496e-07 | 6.9141e-07 | 1.0298e-07 | 8.5323e-07 | 4.2661e-07 | 7.5025e-07 | 8.9736e-07 | 4.2661e-07 | 1.118e-06 | 6.4728e-07 | 9.2678e-07 | 7.9439e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0422479E+00 | DC220MV |  | 4.4166968045500001e-07 | 2.923409950418241e-07 | 1517793279.342 | 1517793294.821 | 4.5639e-07 | 2.7972e-07 | 8.8334e-07 | 7.8028e-07 | 5.7417e-07 | 2.9445e-07 | 6.1834e-07 | 2.7972e-07 | 2.65e-07 | 0 | 7.6556e-07 | -1.7667e-07 | 2.2083e-07 | 8.5389e-07 | 6.0362e-07 | 1.0306e-07 | 5.3e-07 | 2.5028e-07 | 8.3917e-07 | 4.1223e-07
| 20 | 0.05 | 0.2 | 0.05 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9867158E+00 | DC220MV |  | 7.0050539620499996e-07 | 2.2638760240549303e-07 | 1517793309.675 | 1517793325.145 | 1.1196e-06 | 7.2187e-07 | 4.2723e-07 | 4.4196e-07 | 9.4285e-07 | 8.3972e-07 | 4.2723e-07 | 8.2499e-07 | 6.6294e-07 | 5.3035e-07 | 8.8392e-07 | 1.0902e-06 | 5.5982e-07 | 4.2723e-07 | 6.1874e-07 | 6.4821e-07 | 8.5445e-07 | 4.4196e-07 | 1.0165e-06 | 5.3035e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0601052E+00 | DC220MV |  | 5.2469073050499995e-07 | 1.9807791411874514e-07 | 1517793339.987 | 1517793355.481 | 2.9436e-07 | 8.3892e-07 | 4.7097e-07 | 5.74e-07 | 8.0948e-07 | 9.5666e-07 | 6.4758e-07 | 4.5625e-07 | 4.2682e-07 | 5.4456e-07 | 2.502e-07 | 2.502e-07 | 6.4758e-07 | 2.6492e-07 | 3.9738e-07 | 5.0041e-07 | 6.0343e-07 | 7.3589e-07 | 3.6795e-07 | 4.5625e-07
| 20 | 0.05 | 0.2 | 0.0 | 10 | 0.05 | 20.0 | 10.0 | 0.0 | 50.0109788E+00 | DC220MV |  | 0.050001039313500008 | 3.9062200340427019e-07 | 1517793370.319 | 1517793385.733 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.05 | 0.050002 | 0.050001 | 0.05 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001 | 0.050001
| 20 | -0.05 | 0.2 | 0.0 | 10 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9668192E+00 | DC220MV |  | -0.049999104129999999 | 2.5953448616785579e-07 | 1517793400.565 | 1517793416.018 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.05 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.05 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9385050E+00 | DC220MV |  | 9.026682374899998e-07 | 2.2247326460750226e-07 | 1517793430.845 | 1517793446.343 | 7.2096e-07 | 7.651e-07 | 6.621e-07 | 5.2968e-07 | 7.651e-07 | 8.9752e-07 | 1.0447e-06 | 1.2065e-06 | 1.1329e-06 | 1.2654e-06 | 7.3567e-07 | 8.9752e-07 | 9.4166e-07 | 1.1035e-06 | 8.0924e-07 | 9.4166e-07 | 8.5338e-07 | 1.3242e-06 | 5.0026e-07 | 9.5637e-07
| 20 | -0.05 | 0.2 | -0.05 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0023209E+00 | DC220MV |  | 1.1834209240100001e-06 | 3.4784030964472244e-07 | 1517793461.2 | 1517793476.733 | 8.5212e-07 | 9.4027e-07 | 5.8767e-07 | 1.6308e-06 | 1.1166e-06 | 1.3223e-06 | 1.3076e-06 | 1.2194e-06 | 1.4986e-06 | 1.3516e-06 | 9.9904e-07 | 9.5496e-07 | 1.2488e-06 | 8.8151e-07 | 1.4692e-06 | 9.8435e-07 | 1.3369e-06 | 2.1303e-06 | 1.2047e-06 | 6.3175e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9927086E+00 | DC220MV |  | 6.6990253642500008e-07 | 3.174363400352384e-07 | 1517793491.559 | 1517793507.043 | 1.7668e-07 | 1.0012e-06 | 6.1837e-07 | 9.1283e-07 | 8.8339e-07 | 7.5088e-07 | 9.2756e-07 | 1.5754e-06 | 5.1531e-07 | 6.0365e-07 | 8.6866e-07 | 5.5948e-07 | 3.2391e-07 | 6.0365e-07 | 1.0306e-07 | 4.7114e-07 | 6.3309e-07 | 7.8033e-07 | 7.0671e-07 | 3.828e-07
| 20 | -0.05 | 0.2 | -0.05 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9759974E+00 | DC220MV |  | 1.1288351809649999e-06 | 2.2848237445496001e-07 | 1517793521.89 | 1517793537.378 | 1.1333e-06 | 1.0449e-06 | 1.0449e-06 | 1.2363e-06 | 1.4718e-06 | 1.3099e-06 | 1.3834e-06 | 7.6531e-07 | 8.9777e-07 | 1.6484e-06 | 6.9172e-07 | 1.0302e-06 | 8.9777e-07 | 1.251e-06 | 1.0008e-06 | 1.1038e-06 | 1.2363e-06 | 1.1774e-06 | 1.2804e-06 | 9.7136e-07
| 20 | -0.05 | 0.2 | 0.0 | 10 | -0.05 | 20.0 | 10.0 | 0.0 | 50.0408116E+00 | DC220MV |  | -0.049999222189000006 | 2.0264072507570969e-07 | 1517793552.205 | 1517793567.687 | -0.049999 | -0.049999 | -0.05 | -0.049999 | -0.05 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999
| 20 | -0.1 | 0.2 | -0.05 | 10 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9270312E+00 | DC220MV |  | -0.049998876569999989 | 3.3681688563715475e-07 | 1517793582.529 | 1517793597.959 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049998 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.05 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049998 | -0.05
| 20 | -0.1 | 0.2 | 0.0 | 10 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9729345E+00 | DC220MV |  | -0.099999059953500002 | 2.7654602004502556e-07 | 1517793612.789 | 1517793628.285 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099998 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999
| 20 | -0.15 | 0.2 | -0.05 | 10 | -0.09999999999999999 | 20.0 | 10.0 | 0.0 | 50.0435095E+00 | DC220MV |  | -0.099999139442999999 | 3.3979083301587197e-07 | 1517793643.117 | 1517793658.587 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099998 | -0.099999 | -0.099999 | -0.099999
| 20 | -0.15 | 0.2 | 0.0 | 10 | -0.15 | 20.0 | 10.0 | 0.0 | 49.9759738E+00 | DC220MV |  | -0.149999381855 | 3.5710005807009683e-07 | 1517793673.424 | 1517793688.849 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15
| 20 | -0.2 | 0.2 | -0.05 | 10 | -0.15 | 20.0 | 10.0 | 0.0 | 49.9808894E+00 | DC220MV |  | -0.14999945305999998 | 2.4723469295430729e-07 | 1517793703.67 | 1517793719.119 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15
| 20 | -0.2 | 0.2 | 0.0 | 10 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9558974E+00 | DC220MV |  | -0.19999973270499999 | 2.423629663895564e-07 | 1517793733.961 | 1517793749.412 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 20 | -0.25 | 0.2 | -0.05 | 10 | -0.2 | 20.0 | 10.0 | 0.0 | 50.0676291E+00 | DC220MV |  | -0.19999962817000003 | 3.4957623503282135e-07 | 1517793764.235 | 1517793779.7 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 20 | -0.25 | 0.2 | 0.0 | 10 | -0.25 | 20.0 | 10.0 | 0.0 | 49.9769305E+00 | DC220MV |  | -0.24999969988499995 | 3.3805395467837067e-07 | 1517793794.536 | 1517793809.952 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25
| 20 | -0.3 | 0.2 | -0.05 | 10 | -0.25 | 20.0 | 10.0 | 0.0 | 50.0468552E+00 | DC220MV |  | -0.249999931905 | 2.7451359251592887e-07 | 1517793824.807 | 1517793840.262 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25
| 20 | -0.3 | 0.2 | 0.0 | 10 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9726282E+00 | DC220MV |  | -0.30000010021000001 | 3.268677513988475e-07 | 1517793855.093 | 1517793870.505 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 20 | -0.35 | 0.2 | -0.05 | 10 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0288138E+00 | DC220MV |  | -0.29999982521500002 | 3.880846315912276e-07 | 1517793885.359 | 1517793900.792 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 20 | -0.35 | 0.2 | 0.0 | 10 | -0.35 | 20.0 | 10.0 | 0.0 | 50.0152877E+00 | DC220MV |  | -0.34999992392000001 | 3.2030868642539698e-07 | 1517793915.669 | 1517793931.099 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35
| 20 | -0.4 | 0.2 | -0.05 | 10 | -0.35 | 20.0 | 10.0 | 0.0 | 49.9611528E+00 | DC220MV |  | -0.35000020075999999 | 3.4792078321926024e-07 | 1517793945.924 | 1517793961.344 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35
| 20 | -0.4 | 0.2 | 0.0 | 10 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0086713E+00 | DC220MV |  | -0.40000029845000001 | 3.3587965627712852e-07 | 1517793976.202 | 1517793991.651 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 20 | -0.45 | 0.2 | -0.05 | 10 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0002642E+00 | DC220MV |  | -0.40000006310499997 | 2.3569841000242691e-07 | 1517794006.501 | 1517794021.931 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 20 | -0.45 | 0.2 | 0.0 | 10 | -0.45 | 20.0 | 10.0 | 0.0 | 49.9846179E+00 | DC220MV |  | -0.45000034417000007 | 3.4059951277273698e-07 | 1517794036.779 | 1517794052.209 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45
| 20 | -0.5 | 0.2 | -0.05 | 10 | -0.45 | 20.0 | 10.0 | 0.0 | 50.0467181E+00 | DC220MV |  | -0.45000045450500015 | 2.9975959279944847e-07 | 1517794067.042 | 1517794082.487 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45
| 20 | -0.5 | 0.2 | 0.0 | 10 | -0.5 | 20.0 | 10.0 | 0.0 | 50.0303250E+00 | DC220MV |  | -0.50000060494500009 | 2.7522865308277375e-07 | 1517794097.322 | 1517794112.736 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 20 | -0.5 | 0.2 | -0.05 | 10 | -0.45 | 20.0 | 10.0 | 0.0 | 49.9529500E+00 | DC220MV |  | -0.45000052301000004 | 3.0064249184414928e-07 | 1517794127.556 | 1517794142.985 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45
| 20 | -0.45 | 0.2 | 0.0 | 10 | -0.45 | 20.0 | 10.0 | 0.0 | 50.0130270E+00 | DC220MV |  | -0.45000035265999994 | 3.311973073603802e-07 | 1517794157.832 | 1517794173.278 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45 | -0.45
| 20 | -0.45 | 0.2 | -0.05 | 10 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0166659E+00 | DC220MV |  | -0.40000029556499994 | 3.248267742627209e-07 | 1517794188.106 | 1517794203.532 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 20 | -0.4 | 0.2 | 0.0 | 10 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0433441E+00 | DC220MV |  | -0.40000047287999996 | 2.8583626711541537e-07 | 1517794218.38 | 1517794233.81 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 20 | -0.4 | 0.2 | -0.05 | 10 | -0.35 | 20.0 | 10.0 | 0.0 | 49.9980425E+00 | DC220MV |  | -0.35000003531000001 | 2.6302234677094606e-07 | 1517794248.651 | 1517794264.096 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35
| 20 | -0.35 | 0.2 | 0.0 | 10 | -0.35 | 20.0 | 10.0 | 0.0 | 50.0571360E+00 | DC220MV |  | -0.35000006515499998 | 3.7517093367297402e-07 | 1517794278.927 | 1517794294.393 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35 | -0.35
| 20 | -0.35 | 0.2 | -0.05 | 10 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9468067E+00 | DC220MV |  | -0.29999978492000001 | 3.1056845397006523e-07 | 1517794309.23 | 1517794324.629 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 20 | -0.3 | 0.2 | 0.0 | 10 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0476397E+00 | DC220MV |  | -0.30000009421500001 | 2.3136769280791027e-07 | 1517794339.467 | 1517794354.91 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 20 | -0.3 | 0.2 | -0.05 | 10 | -0.25 | 20.0 | 10.0 | 0.0 | 49.9112333E+00 | DC220MV |  | -0.24999993421499997 | 2.5300970787013769e-07 | 1517794369.758 | 1517794385.178 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25
| 20 | -0.25 | 0.2 | 0.0 | 10 | -0.25 | 20.0 | 10.0 | 0.0 | 50.0299802E+00 | DC220MV |  | -0.24999986574499999 | 2.6654752948654737e-07 | 1517794400.022 | 1517794415.507 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25 | -0.25
| 20 | -0.25 | 0.2 | -0.05 | 10 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9827512E+00 | DC220MV |  | -0.19999963810499999 | 2.9849286000945698e-07 | 1517794430.327 | 1517794445.758 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 20 | -0.2 | 0.2 | 0.0 | 10 | -0.2 | 20.0 | 10.0 | 0.0 | 50.0403958E+00 | DC220MV |  | -0.19999970677999998 | 3.082085926752633e-07 | 1517794460.614 | 1517794476.069 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 20 | -0.2 | 0.2 | -0.05 | 10 | -0.15 | 20.0 | 10.0 | 0.0 | 49.9596788E+00 | DC220MV |  | -0.14999967864000002 | 2.603614936957193e-07 | 1517794490.92 | 1517794506.342 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15
| 20 | -0.15 | 0.2 | 0.0 | 10 | -0.15 | 20.0 | 10.0 | 0.0 | 49.9982595E+00 | DC220MV |  | -0.14999938324000001 | 2.9513603371663047e-07 | 1517794521.18 | 1517794536.654 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15 | -0.15
| 20 | -0.15 | 0.2 | -0.05 | 10 | -0.09999999999999999 | 20.0 | 10.0 | 0.0 | 49.8998378E+00 | DC220MV |  | -0.099999227467500007 | 2.8948317351588048e-07 | 1517794551.476 | 1517794566.932 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999
| 20 | -0.1 | 0.2 | 0.0 | 10 | -0.1 | 20.0 | 10.0 | 0.0 | 50.1023506E+00 | DC220MV |  | -0.099999071562499986 | 2.737403149865103e-07 | 1517794581.764 | 1517794597.258 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999
| 20 | -0.1 | 0.2 | -0.05 | 10 | -0.05 | 20.0 | 10.0 | 0.0 | 49.9551252E+00 | DC220MV |  | -0.0499988250955 | 2.8382199110449457e-07 | 1517794612.126 | 1517794627.55 | -0.049999 | -0.049999 | -0.049999 | -0.049998 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049998 | -0.049998 | -0.049999
| 20 | -0.05 | 0.2 | 0.0 | 10 | -0.05 | 20.0 | 10.0 | 0.0 | 50.0199089E+00 | DC220MV |  | -0.049999075773000008 | 3.6411733887726433e-07 | 1517794642.378 | 1517794657.834 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049998 | -0.05 | -0.049999 | -0.049999 | -0.05 | -0.049999 | -0.049999 | -0.049999 | -0.05 | -0.049999 | -0.049999
| 20 | -0.05 | 0.2 | -0.05 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9426884E+00 | DC220MV |  | 9.9960571528499991e-07 | 3.2766690127292847e-07 | 1517794672.681 | 1517794688.162 | 1.0305e-06 | 1.2513e-06 | 1.2955e-06 | 6.3303e-07 | 1.3986e-06 | 1.0011e-06 | 1.5605e-06 | 7.9497e-07 | 6.1831e-07 | 1.4427e-06 | 8.097e-07 | 7.6553e-07 | 1.0158e-06 | 1.2808e-06 | 6.1831e-07 | 1.1189e-06 | 4.4165e-07 | 1.0452e-06 | 1.3544e-06 | 5.1526e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0115310E+00 | DC220MV |  | 9.1909252763499986e-07 | 3.2556950658004962e-07 | 1517794702.998 | 1517794718.512 | 9.9997e-07 | 1.5882e-06 | 1.2647e-06 | 1.4411e-06 | 1.1617e-06 | 7.941e-07 | 8.3821e-07 | 9.2645e-07 | 9.4115e-07 | 5.1469e-07 | 5.7351e-07 | 5.294e-07 | 5.441e-07 | 1.0588e-06 | 7.7939e-07 | 1.1764e-06 | 2.794e-07 | 1.0294e-06 | 1.1764e-06 | 7.6468e-07
| 20 | -0.05 | 0.2 | -0.05 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0187098E+00 | DC220MV |  | 1.061700520315e-06 | 2.9669529510234973e-07 | 1517794733.346 | 1517794748.839 | 1.81e-06 | 1.0448e-06 | 6.1804e-07 | 8.8291e-07 | 9.712e-07 | 8.3877e-07 | 1.2361e-06 | 1.1036e-06 | 9.4177e-07 | 1.0006e-06 | 9.8592e-07 | 9.4177e-07 | 1.604e-06 | 6.4747e-07 | 1.0448e-06 | 1.4568e-06 | 1.2802e-06 | 1.2214e-06 | 8.8291e-07 | 7.2104e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9378839E+00 | DC220MV |  | 8.4810758635250014e-07 | 3.4497760562817997e-07 | 1517794763.685 | 1517794779.164 | 1.4577e-06 | 1.2221e-06 | 1.3841e-06 | 1.1927e-06 | 8.0982e-07 | 1.0307e-06 | 5.8896e-07 | 4.27e-07 | 1.2957e-06 | 5.0062e-07 | 7.362e-07 | 6.7731e-07 | 5.4479e-07 | 6.4786e-07 | 1.1043e-06 | 8.54e-07 | 9.5707e-07 | 7.5093e-07 | 8.8345e-08 | 6.9203e-07
| 20 | -0.05 | 0.2 | 0.0 | 10 | -0.05 | 20.0 | 10.0 | 0.0 | 50.0306556E+00 | DC220MV |  | -0.049999095601000007 | 3.1754997930915174e-07 | 1517794794.004 | 1517794809.482 | -0.049999 | -0.049999 | -0.049998 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.05 | -0.05 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999 | -0.049999
| 20 | 0.1 | 0.2 | 0.0 | 10 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9575691E+00 | DC220MV |  | 0.10000080868000001 | 3.4234589467441396e-07 | 1517794824.323 | 1517794839.744 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0174825E+00 | DC220MV |  | 2.146650928115e-06 | 2.381356248291253e-06 | 1517794854.586 | 1517794870.093 | 6.3245e-07 | 5.1478e-07 | 5.7362e-07 | 7.9424e-07 | 6.3245e-07 | 2.6475e-07 | 2.9416e-07 | 8.8249e-07 | 4.2654e-07 | 7.3541e-07 | 7.2217e-06 | 5.0302e-06 | 5.9421e-06 | 6.1627e-06 | 5.0155e-06 | 4.7948e-06 | 1.4855e-06 | 6.1774e-07 | 5.0008e-07 | 4.1183e-07
| 20 | 0.1 | 0.2 | 0.1 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9380674E+00 | DC220MV |  | 5.0080693631999994e-07 | 2.7974638908709339e-07 | 1517794884.929 | 1517794900.402 | 5.3027e-07 | 8.3959e-07 | 2.9459e-07 | 2.7986e-07 | 2.9459e-07 | 8.9851e-07 | 1.2373e-06 | 4.7135e-07 | 2.9459e-07 | 2.2094e-07 | 1.0311e-07 | 7.954e-07 | 7.8067e-07 | 6.6283e-07 | 4.5662e-07 | 2.504e-07 | 3.2405e-07 | 5.5973e-07 | 4.1243e-07 | 3.0932e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0067320E+00 | DC220MV |  | 4.5483015144300011e-07 | 2.4867681227171039e-07 | 1517794915.23 | 1517794930.762 | 5.2904e-07 | 2.4983e-07 | 3.8209e-07 | -7.3478e-08 | 5.4374e-07 | 3.6739e-07 | 1.6165e-07 | 6.1722e-07 | 7.3478e-07 | 9.8461e-07 | 1.6165e-07 | 6.1722e-07 | 3.8209e-07 | 5.1435e-07 | 6.3191e-07 | 4.4087e-08 | 4.5556e-07 | 7.2009e-07 | 6.1722e-07 | 4.5556e-07
| 20 | 0.1 | 0.2 | 0.1 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0514302E+00 | DC220MV |  | 4.4978270450000001e-07 | 3.3132846823446178e-07 | 1517794945.591 | 1517794961.072 | 2.9446e-07 | 2.9446e-07 | 5.4475e-07 | 4.4169e-08 | 2.2084e-07 | 6.7725e-07 | 4.1224e-07 | 2.5029e-07 | 8.392e-07 | 3.239e-07 | 1.0748e-06 | 4.5641e-07 | 8.5392e-07 | 1.7667e-07 | 1.0306e-07 | 1.1778e-06 | 1.3251e-07 | 5.4475e-07 | -4.4169e-08 | 6.1836e-07
| 20 | 0.1 | 0.2 | 0.0 | 10 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0150281E+00 | DC220MV |  | 0.10000080893 | 3.2572717893961569e-07 | 1517794975.913 | 1517794991.328 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.2 | 0.2 | 0.1 | 10 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9873145E+00 | DC220MV |  | 0.10000093185499999 | 3.0457563834758526e-07 | 1517795006.183 | 1517795021.594 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.2 | 0.2 | 0.0 | 10 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9893608E+00 | DC220MV |  | 0.20000133292999997 | 2.9508999661424756e-07 | 1517795036.435 | 1517795051.879 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 20 | 0.3 | 0.2 | 0.1 | 10 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0042456E+00 | DC220MV |  | 0.20000123881000001 | 2.6571014075515538e-07 | 1517795066.701 | 1517795082.132 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 20 | 0.3 | 0.2 | 0.0 | 10 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9660843E+00 | DC220MV |  | 0.30000154334499995 | 3.7679644037010532e-07 | 1517795096.968 | 1517795112.416 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 20 | 0.4 | 0.2 | 0.1 | 10 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0421062E+00 | DC220MV |  | 0.30000138961 | 2.4663864640932211e-07 | 1517795127.28 | 1517795142.736 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 20 | 0.4 | 0.2 | 0.0 | 10 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9673892E+00 | DC220MV |  | 0.40000163098500002 | 2.4128414427149555e-07 | 1517795157.558 | 1517795172.973 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 20 | 0.5 | 0.2 | 0.1 | 10 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0189081E+00 | DC220MV |  | 0.40000156679999999 | 2.8607847349398914e-07 | 1517795187.822 | 1517795203.299 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 20 | 0.5 | 0.2 | 0.0 | 10 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0120359E+00 | DC220MV |  | 0.50000197202999996 | 3.3789598710042145e-07 | 1517795218.121 | 1517795233.543 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 20 | 0.6 | 0.2 | 0.1 | 10 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0060243E+00 | DC220MV |  | 0.50000186540500002 | 3.399032192875032e-07 | 1517795248.417 | 1517795263.857 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 20 | 0.6 | 0.2 | 0.0 | 10 | 0.6 | 20.0 | 10.0 | 0.0 | 49.9969436E+00 | DC220MV |  | 0.60000224507499988 | 2.0931270595362493e-07 | 1517795278.715 | 1517795294.152 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
| 20 | 0.7 | 0.2 | 0.1 | 10 | 0.6 | 20.0 | 10.0 | 0.0 | 49.9552947E+00 | DC220MV |  | 0.60000213860499996 | 3.1967394087280467e-07 | 1517795308.992 | 1517795324.428 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
| 20 | 0.7 | 0.2 | 0.0 | 10 | 0.7 | 20.0 | 10.0 | 0.0 | 50.0381423E+00 | DC220MV |  | 0.70000250190500002 | 3.7994735750331159e-07 | 1517795339.254 | 1517795354.711 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7
| 20 | 0.8 | 0.2 | 0.1 | 10 | 0.7000000000000001 | 20.0 | 10.0 | 0.0 | 49.9704420E+00 | DC220MV |  | 0.70000237362999995 | 2.7228484735283084e-07 | 1517795369.56 | 1517795384.971 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7
| 20 | 0.8 | 0.2 | 0.0 | 10 | 0.8 | 20.0 | 10.0 | 0.0 | 49.9888421E+00 | DC220MV |  | 0.800002736745 | 2.6666980982625781e-07 | 1517795399.812 | 1517795415.273 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| 20 | 0.9 | 0.2 | 0.1 | 10 | 0.8 | 20.0 | 10.0 | 0.0 | 49.9823600E+00 | DC220MV |  | 0.80000269198500007 | 2.5411430159981577e-07 | 1517795430.104 | 1517795445.534 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| 20 | 0.9 | 0.2 | 0.0 | 10 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0196067E+00 | DC220MV |  | 0.90000290499499991 | 2.7308445848676509e-07 | 1517795460.389 | 1517795475.827 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9
| 20 | 1.0 | 0.2 | 0.1 | 10 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0389596E+00 | DC220MV |  | 0.90000294912500001 | 2.7598183975115307e-07 | 1517795490.651 | 1517795506.069 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9
| 20 | 1.0 | 0.2 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9906150E+00 | DC220MV |  | 1.0000032598500002 | 3.2623262790002542e-07 | 1517795520.901 | 1517795536.32 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 1.0 | 0.2 | 0.1 | 10 | 0.9 | 20.0 | 10.0 | 0.0 | 50.0242618E+00 | DC220MV |  | 0.90000298970500003 | 2.8181608447348528e-07 | 1517795551.161 | 1517795566.585 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9
| 20 | 0.9 | 0.2 | 0.0 | 10 | 0.9 | 20.0 | 10.0 | 0.0 | 49.9648596E+00 | DC220MV |  | 0.90000311241499986 | 2.6951632470294947e-07 | 1517795581.41 | 1517795596.831 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9
| 20 | 0.9 | 0.2 | 0.1 | 10 | 0.8 | 20.0 | 10.0 | 0.0 | 50.0223921E+00 | DC220MV |  | 0.80000265305999996 | 3.2699898532788966e-07 | 1517795611.657 | 1517795627.111 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| 20 | 0.8 | 0.2 | 0.0 | 10 | 0.8 | 20.0 | 10.0 | 0.0 | 49.9832509E+00 | DC220MV |  | 0.8000028318650001 | 3.2621104254219879e-07 | 1517795641.932 | 1517795657.357 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| 20 | 0.8 | 0.2 | 0.1 | 10 | 0.7000000000000001 | 20.0 | 10.0 | 0.0 | 50.0562188E+00 | DC220MV |  | 0.70000237156499989 | 2.7998754128178847e-07 | 1517795672.213 | 1517795687.655 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7
| 20 | 0.7 | 0.2 | 0.0 | 10 | 0.7 | 20.0 | 10.0 | 0.0 | 50.0017218E+00 | DC220MV |  | 0.70000243937500006 | 2.526781804513086e-07 | 1517795702.497 | 1517795717.918 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7
| 20 | 0.7 | 0.2 | 0.1 | 10 | 0.6 | 20.0 | 10.0 | 0.0 | 49.9967172E+00 | DC220MV |  | 0.60000201212499993 | 3.4531150990859465e-07 | 1517795732.76 | 1517795748.189 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
| 20 | 0.6 | 0.2 | 0.0 | 10 | 0.6 | 20.0 | 10.0 | 0.0 | 49.9859992E+00 | DC220MV |  | 0.60000217790499977 | 2.6881530922487795e-07 | 1517795763.038 | 1517795778.471 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6
| 20 | 0.6 | 0.2 | 0.1 | 10 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0162553E+00 | DC220MV |  | 0.50000204012999994 | 3.527457386024051e-07 | 1517795793.297 | 1517795808.731 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 20 | 0.5 | 0.2 | 0.0 | 10 | 0.5 | 20.0 | 10.0 | 0.0 | 50.0037408E+00 | DC220MV |  | 0.50000192468499993 | 3.0710626869697361e-07 | 1517795823.566 | 1517795839.002 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
| 20 | 0.5 | 0.2 | 0.1 | 10 | 0.4 | 20.0 | 10.0 | 0.0 | 50.0125503E+00 | DC220MV |  | 0.40000164160000001 | 3.0289173642969185e-07 | 1517795853.835 | 1517795869.279 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 20 | 0.4 | 0.2 | 0.0 | 10 | 0.4 | 20.0 | 10.0 | 0.0 | 49.9997972E+00 | DC220MV |  | 0.40000170751500008 | 3.0125652404229694e-07 | 1517795884.107 | 1517795899.556 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
| 20 | 0.4 | 0.2 | 0.1 | 10 | 0.3 | 20.0 | 10.0 | 0.0 | 50.0141172E+00 | DC220MV |  | 0.30000130206999992 | 3.2270952899025446e-07 | 1517795914.383 | 1517795929.798 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 20 | 0.3 | 0.2 | 0.0 | 10 | 0.3 | 20.0 | 10.0 | 0.0 | 49.9926237E+00 | DC220MV |  | 0.30000135358499996 | 3.3112013118380616e-07 | 1517795944.634 | 1517795960.051 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3
| 20 | 0.3 | 0.2 | 0.1 | 10 | 0.2 | 20.0 | 10.0 | 0.0 | 49.9501394E+00 | DC220MV |  | 0.20000113414500001 | 2.8383495463993714e-07 | 1517795974.887 | 1517795990.324 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 20 | 0.2 | 0.2 | 0.0 | 10 | 0.2 | 20.0 | 10.0 | 0.0 | 50.0237566E+00 | DC220MV |  | 0.20000129841 | 3.5809653712074602e-07 | 1517796005.173 | 1517796020.624 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
| 20 | 0.2 | 0.2 | 0.1 | 10 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0233316E+00 | DC220MV |  | 0.10000083564999999 | 3.1473277792422531e-07 | 1517796035.444 | 1517796050.86 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.1 | 0.2 | 0.0 | 10 | 0.1 | 20.0 | 10.0 | 0.0 | 50.0751836E+00 | DC220MV |  | 0.10000064968 | 2.2959013829049435e-07 | 1517796065.709 | 1517796081.131 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | 0.1 | 0.2 | 0.1 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0811314E+00 | DC220MV |  | 3.0307273321499993e-07 | 3.5402959359406439e-07 | 1517796095.973 | 1517796111.417 | 1.3273e-07 | -1.1798e-07 | 3.687e-07 | 3.5395e-07 | 1.1798e-07 | 9.4388e-07 | -2.8021e-07 | 3.982e-07 | -2.5072e-07 | 4.5719e-07 | -2.6547e-07 | 5.0143e-07 | 5.4568e-07 | 1.1798e-07 | 2.3597e-07 | 9.2913e-07 | 4.8669e-07 | 1.7698e-07 | 8.1114e-07 | 3.982e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0869670E+00 | DC220MV |  | 3.0439960856000001e-07 | 3.4684294159338282e-07 | 1517796126.267 | 1517796141.717 | 1.1793e-07 | 3.8326e-07 | 8.5497e-07 | -1.0319e-07 | 4.8645e-07 | 4.4223e-07 | 9.1394e-07 | 1.0319e-07 | -5.4541e-07 | 3.5378e-07 | 4.7171e-07 | 4.5697e-07 | 6.3386e-07 | 0 | 2.2111e-07 | -2.506e-07 | 2.506e-07 | 4.7171e-07 | 2.0637e-07 | 6.1912e-07
| 20 | 0.1 | 0.2 | 0.1 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.8742369E+00 | DC220MV |  | 2.4916864362950001e-07 | 2.8588727866914401e-07 | 1517796156.556 | 1517796172.013 | 3.0962e-07 | 3.6859e-07 | 3.0962e-07 | 5.75e-07 | 1.4744e-08 | -1.6218e-07 | 5.8975e-08 | 7.8142e-07 | 7.077e-07 | 2.9487e-07 | 4.8654e-07 | 8.8462e-08 | 5.8975e-08 | 4.718e-07 | 2.9487e-08 | 5.6026e-07 | -1.0321e-07 | -1.3269e-07 | -1.3269e-07 | 3.9808e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0058073E+00 | DC220MV |  | 3.0504423861549999e-07 | 3.3571185462764293e-07 | 1517796186.851 | 1517796202.379 | -1.3231e-07 | 3.3812e-07 | -3.0872e-07 | -8.8206e-08 | 3.5282e-07 | 3.8222e-07 | 8.9676e-07 | 5.2923e-07 | 3.3812e-07 | 8.3795e-07 | -2.4992e-07 | 4.2633e-07 | 6.3214e-07 | 2.7932e-07 | 7.0564e-07 | 6.3214e-07 | 4.4103e-08 | 1.4701e-07 | 1.0291e-07 | 2.3521e-07
| 20 | 0.1 | 0.2 | 0.0 | 10 | 0.1 | 20.0 | 10.0 | 0.0 | 49.9948071E+00 | DC220MV |  | 0.100000527871 | 3.1171207793166212e-07 | 1517796217.207 | 1517796232.639 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1
| 20 | -0.1 | 0.2 | 0.0 | 10 | -0.1 | 20.0 | 10.0 | 0.0 | 50.0471624E+00 | DC220MV |  | -0.099999137560500023 | 2.9665230721143245e-07 | 1517796247.476 | 1517796262.934 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0115404E+00 | DC220MV |  | 5.8509364789700011e-07 | 2.9732910002926228e-07 | 1517796277.764 | 1517796293.229 | 8.1058e-07 | 3.6845e-07 | 7.8111e-07 | 7.9585e-07 | 8.548e-07 | 6.7794e-07 | 2.9476e-07 | 5.453e-07 | 2.5054e-07 | 1.6212e-07 | -1.4738e-08 | 3.8318e-07 | 8.1058e-07 | 1.0759e-06 | 5.0109e-07 | 5.3056e-07 | 6.4847e-07 | 5.3056e-07 | 1.2232e-06 | 4.7161e-07
| 20 | -0.1 | 0.2 | -0.1 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9215290E+00 | DC220MV |  | 1.3131255344550001e-06 | 2.6522216926275679e-07 | 1517796308.086 | 1517796323.557 | 9.7214e-07 | 1.6644e-06 | 1.4435e-06 | 1.1194e-06 | 1.3993e-06 | 1.1047e-06 | 8.2485e-07 | 1.1784e-06 | 1.4435e-06 | 1.7233e-06 | 1.2815e-06 | 1.3993e-06 | 1.1636e-06 | 1.2667e-06 | 1.1047e-06 | 1.5613e-06 | 1.4582e-06 | 1.3551e-06 | 1.8706e-06 | 9.2795e-07
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9551346E+00 | DC220MV |  | 5.9140220118500005e-07 | 2.7334429164947055e-07 | 1517796338.387 | 1517796353.925 | 2.3509e-07 | 7.4935e-07 | 7.3466e-07 | 7.4935e-07 | 3.0856e-07 | 7.1997e-07 | 3.8202e-07 | 8.8159e-07 | 2.6448e-07 | 9.4037e-07 | 4.1141e-07 | -1.3224e-07 | 5.7304e-07 | 7.1997e-07 | 8.9629e-07 | 8.8159e-07 | 3.9672e-07 | 7.7874e-07 | 6.0242e-07 | 7.3466e-07
| 20 | -0.1 | 0.2 | -0.1 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0111959E+00 | DC220MV |  | 1.2910186439599998e-06 | 2.9395668436720801e-07 | 1517796368.765 | 1517796384.278 | 1.5586e-06 | 1.4263e-06 | 8.8225e-07 | 1.838e-06 | 1.6322e-06 | 8.2343e-07 | 1.0881e-06 | 7.9402e-07 | 1.0734e-06 | 1.044e-06 | 1.3675e-06 | 1.1616e-06 | 1.4557e-06 | 1.3234e-06 | 1.5586e-06 | 1.1322e-06 | 1.4704e-06 | 1.7645e-06 | 1.3234e-06 | 1.1028e-06
| 20 | -0.1 | 0.2 | 0.0 | 10 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9645911E+00 | DC220MV |  | -0.099999194603000024 | 4.366478623587047e-07 | 1517796399.125 | 1517796414.576 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999
| 20 | -0.2 | 0.2 | -0.1 | 10 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9648266E+00 | DC220MV |  | -0.099999332841500005 | 3.4629267616125292e-07 | 1517796429.406 | 1517796444.891 | -0.1 | -0.1 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.1 | -0.1 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999
| 20 | -0.2 | 0.2 | 0.0 | 10 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9845708E+00 | DC220MV |  | -0.19999969101500001 | 2.8173531776399231e-07 | 1517796459.731 | 1517796475.192 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 20 | -0.3 | 0.2 | -0.1 | 10 | -0.2 | 20.0 | 10.0 | 0.0 | 50.0274207E+00 | DC220MV |  | -0.19999953446000002 | 3.2900275743421597e-07 | 1517796490.022 | 1517796505.464 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 20 | -0.3 | 0.2 | 0.0 | 10 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9549416E+00 | DC220MV |  | -0.29999997982999999 | 2.9620959150409819e-07 | 1517796520.312 | 1517796535.736 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 20 | -0.4 | 0.2 | -0.1 | 10 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9522061E+00 | DC220MV |  | -0.29999994071999997 | 2.6531101861415773e-07 | 1517796550.585 | 1517796566.052 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 20 | -0.4 | 0.2 | 0.0 | 10 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0307595E+00 | DC220MV |  | -0.40000032065499996 | 2.6636100591624216e-07 | 1517796580.908 | 1517796596.359 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 20 | -0.5 | 0.2 | -0.1 | 10 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0128193E+00 | DC220MV |  | -0.40000026305500003 | 3.3691742234271095e-07 | 1517796611.177 | 1517796626.603 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 20 | -0.5 | 0.2 | 0.0 | 10 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9715539E+00 | DC220MV |  | -0.5000006452850001 | 2.7001808693601166e-07 | 1517796641.429 | 1517796656.865 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 20 | -0.6 | 0.2 | -0.1 | 10 | -0.5 | 20.0 | 10.0 | 0.0 | 50.0230342E+00 | DC220MV |  | -0.50000065161500007 | 3.1258537117775189e-07 | 1517796671.706 | 1517796687.145 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 20 | -0.6 | 0.2 | 0.0 | 10 | -0.6 | 20.0 | 10.0 | 0.0 | 50.0284690E+00 | DC220MV |  | -0.60000110318000011 | 2.7763799920384328e-07 | 1517796701.972 | 1517796717.424 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 20 | -0.7 | 0.2 | -0.1 | 10 | -0.6 | 20.0 | 10.0 | 0.0 | 49.9850941E+00 | DC220MV |  | -0.60000091043500003 | 3.2409666656604965e-07 | 1517796732.288 | 1517796747.689 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 20 | -0.7 | 0.2 | 0.0 | 10 | -0.7 | 20.0 | 10.0 | 0.0 | 49.9291195E+00 | DC220MV |  | -0.70000113938500008 | 2.3419403124407074e-07 | 1517796762.542 | 1517796777.983 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 20 | -0.8 | 0.2 | -0.1 | 10 | -0.7000000000000001 | 20.0 | 10.0 | 0.0 | 50.0291727E+00 | DC220MV |  | -0.7000012766799999 | 2.7491540080359341e-07 | 1517796792.809 | 1517796808.282 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 20 | -0.8 | 0.2 | 0.0 | 10 | -0.8 | 20.0 | 10.0 | 0.0 | 50.0302872E+00 | DC220MV |  | -0.80000158663499987 | 3.6704806806846078e-07 | 1517796823.131 | 1517796838.55 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 20 | -0.9 | 0.2 | -0.1 | 10 | -0.8 | 20.0 | 10.0 | 0.0 | 50.0492891E+00 | DC220MV |  | -0.80000148523499992 | 3.0043168321544311e-07 | 1517796853.39 | 1517796868.807 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 20 | -0.9 | 0.2 | 0.0 | 10 | -0.9 | 20.0 | 10.0 | 0.0 | 50.0732818E+00 | DC220MV |  | -0.90000188815999993 | 2.8087866490969876e-07 | 1517796883.66 | 1517796899.085 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 20 | -1.0 | 0.2 | -0.1 | 10 | -0.9 | 20.0 | 10.0 | 0.0 | 49.9538822E+00 | DC220MV |  | -0.90000189986500012 | 3.2267923745412757e-07 | 1517796913.92 | 1517796929.328 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 20 | -1.0 | 0.2 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9980284E+00 | DC220MV |  | -1.0000022652999998 | 2.8383659030405131e-07 | 1517796944.156 | 1517796959.593 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -1.0 | 0.2 | -0.1 | 10 | -0.9 | 20.0 | 10.0 | 0.0 | 49.9962597E+00 | DC220MV |  | -0.90000189404000008 | 2.7641342659151748e-07 | 1517796974.422 | 1517796989.842 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 20 | -0.9 | 0.2 | 0.0 | 10 | -0.9 | 20.0 | 10.0 | 0.0 | 50.0168547E+00 | DC220MV |  | -0.90000199385999979 | 2.337969854543268e-07 | 1517797004.698 | 1517797020.138 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | -0.9
| 20 | -0.9 | 0.2 | -0.1 | 10 | -0.8 | 20.0 | 10.0 | 0.0 | 50.0096056E+00 | DC220MV |  | -0.8000014229450001 | 3.3282493368042393e-07 | 1517797034.973 | 1517797050.393 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 20 | -0.8 | 0.2 | 0.0 | 10 | -0.8 | 20.0 | 10.0 | 0.0 | 49.9626411E+00 | DC220MV |  | -0.80000177619500001 | 3.2209425247907867e-07 | 1517797065.246 | 1517797080.677 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8 | -0.8
| 20 | -0.8 | 0.2 | -0.1 | 10 | -0.7000000000000001 | 20.0 | 10.0 | 0.0 | 49.9963588E+00 | DC220MV |  | -0.70000112472500009 | 3.15614750712047e-07 | 1517797095.507 | 1517797110.946 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 20 | -0.7 | 0.2 | 0.0 | 10 | -0.7 | 20.0 | 10.0 | 0.0 | 49.9514764E+00 | DC220MV |  | -0.700001277915 | 3.0928054298671544e-07 | 1517797125.773 | 1517797141.219 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7 | -0.7
| 20 | -0.7 | 0.2 | -0.1 | 10 | -0.6 | 20.0 | 10.0 | 0.0 | 50.0244601E+00 | DC220MV |  | -0.60000095325999991 | 3.8089390043563601e-07 | 1517797156.042 | 1517797171.487 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 20 | -0.6 | 0.2 | 0.0 | 10 | -0.6 | 20.0 | 10.0 | 0.0 | 49.9727601E+00 | DC220MV |  | -0.60000122201499984 | 2.9608593224866983e-07 | 1517797186.341 | 1517797201.754 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6 | -0.6
| 20 | -0.6 | 0.2 | -0.1 | 10 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9984387E+00 | DC220MV |  | -0.50000071633499998 | 3.7603511841167103e-07 | 1517797216.597 | 1517797232.031 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 20 | -0.5 | 0.2 | 0.0 | 10 | -0.5 | 20.0 | 10.0 | 0.0 | 49.9730476E+00 | DC220MV |  | -0.50000080008500003 | 2.7189736533253838e-07 | 1517797246.852 | 1517797262.27 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | -0.5
| 20 | -0.5 | 0.2 | -0.1 | 10 | -0.4 | 20.0 | 10.0 | 0.0 | 50.0534013E+00 | DC220MV |  | -0.40000023056 | 2.514884001332485e-07 | 1517797277.099 | 1517797292.541 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 20 | -0.4 | 0.2 | 0.0 | 10 | -0.4 | 20.0 | 10.0 | 0.0 | 49.9344728E+00 | DC220MV |  | -0.40000046559999997 | 2.3323436710695532e-07 | 1517797307.378 | 1517797322.773 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4 | -0.4
| 20 | -0.4 | 0.2 | -0.1 | 10 | -0.3 | 20.0 | 10.0 | 0.0 | 49.9875173E+00 | DC220MV |  | -0.29999998533000005 | 2.5515829616274661e-07 | 1517797337.602 | 1517797353.063 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 20 | -0.3 | 0.2 | 0.0 | 10 | -0.3 | 20.0 | 10.0 | 0.0 | 50.0609847E+00 | DC220MV |  | -0.29999995798000001 | 2.6003038784066115e-07 | 1517797367.913 | 1517797383.341 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3 | -0.3
| 20 | -0.3 | 0.2 | -0.1 | 10 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9098515E+00 | DC220MV |  | -0.19999965624999999 | 2.9898427801496347e-07 | 1517797398.167 | 1517797413.573 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 20 | -0.2 | 0.2 | 0.0 | 10 | -0.2 | 20.0 | 10.0 | 0.0 | 49.9229961E+00 | DC220MV |  | -0.19999974784500002 | 4.0301888103990699e-07 | 1517797428.411 | 1517797443.893 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2 | -0.2
| 20 | -0.2 | 0.2 | -0.1 | 10 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9609739E+00 | DC220MV |  | -0.099999358197499999 | 2.9606950259913353e-07 | 1517797458.742 | 1517797474.211 | -0.1 | -0.099999 | -0.1 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.1 | -0.1 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999
| 20 | -0.1 | 0.2 | 0.0 | 10 | -0.1 | 20.0 | 10.0 | 0.0 | 50.0187854E+00 | DC220MV |  | -0.099999124574500026 | 2.8795691515375599e-07 | 1517797489.048 | 1517797504.505 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.1 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999
| 20 | -0.1 | 0.2 | -0.1 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0022548E+00 | DC220MV |  | 1.3388077651449999e-06 | 2.7908291655881836e-07 | 1517797519.346 | 1517797534.82 | 1.4867e-06 | 9.421e-07 | 1.2807e-06 | 1.6634e-06 | 7.8017e-07 | 1.4279e-06 | 1.2659e-06 | 1.0893e-06 | 1.1629e-06 | 1.7664e-06 | 1.1187e-06 | 1.001e-06 | 1.8548e-06 | 1.6928e-06 | 1.4426e-06 | 1.3101e-06 | 1.2954e-06 | 1.3395e-06 | 1.2218e-06 | 1.6339e-06
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9609739E+00 | DC220MV |  | 5.5030810628e-07 | 1.6599429689320859e-07 | 1517797549.637 | 1517797565.094 | 4.1255e-07 | 3.9781e-07 | 8.1036e-07 | 5.8935e-07 | 4.5675e-07 | 6.0409e-07 | 6.4829e-07 | 6.3355e-07 | 4.8622e-07 | 2.9468e-07 | 5.8935e-07 | 5.5989e-07 | 7.0722e-07 | 6.6302e-07 | 7.2196e-07 | 1.9154e-07 | 5.5989e-07 | 6.9249e-07 | 2.5047e-07 | 7.3669e-07
| 20 | -0.1 | 0.2 | -0.1 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0177799E+00 | DC220MV |  | 1.38590013471e-06 | 2.403832103192503e-07 | 1517797579.947 | 1517797595.432 | 1.9715e-06 | 1.6919e-06 | 1.2947e-06 | 1.3388e-06 | 1.3241e-06 | 1.1034e-06 | 1.3977e-06 | 1.6036e-06 | 1.3094e-06 | 1.383e-06 | 1.2505e-06 | 9.7101e-07 | 1.2653e-06 | 1.6184e-06 | 1.0004e-06 | 1.1476e-06 | 1.5301e-06 | 1.6184e-06 | 1.5595e-06 | 1.3388e-06
| 20 | 0.0 | 0.2 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9662115E+00 | DC220MV |  | 9.1266866922999996e-07 | 3.4055222507527639e-07 | 1517797610.272 | 1517797625.729 | 1.1197e-06 | 6.4822e-07 | 1.3554e-06 | 8.2501e-07 | 1.046e-06 | 1.0165e-06 | 1.3259e-06 | 8.5448e-07 | 6.7769e-07 | 1.046e-06 | 1.3112e-06 | 5.451e-07 | 1.6206e-07 | 1.1197e-06 | 7.2188e-07 | 1.0018e-06 | 2.0625e-07 | 1.2375e-06 | 1.2964e-06 | 7.3662e-07
| 20 | -0.1 | 0.2 | 0.0 | 10 | -0.1 | 20.0 | 10.0 | 0.0 | 49.9656510E+00 | DC220MV |  | -0.099999208149000013 | 3.2818997679514863e-07 | 1517797640.564 | 1517797656.033 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.1 | -0.099999 | -0.099999 | -0.099999 | -0.099999 | -0.099999
| 20 | 1.0 | 2.0 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 50.0282943E+00 | DC2_2V |  | 1.00000355905 | 3.4424373270933927e-07 | 1517797671.184 | 1517797686.616 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0615521E+00 | DC2_2V |  | 6.7446115840000007e-07 | 3.3055703721242407e-07 | 1517797701.457 | 1517797716.926 | 7.6576e-07 | 9.572e-07 | 7.3631e-07 | 1.0603e-06 | 8.3939e-07 | 8.5412e-07 | 4.7124e-07 | 6.185e-07 | 7.2159e-07 | 1.2517e-06 | 3.0925e-07 | 2.9452e-07 | -1.3254e-07 | 6.0378e-07 | 5.3014e-07 | 5.4487e-07 | 1.3254e-06 | 8.0994e-07 | 5.0069e-07 | 4.2706e-07
| 20 | 1.0 | 2.0 | 1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9863575E+00 | DC2_2V |  | -2.5784300044999999e-06 | 2.2724763859596506e-07 | 1517797731.749 | 1517797747.19 | -2.5644e-06 | -2.6086e-06 | -2.7118e-06 | -2.8592e-06 | -2.5792e-06 | -2.2255e-06 | -2.3876e-06 | -2.5644e-06 | -3.0655e-06 | -2.4318e-06 | -2.8739e-06 | -2.5055e-06 | -2.2107e-06 | -2.6086e-06 | -2.6971e-06 | -2.5055e-06 | -2.6086e-06 | -2.7118e-06 | -2.7413e-06 | -2.1075e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0189789E+00 | DC2_2V |  | 6.5054771449900006e-07 | 3.7329326920837919e-07 | 1517797762.049 | 1517797777.535 | 8.831e-07 | 9.8612e-07 | 7.8007e-07 | 1.0597e-06 | 9.2725e-07 | 8.6838e-07 | 1.4718e-08 | 2.9437e-07 | 8.2422e-07 | 8.831e-07 | 1.1922e-06 | 2.9437e-07 | -1.0303e-07 | 9.5669e-07 | 2.0606e-07 | 1.7662e-07 | 5.0042e-07 | 4.5627e-07 | 9.4197e-07 | 8.6838e-07
| 20 | 1.0 | 2.0 | 1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0155426E+00 | DC2_2V |  | -2.6327438166999998e-06 | 3.80065039046059e-07 | 1517797792.353 | 1517797807.833 | -2.4135e-06 | -2.7372e-06 | -2.6931e-06 | -2.8108e-06 | -2.4871e-06 | -3.4878e-06 | -3.0168e-06 | -3.0316e-06 | -1.9131e-06 | -2.1339e-06 | -2.8255e-06 | -2.2222e-06 | -2.2074e-06 | -2.384e-06 | -2.5754e-06 | -2.8402e-06 | -2.8108e-06 | -2.3105e-06 | -3.2082e-06 | -2.5459e-06
| 20 | 1.0 | 2.0 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 50.0273451E+00 | DC2_2V |  | 1.0000033260000003 | 2.7695107869181643e-07 | 1517797822.65 | 1517797838.054 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 2.0 | 2.0 | 1.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9857163E+00 | DC2_2V |  | 1.0000005452849998 | 3.427967083817426e-07 | 1517797852.899 | 1517797868.342 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 2.0 | 2.0 | 0.0 | 10 | 2.0 | 20.0 | 10.0 | 0.0 | 50.0819549E+00 | DC2_2V |  | 2.0000069329999999 | 3.8300182770687937e-07 | 1517797883.177 | 1517797898.621 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 20 | 3.0 | 2.0 | 1.0 | 10 | 2.0 | 20.0 | 10.0 | 0.0 | 49.9733067E+00 | DC2_2V |  | 2.0000036567999997 | 2.8114490927326597e-07 | 1517797913.456 | 1517797928.834 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 20 | 3.0 | 2.0 | 0.0 | 10 | 3.0 | 20.0 | 10.0 | 0.0 | 49.9433238E+00 | DC2_2V |  | 3.0000100698999992 | 3.2803550114001471e-07 | 1517797943.675 | 1517797959.105 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
| 20 | 4.0 | 2.0 | 1.0 | 10 | 3.0 | 20.0 | 10.0 | 0.0 | 49.9381615E+00 | DC2_2V |  | 3.0000069381000003 | 3.355890195169776e-07 | 1517797973.928 | 1517797989.363 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
| 20 | 4.0 | 2.0 | 0.0 | 10 | 4.0 | 20.0 | 10.0 | 0.0 | 50.0338531E+00 | DC2_2V |  | 4.0000132633500005 | 3.1397583906493683e-07 | 1517798004.186 | 1517798019.627 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
| 20 | 5.0 | 2.0 | 1.0 | 10 | 4.0 | 20.0 | 10.0 | 0.0 | 50.0265896E+00 | DC2_2V |  | 4.0000101756000008 | 3.3642062948538227e-07 | 1517798034.445 | 1517798049.841 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
| 20 | 5.0 | 2.0 | 0.0 | 10 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0390210E+00 | DC2_2V |  | 5.0000163272500009 | 2.5077537264221008e-07 | 1517798064.66 | 1517798080.079 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 20 | 5.0 | 2.0 | 1.0 | 10 | 4.0 | 20.0 | 10.0 | 0.0 | 50.0465858E+00 | DC2_2V |  | 4.0000102832000008 | 3.1852764398416351e-07 | 1517798094.928 | 1517798110.322 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
| 20 | 4.0 | 2.0 | 0.0 | 10 | 4.0 | 20.0 | 10.0 | 0.0 | 49.9890401E+00 | DC2_2V |  | 4.0000134525000011 | 3.3652615058756015e-07 | 1517798125.139 | 1517798140.542 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
| 20 | 4.0 | 2.0 | 1.0 | 10 | 3.0 | 20.0 | 10.0 | 0.0 | 50.0400320E+00 | DC2_2V |  | 3.0000074412000002 | 3.0051632230958678e-07 | 1517798155.371 | 1517798170.784 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
| 20 | 3.0 | 2.0 | 0.0 | 10 | 3.0 | 20.0 | 10.0 | 0.0 | 50.0445680E+00 | DC2_2V |  | 3.0000102397000004 | 2.7626040252038506e-07 | 1517798185.611 | 1517798201.018 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
| 20 | 3.0 | 2.0 | 1.0 | 10 | 2.0 | 20.0 | 10.0 | 0.0 | 50.0502864E+00 | DC2_2V |  | 2.0000039415000002 | 3.3671553275464432e-07 | 1517798215.855 | 1517798231.249 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 20 | 2.0 | 2.0 | 0.0 | 10 | 2.0 | 20.0 | 10.0 | 0.0 | 50.0126022E+00 | DC2_2V |  | 2.0000071981499996 | 2.2344871341161637e-07 | 1517798246.071 | 1517798261.443 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 20 | 2.0 | 2.0 | 1.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9537221E+00 | DC2_2V |  | 1.0000008408999999 | 2.5527179633198708e-07 | 1517798276.272 | 1517798291.676 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 1.0 | 2.0 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9989481E+00 | DC2_2V |  | 1.0000034273 | 3.172990860557636e-07 | 1517798306.538 | 1517798321.981 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 1.0 | 2.0 | 1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0350529E+00 | DC2_2V |  | -2.55466795665e-06 | 2.5115634954375563e-07 | 1517798336.833 | 1517798352.299 | -2.4148e-06 | -2.3853e-06 | -2.4148e-06 | -2.4148e-06 | -2.5915e-06 | -2.4442e-06 | -2.8418e-06 | -2.6945e-06 | -2.0172e-06 | -2.9154e-06 | -2.8123e-06 | -2.562e-06 | -2.6504e-06 | -2.886e-06 | -2.6798e-06 | -2.1497e-06 | -2.1497e-06 | -2.7387e-06 | -2.8271e-06 | -2.5031e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9478987E+00 | DC2_2V |  | 5.3884240003499993e-07 | 2.5420892851873984e-07 | 1517798367.116 | 1517798382.595 | 6.9196e-07 | 4.564e-07 | 6.3307e-07 | 3.9751e-07 | 3.5334e-07 | 1.2072e-06 | 2.65e-07 | 7.6557e-07 | 7.9501e-07 | 8.6863e-07 | 8.3918e-07 | 3.8278e-07 | 2.3556e-07 | 6.0362e-07 | 5.7418e-07 | 2.3556e-07 | 3.5334e-07 | 2.9445e-07 | 5.1529e-07 | 3.0917e-07
| 20 | 1.0 | 2.0 | 1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9323793E+00 | DC2_2V |  | -2.7347089896499997e-06 | 3.6719158066635056e-07 | 1517798397.431 | 1517798412.947 | -3.659e-06 | -2.5569e-06 | -2.5128e-06 | -3.1447e-06 | -2.4687e-06 | -2.8655e-06 | -3.2035e-06 | -2.601e-06 | -1.9838e-06 | -2.41e-06 | -2.5569e-06 | -3.1594e-06 | -3.1594e-06 | -2.8361e-06 | -2.6892e-06 | -2.5716e-06 | -2.7479e-06 | -2.6598e-06 | -2.4981e-06 | -2.41e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0356386E+00 | DC2_2V |  | 4.7571516287850001e-07 | 2.9892849608155825e-07 | 1517798427.79 | 1517798443.291 | 3.3822e-07 | 4.4116e-07 | 3.5293e-07 | 6.6174e-07 | 6.4703e-07 | 4.9998e-07 | 7.3526e-08 | 2.6469e-07 | 4.5586e-07 | 3.9704e-07 | 1.2647e-06 | 4.2645e-07 | 5.588e-07 | 7.3526e-07 | 2.4999e-07 | 6.7644e-07 | 9.1173e-07 | 0 | -2.9411e-08 | 5.8821e-07
| 20 | 1.0 | 2.0 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9750879E+00 | DC2_2V |  | 1.00000333345 | 3.4185193798728614e-07 | 1517798458.133 | 1517798473.523 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9984529E+00 | DC2_2V |  | -1.0000023188000002 | 3.0301726024387861e-07 | 1517798488.344 | 1517798503.77 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9836374E+00 | DC2_2V |  | 1.0023294999049997e-06 | 2.4522531947535001e-07 | 1517798518.602 | 1517798534.082 | 7.3592e-07 | 1.4866e-06 | 8.6839e-07 | 9.8614e-07 | 1.0156e-06 | 1.0892e-06 | 8.2424e-07 | 8.6839e-07 | 1.2364e-06 | 7.8008e-07 | 1.2511e-06 | 1.2805e-06 | 6.7705e-07 | 1.2511e-06 | 7.6536e-07 | 8.3895e-07 | 1.3247e-06 | 6.329e-07 | 1.2658e-06 | 8.6839e-07
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9264198E+00 | DC2_2V |  | 4.2398078557500001e-06 | 2.5859348546478607e-07 | 1517798548.907 | 1517798564.403 | 4.192e-06 | 4.0008e-06 | 4.6333e-06 | 4.1626e-06 | 4.4126e-06 | 4.2067e-06 | 4.0302e-06 | 4.089e-06 | 4.648e-06 | 4.3244e-06 | 4.3391e-06 | 4.3244e-06 | 3.7066e-06 | 4.3832e-06 | 4.1038e-06 | 4.6186e-06 | 4.3979e-06 | 4.2655e-06 | 4.295e-06 | 3.6625e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0385013E+00 | DC2_2V |  | 7.8878352685149991e-07 | 3.0672294832492865e-07 | 1517798579.259 | 1517798594.768 | 6.175e-07 | 8.6744e-07 | 3.0875e-07 | 1.0439e-06 | 1.2938e-06 | 8.8214e-08 | 1.235e-06 | 9.9976e-07 | 3.8226e-07 | 7.7923e-07 | 9.4095e-07 | 1.1027e-06 | 9.8506e-07 | 7.3512e-07 | 5.7339e-07 | 6.322e-07 | 7.3512e-07 | 8.0863e-07 | 1.1027e-06 | 5.4399e-07
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9677896E+00 | DC2_2V |  | 4.4565338577499993e-06 | 3.3486522408899486e-07 | 1517798609.621 | 1517798625.074 | 3.9925e-06 | 4.7438e-06 | 4.3608e-06 | 4.405e-06 | 4.0956e-06 | 4.2576e-06 | 4.6554e-06 | 4.8175e-06 | 4.9501e-06 | 4.4344e-06 | 4.0809e-06 | 4.346e-06 | 4.405e-06 | 4.6554e-06 | 5.0237e-06 | 4.1693e-06 | 4.8617e-06 | 3.7126e-06 | 4.567e-06 | 4.5965e-06
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1 | 20.0 | 10.0 | 0.0 | 49.9958400E+00 | DC2_2V |  | -1.0000022449999999 | 3.1716194601877087e-07 | 1517798639.908 | 1517798655.333 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -2.0 | 2.0 | -1.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9481576E+00 | DC2_2V |  | -0.99999934859500006 | 2.6412069300316794e-07 | 1517798670.174 | 1517798685.604 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -2.0 | 2.0 | 0.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 49.9359548E+00 | DC2_2V |  | -2.0000061084499996 | 3.0333751416352257e-07 | 1517798700.429 | 1517798715.878 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -3.0 | 2.0 | -1.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 50.0011746E+00 | DC2_2V |  | -2.0000027495000001 | 2.4777661313803843e-07 | 1517798730.712 | 1517798746.164 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -3.0 | 2.0 | 0.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 49.9430602E+00 | DC2_2V |  | -3.0000095367999999 | 2.8641134755794415e-07 | 1517798760.996 | 1517798776.429 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -4.0 | 2.0 | -1.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 50.0030568E+00 | DC2_2V |  | -3.00000611905 | 2.8897291833919966e-07 | 1517798791.278 | 1517798806.752 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -4.0 | 2.0 | 0.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 50.0294702E+00 | DC2_2V |  | -4.0000130457000012 | 2.9225897760029071e-07 | 1517798821.602 | 1517798837.04 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -5.0 | 2.0 | -1.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9465289E+00 | DC2_2V |  | -4.0000095816999988 | 3.4610895691356002e-07 | 1517798851.884 | 1517798867.311 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -5.0 | 2.0 | 0.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 50.0763333E+00 | DC2_2V |  | -5.0000165805000005 | 3.2315623160882058e-07 | 1517798882.15 | 1517798897.631 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -5.0 | 2.0 | -1.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9996981E+00 | DC2_2V |  | -4.0000099672499996 | 2.7954371297148167e-07 | 1517798912.546 | 1517798927.946 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -4.0 | 2.0 | 0.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 50.0061140E+00 | DC2_2V |  | -4.0000135968999988 | 3.9687276303104143e-07 | 1517798942.79 | 1517798958.228 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -4.0 | 2.0 | -1.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 50.0414353E+00 | DC2_2V |  | -3.0000066040499997 | 2.2746526661335397e-07 | 1517798973.088 | 1517798988.537 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -3.0 | 2.0 | 0.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 49.9248066E+00 | DC2_2V |  | -3.0000099179499999 | 2.9011954004939561e-07 | 1517799003.384 | 1517799018.795 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -3.0 | 2.0 | -1.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 50.0549140E+00 | DC2_2V |  | -2.0000030328999996 | 3.2898022733827971e-07 | 1517799033.623 | 1517799049.09 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -2.0 | 2.0 | 0.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 50.0542286E+00 | DC2_2V |  | -2.0000065062000001 | 2.2353805047849058e-07 | 1517799063.938 | 1517799079.353 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -2.0 | 2.0 | -1.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9981180E+00 | DC2_2V |  | -0.99999966085999981 | 3.6480129579696988e-07 | 1517799094.203 | 1517799109.646 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 50.0862239E+00 | DC2_2V |  | -1.0000025792499998 | 3.7990852518038362e-07 | 1517799124.472 | 1517799139.93 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0479043E+00 | DC2_2V |  | 4.2367619783499997e-06 | 2.7869416788461817e-07 | 1517799154.767 | 1517799170.243 | 4.8016e-06 | 3.8295e-06 | 4.507e-06 | 4.6396e-06 | 4.5218e-06 | 3.9621e-06 | 4.0799e-06 | 4.0652e-06 | 4.0063e-06 | 4.0357e-06 | 4.1977e-06 | 4.4187e-06 | 3.8737e-06 | 4.0063e-06 | 4.5365e-06 | 4.5365e-06 | 4.3892e-06 | 4.2419e-06 | 4.1683e-06 | 3.9179e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0023680E+00 | DC2_2V |  | 9.7854592542999995e-07 | 2.6772962073935179e-07 | 1517799185.06 | 1517799200.527 | 8.9897e-07 | 7.8107e-07 | 8.9897e-07 | 9.8739e-07 | 6.1896e-07 | 7.3686e-07 | 1.1053e-06 | 1.3263e-06 | 1.0758e-06 | 6.0422e-07 | 1.2232e-06 | 1.1495e-06 | 1.3116e-06 | 1.1348e-06 | 1.5032e-06 | 1.0169e-06 | 4.7159e-07 | 9.4318e-07 | 6.337e-07 | 1.1495e-06
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0133621E+00 | DC2_2V |  | 4.4268529411500004e-06 | 3.3905329262025172e-07 | 1517799215.372 | 1517799230.866 | 4.3422e-06 | 3.989e-06 | 4.1656e-06 | 4.666e-06 | 4.666e-06 | 4.7396e-06 | 4.4158e-06 | 4.7249e-06 | 4.092e-06 | 4.2097e-06 | 4.46e-06 | 4.2392e-06 | 4.092e-06 | 4.0331e-06 | 4.0478e-06 | 5.1371e-06 | 4.6072e-06 | 5.0635e-06 | 4.7249e-06 | 4.1214e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0186201E+00 | DC2_2V |  | 6.0203460275999996e-07 | 2.601637041360128e-07 | 1517799245.705 | 1517799261.204 | 7.507e-07 | 6.6239e-07 | 8.6846e-07 | 5.7407e-07 | 3.5327e-07 | 4.2687e-07 | 1.0304e-06 | 4.2687e-07 | 5.8879e-07 | 5.7407e-07 | 2.9439e-07 | 5.5935e-07 | 3.2383e-07 | 7.507e-07 | 6.3295e-07 | 2.7967e-07 | 1.2953e-06 | 6.3295e-07 | 7.8014e-07 | 2.3551e-07
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9927275E+00 | DC2_2V |  | -1.0000024726999999 | 2.8866106423481991e-07 | 1517799276.038 | 1517799291.462 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | 1.0 | 2.0 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 50.0029766E+00 | DC2_2V |  | 1.0000034915500002 | 3.3317389979847114e-07 | 1517799306.304 | 1517799321.744 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0095396E+00 | DC2_2V |  | 6.2271376234600008e-07 | 3.5122413958891964e-07 | 1517799336.604 | 1517799352.095 | 5.1525e-07 | 1.2219e-06 | 4.7108e-07 | 4.7108e-07 | 0 | 8.98e-07 | 8.244e-07 | 3.8276e-07 | 2.061e-07 | -2.9443e-08 | 8.98e-07 | 1.2219e-06 | 5.8885e-07 | 8.98e-07 | 7.2135e-07 | 4.7108e-07 | 8.0968e-07 | 2.061e-07 | 6.919e-07 | 9.8633e-07
| 20 | 1.0 | 2.0 | 1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0073879E+00 | DC2_2V |  | -2.6989257044000001e-06 | 2.2330673602757269e-07 | 1517799366.93 | 1517799382.413 | -2.562e-06 | -2.9448e-06 | -2.5914e-06 | -2.9006e-06 | -2.6209e-06 | -2.297e-06 | -2.827e-06 | -2.297e-06 | -2.9301e-06 | -2.8712e-06 | -2.7829e-06 | -2.989e-06 | -2.6503e-06 | -3.1068e-06 | -2.4442e-06 | -2.8417e-06 | -2.5473e-06 | -2.6945e-06 | -2.6356e-06 | -2.4442e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9497063E+00 | DC2_2V |  | 4.9980035807499999e-07 | 2.3143812124903068e-07 | 1517799397.237 | 1517799412.726 | 3.6804e-07 | 4.1221e-07 | 3.2388e-07 | 3.6804e-07 | 8.9802e-07 | 5.447e-07 | 5.447e-07 | 7.508e-07 | 5.5942e-07 | 5.7414e-07 | 6.0359e-07 | 5.7414e-07 | 0 | 8.2441e-07 | 2.3555e-07 | 2.6499e-07 | 8.0969e-07 | 6.1831e-07 | 1.1777e-07 | 6.0359e-07
| 20 | 1.0 | 2.0 | 1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0056893E+00 | DC2_2V |  | -2.8650390060499999e-06 | 2.8992650376700249e-07 | 1517799427.555 | 1517799443.077 | -2.5431e-06 | -2.6313e-06 | -2.3961e-06 | -2.8077e-06 | -3.2487e-06 | -2.5431e-06 | -2.5725e-06 | -3.3075e-06 | -3.2193e-06 | -2.7636e-06 | -2.6607e-06 | -3.2928e-06 | -2.6901e-06 | -2.8665e-06 | -2.6901e-06 | -2.9547e-06 | -3.1458e-06 | -2.7195e-06 | -2.8959e-06 | -3.3516e-06
| 20 | 1.0 | 2.0 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9398602E+00 | DC2_2V |  | 1.0000032399000001 | 2.6270951638178479e-07 | 1517799457.926 | 1517799473.337 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 2.0 | 2.0 | 1.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 50.1151419E+00 | DC2_2V |  | 1.0000006762 | 2.881712338076062e-07 | 1517799488.17 | 1517799503.627 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 2.0 | 2.0 | 0.0 | 10 | 2.0 | 20.0 | 10.0 | 0.0 | 50.0256452E+00 | DC2_2V |  | 2.0000068501499997 | 3.0139397387772427e-07 | 1517799518.475 | 1517799533.858 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 20 | 3.0 | 2.0 | 1.0 | 10 | 2.0 | 20.0 | 10.0 | 0.0 | 49.9429425E+00 | DC2_2V |  | 2.0000036454499996 | 3.5702261481793192e-07 | 1517799548.704 | 1517799564.118 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 20 | 3.0 | 2.0 | 0.0 | 10 | 3.0 | 20.0 | 10.0 | 0.0 | 49.9133532E+00 | DC2_2V |  | 3.0000098386000005 | 3.047310945690739e-07 | 1517799578.97 | 1517799594.437 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
| 20 | 4.0 | 2.0 | 1.0 | 10 | 3.0 | 20.0 | 10.0 | 0.0 | 49.9821950E+00 | DC2_2V |  | 3.0000068685999994 | 3.8286308260843291e-07 | 1517799609.275 | 1517799624.738 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
| 20 | 4.0 | 2.0 | 0.0 | 10 | 4.0 | 20.0 | 10.0 | 0.0 | 50.0195879E+00 | DC2_2V |  | 4.0000131919000008 | 3.4311541796436455e-07 | 1517799639.569 | 1517799655.022 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
| 20 | 5.0 | 2.0 | 1.0 | 10 | 4.0 | 20.0 | 10.0 | 0.0 | 50.1185872E+00 | DC2_2V |  | 4.0000102430999993 | 2.8831075246488963e-07 | 1517799669.878 | 1517799685.308 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
| 20 | 5.0 | 2.0 | 0.0 | 10 | 5.0 | 20.0 | 10.0 | 0.0 | 49.9557467E+00 | DC2_2V |  | 5.00001627375 | 2.6188582912902613e-07 | 1517799700.158 | 1517799715.544 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 20 | 6.0 | 2.0 | 1.0 | 10 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0224251E+00 | DC2_2V |  | 5.0000134210999994 | 3.463024545249046e-07 | 1517799730.381 | 1517799745.846 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 20 | 6.0 | 2.0 | 0.0 | 10 | 6.0 | 20.0 | 10.0 | 0.0 | 49.9617934E+00 | DC2_2V |  | 6.0000198939499994 | 3.1881020605704035e-07 | 1517799760.704 | 1517799776.135 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6
| 20 | 7.0 | 2.0 | 1.0 | 10 | 6.0 | 20.0 | 10.0 | 0.0 | 49.9973916E+00 | DC2_2V |  | 6.000017055699999 | 3.3797072354389233e-07 | 1517799790.959 | 1517799806.415 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6
| 20 | 7.0 | 2.0 | 0.0 | 10 | 7.0 | 20.0 | 10.0 | 0.0 | 49.9497063E+00 | DC2_2V |  | 7.0000229556000004 | 4.2329297173678399e-07 | 1517799821.256 | 1517799836.687 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
| 20 | 8.0 | 2.0 | 1.0 | 10 | 7.0 | 20.0 | 10.0 | 0.0 | 49.9878992E+00 | DC2_2V |  | 7.0000200093 | 3.8522682934400941e-07 | 1517799851.514 | 1517799866.969 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
| 20 | 8.0 | 2.0 | 0.0 | 10 | 8.0 | 20.0 | 10.0 | 0.0 | 49.9583790E+00 | DC2_2V |  | 8.000025695699998 | 3.6170749215232917e-07 | 1517799881.826 | 1517799897.257 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8
| 20 | 9.0 | 2.0 | 1.0 | 10 | 8.0 | 20.0 | 10.0 | 0.0 | 49.9880972E+00 | DC2_2V |  | 8.0000229131500014 | 6.6373340092106927e-07 | 1517799912.115 | 1517799927.563 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8
| 20 | 9.0 | 2.0 | 0.0 | 10 | 9.0 | 20.0 | 10.0 | 0.0 | 50.0899303E+00 | DC2_2V |  | 9.0000288547499991 | 3.3067595542663783e-07 | 1517799942.414 | 1517799957.845 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
| 20 | 10.0 | 2.0 | 1.0 | 10 | 9.0 | 20.0 | 10.0 | 0.0 | 50.0095396E+00 | DC2_2V |  | 9.0000257479000005 | 2.919386751462978e-07 | 1517799972.691 | 1517799988.099 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
| 20 | 10.0 | 2.0 | 0.0 | 10 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9750078E+00 | DC2_2V |  | 10.000031805000001 | 4.0145360878068454e-07 | 1517800002.957 | 1517800018.367 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
| 20 | 10.0 | 2.0 | 1.0 | 10 | 9.0 | 20.0 | 10.0 | 0.0 | 50.0209286E+00 | DC2_2V |  | 9.0000199849499989 | 2.4199885827251243e-07 | 1517951464.316 | 1517951479.762 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
| 20 | 9.0 | 2.0 | 0.0 | 10 | 9.0 | 20.0 | 10.0 | 0.0 | 49.9747533E+00 | DC220MV |  | 9.0000242120000014 | 2.9913943256535047e-07 | 1517951494.599 | 1517951510.019 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
| 20 | 9.0 | 2.0 | 1.0 | 10 | 8.0 | 20.0 | 10.0 | 0.0 | 50.0531555E+00 | DC2_2V |  | 8.0000186410500014 | 2.8680245376819811e-07 | 1517951524.874 | 1517951540.329 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8
| 20 | 8.0 | 2.0 | 0.0 | 10 | 8.0 | 20.0 | 10.0 | 0.0 | 49.9809789E+00 | DC220MV |  | 8.0000220663000015 | 2.4321350705351536e-07 | 1517951555.172 | 1517951570.581 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8
| 20 | 8.0 | 2.0 | 1.0 | 10 | 7.0 | 20.0 | 10.0 | 0.0 | 50.0451965E+00 | DC2_2V |  | 7.0000163317499995 | 2.3739795180898987e-07 | 1517951585.434 | 1517951600.874 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
| 20 | 7.0 | 2.0 | 0.0 | 10 | 7.0 | 20.0 | 10.0 | 0.0 | 50.0347931E+00 | DC220MV |  | 7.0000195328499988 | 2.1045789964326464e-07 | 1517951615.736 | 1517951631.14 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
| 20 | 7.0 | 2.0 | 1.0 | 10 | 6.0 | 20.0 | 10.0 | 0.0 | 49.9868619E+00 | DC2_2V |  | 6.0000137680500014 | 2.4192818673369378e-07 | 1517951645.993 | 1517951661.405 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6
| 20 | 6.0 | 2.0 | 0.0 | 10 | 6.0 | 20.0 | 10.0 | 0.0 | 50.0368148E+00 | DC220MV |  | 6.0000168733500008 | 2.8648564982986413e-07 | 1517951676.237 | 1517951691.681 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6
| 20 | 6.0 | 2.0 | 1.0 | 10 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0033115E+00 | DC2_2V |  | 5.0000108080499999 | 2.5451826542588857e-07 | 1517951706.52 | 1517951721.934 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 20 | 5.0 | 2.0 | 0.0 | 10 | 5.0 | 20.0 | 10.0 | 0.0 | 49.9986085E+00 | DC220MV |  | 5.0000138417500013 | 2.1750169547586061e-07 | 1517951736.784 | 1517951752.213 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 20 | 5.0 | 2.0 | 1.0 | 10 | 4.0 | 20.0 | 10.0 | 0.0 | 49.9820489E+00 | DC2_2V |  | 4.0000080413000001 | 1.8646155098214235e-07 | 1517951767.061 | 1517951782.483 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
| 20 | 4.0 | 2.0 | 0.0 | 10 | 4.0 | 20.0 | 10.0 | 0.0 | 49.9937979E+00 | DC220MV |  | 4.0000111931499998 | 1.9618340272058568e-07 | 1517951797.34 | 1517951812.785 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
| 20 | 4.0 | 2.0 | 1.0 | 10 | 3.0 | 20.0 | 10.0 | 0.0 | 49.9661079E+00 | DC2_2V |  | 3.0000051748500001 | 1.3789861312014234e-07 | 1517951827.632 | 1517951843.069 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
| 20 | 3.0 | 2.0 | 0.0 | 10 | 3.0 | 20.0 | 10.0 | 0.0 | 50.0213582E+00 | DC220MV |  | 3.0000081005999997 | 2.3905321583055427e-07 | 1517951857.932 | 1517951873.378 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
| 20 | 3.0 | 2.0 | 1.0 | 10 | 2.0 | 20.0 | 10.0 | 0.0 | 49.9761388E+00 | DC2_2V |  | 2.0000023835 | 1.8312058866593045e-07 | 1517951888.237 | 1517951903.653 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 20 | 2.0 | 2.0 | 0.0 | 10 | 2.0 | 20.0 | 10.0 | 0.0 | 49.9751397E+00 | DC220MV |  | 2.0000039065000004 | 1.6411413712892771e-07 | 1517951918.512 | 1517951933.958 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
| 20 | 2.0 | 2.0 | 1.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9698248E+00 | DC2_2V |  | 0.99999731559499982 | 9.1034073146145761e-08 | 1517951948.809 | 1517951964.249 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 1.0 | 2.0 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9740418E+00 | DC220MV |  | 1.0000018379999998 | 1.8917293674774295e-07 | 1517951979.097 | 1517951994.542 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | 1.0 | 2.0 | 1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9994717E+00 | DC2_2V |  | -4.1744993810500001e-06 | 1.0877034349103284e-07 | 1517952009.392 | 1517952024.923 | -4.2684e-06 | -4.3449e-06 | -4.2371e-06 | -4.1593e-06 | -4.2783e-06 | -4.294e-06 | -4.3405e-06 | -4.0987e-06 | -4.2108e-06 | -4.136e-06 | -4.1933e-06 | -4.029e-06 | -3.8564e-06 | -4.0935e-06 | -4.1259e-06 | -4.1869e-06 | -4.1831e-06 | -4.1527e-06 | -4.1506e-06 | -4.1506e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9723973E+00 | DC220MV |  | -7.3112539893850016e-08 | 5.0246697148077598e-08 | 1517952039.768 | 1517952055.318 | -1.0237e-07 | -6.089e-08 | -1.1178e-07 | -9.0894e-08 | -1.7649e-09 | -1.2016e-07 | -2.9269e-08 | -8.4864e-08 | -7.3392e-08 | -7.3686e-08 | -8.9718e-09 | -7.5598e-08 | -2.618e-08 | -7.3098e-08 | -2.2856e-07 | -1.1972e-07 | -5.7802e-08 | -3.2798e-08 | -1.1766e-08 | -7.8687e-08
| 20 | 1.0 | 2.0 | 1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9825297E+00 | DC2_2V |  | -4.2315525353499997e-06 | 6.2839086902305942e-08 | 1517952070.153 | 1517952085.704 | -4.1383e-06 | -4.2118e-06 | -4.1928e-06 | -4.3456e-06 | -4.2222e-06 | -4.2448e-06 | -4.2373e-06 | -4.1703e-06 | -4.1924e-06 | -4.1101e-06 | -4.3183e-06 | -4.1868e-06 | -4.3161e-06 | -4.262e-06 | -4.3063e-06 | -4.2489e-06 | -4.1903e-06 | -4.1771e-06 | -4.2466e-06 | -4.313e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0076049E+00 | DC220MV |  | -8.5577678448949998e-08 | 4.3665397493866899e-08 | 1517952100.558 | 1517952116.102 | -1.1577e-07 | -8.7085e-08 | -1.5902e-07 | -2.4419e-08 | -8.4585e-08 | -8.2231e-08 | -5.8989e-08 | -8.4438e-08 | -1.1548e-07 | -4.2513e-08 | -1.1857e-07 | -1.7711e-07 | -2.9127e-08 | -4.0748e-08 | -9.3411e-08 | -1.1386e-07 | -1.1401e-07 | 5.0015e-09 | -8.1643e-08 | -9.3558e-08
| 20 | 1.0 | 2.0 | 0.0 | 10 | 1.0 | 20.0 | 10.0 | 0.0 | 49.9884178E+00 | DC220MV |  | 1.00000101995 | 5.5958444411825943e-08 | 1517952130.946 | 1517952146.36 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9823364E+00 | DC220MV |  | -1.00000228665 | 4.8187420557126254e-08 | 1517952161.218 | 1517952176.654 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0374998E+00 | DC220MV |  | -7.3494937901000004e-08 | 6.3116546866822181e-08 | 1517952191.501 | 1517952207.045 | -1.4812e-07 | -1.946e-07 | -7.3399e-08 | -1.0385e-07 | -7.6488e-09 | -1.0473e-07 | 1.1738e-07 | 7.6488e-09 | -7.0899e-08 | -1.3503e-07 | -5.8249e-08 | -9.6346e-08 | -9.0462e-08 | -3.7509e-08 | -9.561e-08 | -1.1208e-07 | -9.0609e-08 | -4.1627e-08 | -4.8541e-08 | -8.5608e-08
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9668051E+00 | DC2_2V |  | 3.4568057901999999e-06 | 9.2940187484444772e-08 | 1517952221.903 | 1517952237.411 | 3.4387e-06 | 3.5369e-06 | 3.4136e-06 | 3.3379e-06 | 3.5986e-06 | 3.5309e-06 | 3.1739e-06 | 3.5109e-06 | 3.4159e-06 | 3.5153e-06 | 3.3695e-06 | 3.4183e-06 | 3.4134e-06 | 3.5291e-06 | 3.5274e-06 | 3.5238e-06 | 3.4336e-06 | 3.4755e-06 | 3.4231e-06 | 3.5497e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0176241E+00 | DC220MV |  | -4.0409474836999994e-08 | 4.1771932917377065e-08 | 1517952252.26 | 1517952267.824 | -8.6597e-08 | -7.2777e-08 | -4.6018e-08 | -6.6896e-08 | -5.2634e-08 | -2.0583e-08 | -6.4396e-08 | -1.0409e-07 | -5.4546e-08 | -4.5871e-08 | -6.5426e-08 | -1.8966e-08 | -5.8221e-08 | -1.8966e-08 | 8.8508e-08 | -5.1458e-08 | 1.1321e-08 | -6.4543e-08 | -3.6462e-08 | 2.0436e-08
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9600555E+00 | DC2_2V |  | 3.5569592235500005e-06 | 7.181013786577813e-08 | 1517952282.672 | 1517952298.194 | 3.7029e-06 | 3.5637e-06 | 3.6084e-06 | 3.44e-06 | 3.5413e-06 | 3.5135e-06 | 3.3988e-06 | 3.5812e-06 | 3.5522e-06 | 3.5881e-06 | 3.5034e-06 | 3.5846e-06 | 3.5793e-06 | 3.6249e-06 | 3.5337e-06 | 3.6149e-06 | 3.5452e-06 | 3.6786e-06 | 3.5118e-06 | 3.4728e-06
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 50.0006887E+00 | DC220MV |  | -1.0000022567500002 | 6.8654114960658292e-08 | 1517952313.057 | 1517952328.493 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -2.0 | 2.0 | -1.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 50.0341271E+00 | DC2_2V |  | -0.99999878270499987 | 6.1739091961485833e-08 | 1517952343.353 | 1517952358.793 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -2.0 | 2.0 | 0.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 49.9571923E+00 | DC220MV |  | -2.0000055182999992 | 1.8776983252485467e-07 | 1517952373.614 | 1517952389.028 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -3.0 | 2.0 | -1.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 49.9992642E+00 | DC2_2V |  | -2.0000035432000001 | 2.2841203997569178e-07 | 1517952403.875 | 1517952419.319 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -3.0 | 2.0 | 0.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 50.0002075E+00 | DC220MV |  | -3.0000093802499999 | 1.7205634976232802e-07 | 1517952434.168 | 1517952449.594 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -4.0 | 2.0 | -1.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 50.0340420E+00 | DC2_2V |  | -3.0000060209499999 | 2.4612283014008133e-07 | 1517952464.443 | 1517952479.868 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -4.0 | 2.0 | 0.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9389238E+00 | DC220MV |  | -4.0000120905500003 | 2.8751721260697806e-07 | 1517952494.727 | 1517952510.143 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -5.0 | 2.0 | -1.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9544707E+00 | DC2_2V |  | -4.0000087056 | 2.3119610728104944e-07 | 1517952524.986 | 1517952540.435 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -5.0 | 2.0 | 0.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 49.9943638E+00 | DC220MV |  | -5.0000151378000002 | 1.5333512316170919e-07 | 1517952555.299 | 1517952570.757 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -6.0 | 2.0 | -1.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 50.0059111E+00 | DC2_2V |  | -5.0000120572000002 | 2.1286934954682329e-07 | 1517952585.65 | 1517952601.094 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -6.0 | 2.0 | 0.0 | 10 | -6.0 | 20.0 | 10.0 | 0.0 | 50.0103134E+00 | DC220MV |  | -6.0000184518499999 | 2.286548218063639e-07 | 1517952615.941 | 1517952631.377 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6
| 20 | -7.0 | 2.0 | -1.0 | 10 | -6.0 | 20.0 | 10.0 | 0.0 | 49.9625799E+00 | DC2_2V |  | -6.0000150767499987 | 2.3373465194823398e-07 | 1517952646.225 | 1517952661.656 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6
| 20 | -7.0 | 2.0 | 0.0 | 10 | -7.0 | 20.0 | 10.0 | 0.0 | 50.0407076E+00 | DC220MV |  | -7.0000213871499994 | 3.054145176119334e-07 | 1517952676.495 | 1517952691.943 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7
| 20 | -8.0 | 2.0 | -1.0 | 10 | -7.0 | 20.0 | 10.0 | 0.0 | 50.0485754E+00 | DC2_2V |  | -7.0000181981500003 | 3.3825305246367151e-07 | 1517952706.801 | 1517952722.223 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7
| 20 | -8.0 | 2.0 | 0.0 | 10 | -8.0 | 20.0 | 10.0 | 0.0 | 50.0183747E+00 | DC220MV |  | -8.0000248957499984 | 3.3811431135592528e-07 | 1517952737.063 | 1517952752.47 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8
| 20 | -9.0 | 2.0 | -1.0 | 10 | -8.0 | 20.0 | 10.0 | 0.0 | 49.9772745E+00 | DC2_2V |  | -8.0000213047000006 | 1.5267845298952118e-07 | 1517952767.29 | 1517952782.714 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8
| 20 | -9.0 | 2.0 | 0.0 | 10 | -9.0 | 20.0 | 10.0 | 0.0 | 49.9924257E+00 | DC220MV |  | -9.0000278492499994 | 3.2379466863245828e-07 | 1517952797.597 | 1517952813.041 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9
| 20 | -10.0 | 2.0 | -1.0 | 10 | -9.0 | 20.0 | 10.0 | 0.0 | 50.0368148E+00 | DC2_2V |  | -9.0000242145499989 | 2.407969010241106e-07 | 1517952827.886 | 1517952843.323 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9
| 20 | -10.0 | 2.0 | 0.0 | 10 | -10.0 | 20.0 | 10.0 | 0.0 | 50.0335602E+00 | DC220MV |  | -10.000031203499999 | 3.2367074307150449e-07 | 1517952858.172 | 1517952873.572 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10
| 20 | -10.0 | 2.0 | -1.0 | 10 | -9.0 | 20.0 | 10.0 | 0.0 | 49.9858248E+00 | DC2_2V |  | -9.0000247328 | 2.6838621435220939e-07 | 1517952888.418 | 1517952903.836 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9
| 20 | -9.0 | 2.0 | 0.0 | 10 | -9.0 | 20.0 | 10.0 | 0.0 | 50.0013585E+00 | DC220MV |  | -9.0000283564499988 | 2.1476998752748057e-07 | 1517952918.694 | 1517952934.135 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9 | -9
| 20 | -9.0 | 2.0 | -1.0 | 10 | -8.0 | 20.0 | 10.0 | 0.0 | 50.0949725E+00 | DC2_2V |  | -8.0000217899999999 | 2.1249776475159356e-07 | 1517952948.985 | 1517952964.426 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8
| 20 | -8.0 | 2.0 | 0.0 | 10 | -8.0 | 20.0 | 10.0 | 0.0 | 50.0432544E+00 | DC220MV |  | -8.0000253951499989 | 2.2927282337007035e-07 | 1517952979.245 | 1517952994.633 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8 | -8
| 20 | -8.0 | 2.0 | -1.0 | 10 | -7.0 | 20.0 | 10.0 | 0.0 | 50.0366684E+00 | DC2_2V |  | -7.0000187796500013 | 2.8721286097512962e-07 | 1517953009.478 | 1517953024.899 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7
| 20 | -7.0 | 2.0 | 0.0 | 10 | -7.0 | 20.0 | 10.0 | 0.0 | 49.9671866E+00 | DC220MV |  | -7.0000221753499998 | 2.0941257716050925e-07 | 1517953039.737 | 1517953055.144 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7 | -7
| 20 | -7.0 | 2.0 | -1.0 | 10 | -6.0 | 20.0 | 10.0 | 0.0 | 50.0657846E+00 | DC2_2V |  | -6.00001523195 | 2.5883092459467077e-07 | 1517953070.004 | 1517953085.462 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6
| 20 | -6.0 | 2.0 | 0.0 | 10 | -6.0 | 20.0 | 10.0 | 0.0 | 49.9856221E+00 | DC220MV |  | -6.00001858135 | 2.7594624022255648e-07 | 1517953100.315 | 1517953115.719 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6 | -6
| 20 | -6.0 | 2.0 | -1.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 50.0600626E+00 | DC2_2V |  | -5.0000119706500001 | 2.0997077772763082e-07 | 1517953130.571 | 1517953146.014 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -5.0 | 2.0 | 0.0 | 10 | -5.0 | 20.0 | 10.0 | 0.0 | 49.9996604E+00 | DC220MV |  | -5.0000152846500008 | 2.7713034384561863e-07 | 1517953160.865 | 1517953176.265 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5 | -5
| 20 | -5.0 | 2.0 | -1.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9791172E+00 | DC2_2V |  | -4.000008833249999 | 1.9729010997137753e-07 | 1517953191.09 | 1517953206.528 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -4.0 | 2.0 | 0.0 | 10 | -4.0 | 20.0 | 10.0 | 0.0 | 49.9709414E+00 | DC220MV |  | -4.0000121665999995 | 2.0612991043262152e-07 | 1517953221.391 | 1517953236.833 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4 | -4
| 20 | -4.0 | 2.0 | -1.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 49.9867252E+00 | DC2_2V |  | -3.00000591485 | 1.819956798703311e-07 | 1517953251.663 | 1517953267.107 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -3.0 | 2.0 | 0.0 | 10 | -3.0 | 20.0 | 10.0 | 0.0 | 49.9611717E+00 | DC220MV |  | -3.0000089479499996 | 1.6742176527722279e-07 | 1517953281.94 | 1517953297.371 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
| 20 | -3.0 | 2.0 | -1.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 49.9673515E+00 | DC2_2V |  | -2.0000027045 | 2.548549587601798e-07 | 1517953312.197 | 1517953327.628 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -2.0 | 2.0 | 0.0 | 10 | -2.0 | 20.0 | 10.0 | 0.0 | 50.0139803E+00 | DC220MV |  | -2.0000052139000002 | 1.3849725628108434e-07 | 1517953342.466 | 1517953357.909 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
| 20 | -2.0 | 2.0 | -1.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 49.9940998E+00 | DC2_2V |  | -0.99999890432000016 | 6.1861665029042315e-08 | 1517953372.761 | 1517953388.196 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 50.0469875E+00 | DC220MV |  | -1.0000026514 | 2.0941571098006633e-07 | 1517953403.042 | 1517953418.485 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9911242E+00 | DC2_2V |  | 3.4170283249500007e-06 | 6.6262063430627916e-08 | 1517953433.334 | 1517953448.848 | 3.3677e-06 | 3.4792e-06 | 3.4898e-06 | 3.4603e-06 | 3.4094e-06 | 3.328e-06 | 3.368e-06 | 3.3814e-06 | 3.4279e-06 | 3.3206e-06 | 3.378e-06 | 3.4132e-06 | 3.371e-06 | 3.4778e-06 | 3.3826e-06 | 3.3028e-06 | 3.5511e-06 | 3.5158e-06 | 3.4758e-06 | 3.44e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0603983E+00 | DC220MV |  | 6.3836710405899986e-08 | 4.2118664929435562e-08 | 1517953463.677 | 1517953479.23 | 8.8537e-08 | 6.8094e-08 | 3.8974e-08 | 1.9266e-08 | 6.8094e-08 | 7.3095e-08 | 2.162e-08 | 3.8974e-08 | 9.192e-08 | 7.133e-08 | 1.3295e-07 | 6.3829e-08 | -2.8973e-08 | 1.0192e-07 | 7.3977e-08 | 1.2501e-07 | 6.0153e-08 | -7.0595e-09 | 4.2504e-08 | 1.3251e-07
| 20 | -1.0 | 2.0 | -1.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9964436E+00 | DC2_2V |  | 3.6228821413000004e-06 | 9.2399400391882004e-08 | 1517953494.067 | 1517953509.567 | 3.775e-06 | 3.6195e-06 | 3.6075e-06 | 3.5759e-06 | 3.7223e-06 | 3.6612e-06 | 3.5849e-06 | 3.3677e-06 | 3.6746e-06 | 3.7148e-06 | 3.5704e-06 | 3.5539e-06 | 3.6541e-06 | 3.4718e-06 | 3.6821e-06 | 3.6996e-06 | 3.725e-06 | 3.6216e-06 | 3.6231e-06 | 3.5524e-06
| 20 | 0.0 | 2.0 | 0.0 | 10 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9722890E+00 | DC220MV |  | 5.1984582989750003e-08 | 5.3183752199464191e-08 | 1517953524.425 | 1517953539.982 | 1.1176e-08 | 1.4147e-07 | 4.7793e-08 | 4.9411e-08 | 8.5881e-08 | 6.6617e-08 | 1.2823e-07 | 3.1764e-08 | 6.5734e-08 | 6.3823e-08 | 4.0294e-08 | 8.9558e-08 | 7.7058e-08 | 3.6029e-08 | 2.9558e-08 | -1.1765e-09 | -1.275e-07 | 6.8234e-08 | 5.7499e-08 | 7.8234e-08
| 20 | -1.0 | 2.0 | 0.0 | 10 | -1.0 | 20.0 | 10.0 | 0.0 | 50.0249700E+00 | DC220MV |  | -1.0000022362499998 | 5.2708514485110661e-08 | 1517953554.813 | 1517953570.231 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1 | -1
| 100 | 5.0 | 10 | 0.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 49.9364817E+00 | DC11V |  | 4.9999857663 | 1.0614904371225877e-05 | 1517953585.789 | 1517953600.902 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0200552E+00 | DC11V |  | -1.1101065788705e-05 | 9.7288844040010689e-06 | 1517953615.74 | 1517953630.924 | 2.9407e-07 | -2.0438e-05 | -1.2645e-05 | -1.7791e-05 | -1.3527e-05 | 5.7343e-06 | -1.1763e-05 | -1.485e-05 | -1.0733e-05 | -2.7936e-05 | -7.7928e-06 | -1.1175e-05 | -1.4703e-05 | 1.1763e-05 | -9.4102e-06 | -8.528e-06 | -2.5731e-05 | -1.4409e-05 | -2.0438e-05 | 2.0585e-06
| 100 | 5.0 | 10 | 5.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0016604E+00 | DC11V |  | -3.4159329889499997e-05 | 7.3552541954810305e-06 | 1517953645.753 | 1517953660.906 | -3.5035e-05 | -3.8863e-05 | -3.4299e-05 | -1.7665e-05 | -3.371e-05 | -3.7243e-05 | -4.5192e-05 | -2.4289e-05 | -2.8264e-05 | -3.2533e-05 | -4.3426e-05 | -4.4015e-05 | -2.8558e-05 | -2.6497e-05 | -4.5487e-05 | -3.268e-05 | -3.268e-05 | -3.4299e-05 | -2.6644e-05 | -4.1807e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9743151E+00 | DC11V |  | -2.0988235102250001e-05 | 8.3019269795900514e-06 | 1517953675.75 | 1517953690.916 | -3.2957e-05 | -1.7509e-05 | -1.898e-05 | -2.2364e-05 | -1.3389e-05 | -2.2952e-05 | -2.7808e-05 | -1.7067e-05 | -2.6631e-05 | -1.589e-05 | -1.2359e-05 | -3.487e-05 | -2.7808e-05 | -2.6484e-05 | -1.6626e-05 | -2.5454e-05 | -1.0593e-05 | -2.3247e-05 | 1.177e-06 | -2.7955e-05
| 100 | 5.0 | 10 | 5.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0447665E+00 | DC11V |  | -0.00014043681030050001 | 4.3967003917897915e-05 | 1517953705.756 | 1517953720.922 | -9.6199e-05 | -0.000144 | -0.00010194 | -0.00012841 | -0.00019652 | -8.4578e-05 | -0.00013547 | -0.0001696 | -0.00015121 | -0.00012753 | -0.00013591 | -0.00014474 | -7.0605e-05 | -0.00013135 | -0.0001824 | -0.00010576 | -0.00027977 | -0.00014047 | -0.00015268 | -0.00012959
| 100 | 5.0 | 10 | 0.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 49.9603193E+00 | DC11V |  | 4.9997793409 | 0.00012052581776043576 | 1517953735.77 | 1517953750.873 | 4.9999 | 4.9997 | 4.9998 | 4.9997 | 4.9998 | 4.9998 | 4.9998 | 4.9997 | 4.9997 | 4.9996 | 4.9997 | 4.9997 | 4.9996 | 4.9998 | 4.9997 | 4.9998 | 5 | 5 | 5 | 5
| 100 | 10.0 | 10 | 5.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0393800E+00 | DC11V |  | 4.9999481575000013 | 2.1555375159081263e-05 | 1517953765.698 | 1517953780.848 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 5 | 4.9999 | 5 | 5 | 5 | 5 | 5 | 5 | 4.9999 | 5 | 5 | 4.9999 | 5
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9565237E+00 | DC11V |  | 9.9999542774999988 | 3.0903672836057091e-05 | 1517953795.683 | 1517953810.789 | 9.9999 | 10 | 10 | 10 | 10 | 10 | 10 | 9.9999 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 9.9998 | 10 | 10 | 10 | 10
| 100 | 15.0 | 10 | 5.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 50.0140511E+00 | DC11V |  | 9.9998950935000011 | 5.4226516646914539e-05 | 1517953825.612 | 1517953840.751 | 9.9999 | 9.9999 | 9.9999 | 10 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9998 | 9.9998 | 9.9998 | 9.9999 | 9.9999 | 9.9998 | 9.9998 | 9.9999 | 9.9999 | 10
| 100 | 15.0 | 10 | 0.0 | 100 | 15.0 | 20.0 | 10.0 | 0.0 | 49.9946892E+00 | DC11V |  | 14.999955931999997 | 8.4848780779886177e-06 | 1517953855.579 | 1517953870.691 | 15 | 15 | 15 | 14.9999 | 14.9999 | 15 | 14.9999 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15
| 100 | 20.0 | 10 | 5.0 | 100 | 15.0 | 20.0 | 10.0 | 0.0 | 49.9774818E+00 | DC11V |  | 14.999942616000002 | 7.6267551423961121e-06 | 1517953885.541 | 1517953900.657 | 14.9999 | 15 | 14.9999 | 14.9999 | 15 | 14.9999 | 15 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 15 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 15
| 100 | 20.0 | 10 | 0.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 50.0040664E+00 | DC11V |  | 19.9999547205 | 7.3937321263215743e-06 | 1517953915.483 | 1517953930.604 | 20 | 20 | 19.9999 | 20 | 19.9999 | 19.9999 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 19.9999 | 19.9999 | 19.9999 | 20 | 19.9999
| 100 | 25.0 | 10 | 5.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9855231E+00 | DC11V |  | 19.999929102499998 | 1.7884988642868916e-05 | 1517953945.461 | 1517953960.566 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 20 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 25.0 | 10 | 0.0 | 100 | 25.0 | 20.0 | 10.0 | 0.0 | 50.0683291E+00 | DC11V |  | 24.999941704000001 | 9.9391837692757562e-06 | 1517953975.417 | 1517953990.54 | 24.9999 | 25 | 24.9999 | 24.9999 | 25 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 25 | 25 | 24.9999 | 24.9999
| 100 | 30.0 | 10 | 5.0 | 100 | 25.0 | 20.0 | 10.0 | 0.0 | 50.0050241E+00 | DC11V |  | 24.999922012499997 | 2.9415257924239914e-05 | 1517954005.37 | 1517954021.281 | 24.9998 | 25 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999
| 100 | 30.0 | 10 | 0.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9812994E+00 | DC11V |  | 29.999942272000006 | 1.2540105900546096e-05 | 1517954036.135 | 1517954051.241 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 30 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 30 | 29.9999 | 29.9999 | 29.9999 | 30 | 29.9999 | 29.9999 | 30 | 29.9999
| 100 | 35.0 | 10 | 5.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9936376E+00 | DC11V |  | 29.999904536499997 | 4.5773275093279398e-05 | 1517954066.072 | 1517954081.194 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9998 | 29.9998 | 29.9998 | 29.9999 | 29.9999 | 29.9999
| 100 | 35.0 | 10 | 0.0 | 100 | 35.0 | 20.0 | 10.0 | 0.0 | 49.9911997E+00 | DC11V |  | 34.999938380999993 | 7.5814410903951846e-06 | 1517954096.04 | 1517954111.151 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 35 | 34.9999 | 34.9999 | 35
| 100 | 40.0 | 10 | 5.0 | 100 | 35.0 | 20.0 | 10.0 | 0.0 | 50.0354780E+00 | DC11V |  | 34.999921062500007 | 9.5692151584665448e-06 | 1517954125.994 | 1517954141.104 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999
| 100 | 40.0 | 10 | 0.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9746591E+00 | DC11V |  | 39.999928728499995 | 8.7561191606936471e-06 | 1517954155.954 | 1517954171.046 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 45.0 | 10 | 5.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 50.0092895E+00 | DC11V |  | 39.999895579499992 | 4.3609968524893703e-05 | 1517954185.894 | 1517954201.011 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9997 | 39.9998 | 39.9999 | 39.9999
| 100 | 45.0 | 10 | 0.0 | 100 | 45.0 | 20.0 | 10.0 | 0.0 | 49.9903368E+00 | DC11V |  | 44.999928765999996 | 1.0420285216391099e-05 | 1517954215.831 | 1517954230.936 | 44.9999 | 44.9999 | 45 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999
| 100 | 50.0 | 10 | 5.0 | 100 | 45.0 | 20.0 | 10.0 | 0.0 | 49.9935433E+00 | DC11V |  | 44.999881629000001 | 7.7970698335716972e-05 | 1517954245.758 | 1517954260.88 | 44.9996 | 44.9998 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999
| 100 | 50.0 | 10 | 0.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 50.0258294E+00 | DC11V |  | 49.999930971000005 | 1.2664632209112555e-05 | 1517954275.712 | 1517954290.828 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 50 | 50 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 50.0 | 10 | 5.0 | 100 | 45.0 | 20.0 | 10.0 | 0.0 | 49.9970238E+00 | DC11V |  | 44.999910624000002 | 1.2456631727248335e-05 | 1517954305.664 | 1517954320.752 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999
| 100 | 45.0 | 10 | 0.0 | 100 | 45.0 | 20.0 | 10.0 | 0.0 | 49.9904877E+00 | DC11V |  | 44.999921755999999 | 8.390701639150359e-06 | 1517954335.6 | 1517954350.719 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999 | 44.9999
| 100 | 45.0 | 10 | 5.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9427543E+00 | DC11V |  | 39.999876911999998 | 8.047870964453706e-05 | 1517954365.552 | 1517954380.696 | 39.9997 | 39.9999 | 39.9999 | 39.9996 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9998 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 40.0 | 10 | 0.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 50.0252722E+00 | DC11V |  | 39.999892829000004 | 7.849911865108697e-05 | 1517954395.535 | 1517954410.675 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9998 | 39.9997 | 39.9997
| 100 | 40.0 | 10 | 5.0 | 100 | 35.0 | 20.0 | 10.0 | 0.0 | 50.0304856E+00 | DC11V |  | 34.998963349999997 | 0.00027154553575823312 | 1517954425.522 | 1517954440.62 | 34.9995 | 34.9988 | 34.9986 | 34.9987 | 34.9988 | 34.9988 | 34.999 | 34.9988 | 34.9987 | 34.9987 | 34.9988 | 34.9995 | 34.9991 | 34.9988 | 34.999 | 34.9989 | 34.9996 | 34.9992 | 34.999 | 34.9991
| 100 | 35.0 | 10 | 0.0 | 100 | 35.0 | 20.0 | 10.0 | 0.0 | 49.8945906E+00 | DC11V |  | 34.999927992499991 | 1.1936585725177348e-05 | 1517954455.462 | 1517954470.559 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999 | 34.9999
| 100 | 35.0 | 10 | 5.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9423684E+00 | DC11V |  | 29.999919498999997 | 8.7651776365356672e-06 | 1517954485.399 | 1517954500.565 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 30.0 | 10 | 0.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9777976E+00 | DC11V |  | 29.999928696000001 | 7.8559094956505906e-06 | 1517954515.391 | 1517954530.54 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 30.0 | 10 | 5.0 | 100 | 25.0 | 20.0 | 10.0 | 0.0 | 49.9446133E+00 | DC11V |  | 24.999917951 | 8.9316862351649585e-06 | 1517954545.393 | 1517954560.519 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999
| 100 | 25.0 | 10 | 0.0 | 100 | 25.0 | 20.0 | 10.0 | 0.0 | 50.0326156E+00 | DC11V |  | 24.999930439000003 | 8.8614665261842988e-06 | 1517954575.397 | 1517954590.529 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999 | 24.9999
| 100 | 25.0 | 10 | 5.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 50.0091385E+00 | DC11V |  | 19.999916133500001 | 1.0911862478392988e-05 | 1517954605.378 | 1517954620.469 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 20.0 | 10 | 0.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 50.0208247E+00 | DC11V |  | 19.999909327499999 | 6.251250137937961e-05 | 1517954635.297 | 1517954650.412 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9997 | 19.9999 | 19.9997 | 19.9998
| 100 | 20.0 | 10 | 5.0 | 100 | 15.0 | 20.0 | 10.0 | 0.0 | 49.9643038E+00 | DC11V |  | 14.999861339999995 | 3.2978053611542812e-05 | 1517954665.237 | 1517954680.343 | 14.9999 | 14.9999 | 14.9998 | 14.9998 | 14.9999 | 14.9998 | 14.9998 | 14.9998 | 14.9998 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9998 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999
| 100 | 15.0 | 10 | 0.0 | 100 | 15.0 | 20.0 | 10.0 | 0.0 | 49.9969247E+00 | DC11V |  | 14.999935026000003 | 8.6262954970542194e-06 | 1517954695.175 | 1517954710.315 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 14.9999 | 15 | 14.9999
| 100 | 15.0 | 10 | 5.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 50.0107476E+00 | DC11V |  | 9.9999195456999992 | 1.0765669756415871e-05 | 1517954725.15 | 1517954740.271 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.8819867E+00 | DC11V |  | 9.9999426006499981 | 1.1386712757874764e-05 | 1517954755.098 | 1517954770.219 | 9.9999 | 9.9999 | 9.9999 | 10 | 10 | 10 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 10 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 10.0 | 10 | 5.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 50.0260324E+00 | DC11V |  | 4.9999282976499995 | 8.839930148339149e-06 | 1517954785.064 | 1517954800.272 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999
| 100 | 5.0 | 10 | 0.0 | 100 | 5.0 | 20.0 | 10.0 | 0.0 | 49.9335695E+00 | DC11V |  | 4.9999382951499998 | 5.512281662565377e-06 | 1517954815.121 | 1517954830.226 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999
| 100 | 5.0 | 10 | 5 | 100 | 0 | 20.0 | 10.0 | 0.0 | 49.9745507E+00 | DC11V |  | -7.0761873933000019e-05 | 1.0500332077447754e-05 | 1517954845.054 | 1517954860.247 | -7.1129e-05 | -8.0388e-05 | -6.0548e-05 | -9.0675e-05 | -4.8497e-05 | -7.2893e-05 | -8.9793e-05 | -7.5979e-05 | -7.9653e-05 | -6.5398e-05 | -6.3487e-05 | -6.2312e-05 | -7.0101e-05 | -7.7889e-05 | -6.5692e-05 | -7.157e-05 | -8.3327e-05 | -5.8197e-05 | -6.4516e-05 | -6.3193e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0254941E+00 | DC11V |  | -5.7312870348999986e-05 | 8.0450360485456539e-06 | 1517954875.079 | 1517954890.246 | -6.0469e-05 | -5.5761e-05 | -6.9296e-05 | -5.0317e-05 | -5.4289e-05 | -5.9586e-05 | -6.1646e-05 | -5.6055e-05 | -5.3407e-05 | -4.914e-05 | -7.0915e-05 | -6.7678e-05 | -4.4579e-05 | -5.782e-05 | -5.8115e-05 | -6.9738e-05 | -6.0469e-05 | -5.6496e-05 | -5.0758e-05 | -3.9724e-05
| 100 | 5.0 | 10 | 5 | 100 | 0 | 20.0 | 10.0 | 0.0 | 50.0470206E+00 | DC11V |  | -7.4040782296500002e-05 | 1.1053109139011094e-05 | 1517954905.102 | 1517954920.245 | -8.0558e-05 | -8.2472e-05 | -8.6007e-05 | -5.7141e-05 | -6.097e-05 | -6.7745e-05 | -6.745e-05 | -6.4652e-05 | -9.5874e-05 | -6.3916e-05 | -7.8938e-05 | -7.0838e-05 | -7.2458e-05 | -7.3194e-05 | -5.8909e-05 | -9.0277e-05 | -7.0249e-05 | -8.3503e-05 | -9.0425e-05 | -6.5241e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9505301E+00 | DC11V |  | -5.9302906994000002e-05 | 9.5292287351700511e-06 | 1517954935.062 | 1517954950.163 | -5.3409e-05 | -6.5802e-05 | -5.636e-05 | -3.8507e-05 | -5.7392e-05 | -6.2851e-05 | -5.7982e-05 | -8.1736e-05 | -6.6392e-05 | -5.3261e-05 | -5.6064e-05 | -6.3884e-05 | -4.4852e-05 | -5.3409e-05 | -5.5769e-05 | -5.1048e-05 | -6.6245e-05 | -6.831e-05 | -5.9163e-05 | -7.3621e-05
reading 0 is more than 0.1% from nominal. reading 1 is more than 0.1% from nominal. reading 2 is more than 0.1% from nominal. reading 3 is more than 0.1% from nominal. reading 4 is more than 0.1% from nominal. reading 5 is more than 0.1% from nominal. reading 6 is more than 0.1% from nominal. reading 7 is more than 0.1% from nominal. reading 8 is more than 0.1% from nominal. reading 9 is more than 0.1% from nominal. reading 10 is more than 0.1% from nominal. reading 11 is more than 0.1% from nominal. reading 12 is more than 0.1% from nominal. reading 13 is more than 0.1% from nominal. reading 14 is more than 0.1% from nominal. reading 15 is more than 0.1% from nominal. reading 16 is more than 0.1% from nominal. reading 17 is more than 0.1% from nominal. reading 18 is more than 0.1% from nominal. reading 19 is more than 0.1% from nominal. | 100 | 5.0 | 10 | 0 | 100 | 0 | 20.0 | 10.0 | 0.0 | 49.9752905E+00 | DC11V |  | 4.9999371228999987 | 8.6466822244481122e-06 | 1517954965.021 | 1517954980.17 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 4.9999 | 5
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9544802E+00 | DC11V |  | 9.9999345229000021 | 7.3382173782714203e-06 | 1517954994.992 | 1517955010.118 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.1502520E+00 | DC11V |  | -5.7588643529499984e-05 | 1.0013710807664705e-05 | 1517955024.949 | 1517955040.14 | -4.9387e-05 | -6.9671e-05 | -6.9083e-05 | -6.4967e-05 | -6.482e-05 | -6.0558e-05 | -5.85e-05 | -5.3208e-05 | -3.072e-05 | -4.6153e-05 | -5.3796e-05 | -5.4972e-05 | -7.0406e-05 | -4.9975e-05 | -5.2327e-05 | -6.6437e-05 | -5.1151e-05 | -6.6878e-05 | -4.9681e-05 | -6.9083e-05
| 100 | 10.0 | 10 | 10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9904782E+00 | DC11V |  | -8.5043733169000004e-05 | 7.9262717547817687e-06 | 1517955054.992 | 1517955070.902 | -8.0783e-05 | -8.8167e-05 | -8.4918e-05 | -8.4771e-05 | -8.5657e-05 | -0.0001019 | -7.842e-05 | -9.2598e-05 | -7.842e-05 | -0.00010471 | -9.1564e-05 | -8.4771e-05 | -8.0045e-05 | -7.9306e-05 | -7.842e-05 | -8.7724e-05 | -8.7281e-05 | -7.1036e-05 | -7.7682e-05 | -8.2703e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0125692E+00 | DC11V |  | -6.1597226959000001e-05 | 8.061393986346641e-06 | 1517955085.721 | 1517955100.867 | -7.407e-05 | -6.3467e-05 | -6.0375e-05 | -7.7015e-05 | -5.6694e-05 | -5.4338e-05 | -6.2142e-05 | -6.1406e-05 | -6.3173e-05 | -5.6105e-05 | -5.4927e-05 | -5.2423e-05 | -6.3026e-05 | -5.8019e-05 | -6.4351e-05 | -7.7751e-05 | -5.8608e-05 | -7.2892e-05 | -5.1098e-05 | -5.0067e-05
| 100 | 10.0 | 10 | 10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0138151E+00 | DC11V |  | -8.8702721433000001e-05 | 8.77616940013124e-06 | 1517955115.71 | 1517955130.859 | -9.9524e-05 | -0.00010041 | -0.00010217 | -7.8176e-05 | -8.4654e-05 | -8.9512e-05 | -7.8618e-05 | -8.966e-05 | -9.1279e-05 | -7.9207e-05 | -0.00010924 | -9.5549e-05 | -8.5243e-05 | -8.4801e-05 | -7.7735e-05 | -9.0396e-05 | -8.5096e-05 | -8.7451e-05 | -7.6704e-05 | -8.8629e-05
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 50.0325259E+00 | DC11V |  | 9.9999316193999981 | 9.6890961517442796e-06 | 1517955145.691 | 1517955160.816 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 20.0 | 10 | 10.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 49.9983302E+00 | DC11V |  | 9.9999113364999985 | 6.008412964425717e-06 | 1517955175.657 | 1517955190.769 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 20.0 | 10 | 0.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 50.0499224E+00 | DC11V |  | 19.999927336999995 | 8.8364054340150262e-06 | 1517955205.617 | 1517955220.724 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 30.0 | 10 | 10.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9664942E+00 | DC11V |  | 19.999903814499994 | 9.3906120540506367e-06 | 1517955235.56 | 1517955250.648 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 30.0 | 10 | 0.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 50.0396257E+00 | DC11V |  | 29.999921224000001 | 1.0064833034191389e-05 | 1517955265.521 | 1517955280.649 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 40.0 | 10 | 10.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9542824E+00 | DC11V |  | 29.999893733500006 | 9.0934301969215201e-06 | 1517955295.502 | 1517955310.589 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 40.0 | 10 | 0.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9977501E+00 | DC11V |  | 39.999913035999995 | 7.6348152569899448e-06 | 1517955325.439 | 1517955340.576 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 50.0 | 10 | 10.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9878614E+00 | DC11V |  | 39.999888767000002 | 9.8162595213315528e-06 | 1517955355.413 | 1517955370.52 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 50.0 | 10 | 0.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 50.0065999E+00 | DC11V |  | 49.999906974500007 | 1.0370889293918595e-05 | 1517955385.371 | 1517955400.496 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 60.0 | 10 | 10.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 50.0230389E+00 | DC11V |  | 49.999885257499997 | 1.0596458784281525e-05 | 1517955415.35 | 1517955430.455 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 60.0 | 10 | 0.0 | 100 | 60.0 | 20.0 | 10.0 | 0.0 | 50.0523614E+00 | DC11V |  | 59.999896801000013 | 9.7600737187778368e-06 | 1517955445.311 | 1517955460.426 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999
| 100 | 70.0 | 10 | 10.0 | 100 | 60.0 | 20.0 | 10.0 | 0.0 | 49.9798053E+00 | DC11V |  | 59.999870771999987 | 6.9750789251732496e-06 | 1517955475.268 | 1517955490.356 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999
| 100 | 70.0 | 10 | 0.0 | 100 | 70.0 | 20.0 | 10.0 | 0.0 | 50.0135320E+00 | DC11V |  | 69.999892265499994 | 9.3007206555926611e-06 | 1517955505.198 | 1517955520.333 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999
| 100 | 80.0 | 10 | 10.0 | 100 | 70.0 | 20.0 | 10.0 | 0.0 | 50.0167651E+00 | DC11V |  | 69.999862256 | 6.4564761288354673e-06 | 1517955535.166 | 1517955550.275 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9998
| 100 | 80.0 | 10 | 0.0 | 100 | 80.0 | 20.0 | 10.0 | 0.0 | 49.9848584E+00 | DC11V |  | 79.999876225000008 | 8.0505238965563614e-06 | 1517955565.101 | 1517955580.202 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999
| 100 | 90.0 | 10 | 10.0 | 100 | 80.0 | 20.0 | 10.0 | 0.0 | 50.0225715E+00 | DC11V |  | 79.999855965999998 | 7.1401858514122275e-06 | 1517955595.032 | 1517955610.155 | 79.9998 | 79.9998 | 79.9999 | 79.9999 | 79.9998 | 79.9999 | 79.9999 | 79.9998 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9998 | 79.9999 | 79.9998 | 79.9999
| 100 | 90.0 | 10 | 0.0 | 100 | 90.0 | 20.0 | 10.0 | 0.0 | 49.9767702E+00 | DC11V |  | 89.999868807500008 | 8.8517251846625375e-06 | 1517955624.998 | 1517955640.098 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9998 | 89.9999 | 89.9999 | 89.9999
| 100 | 100.0 | 10 | 10.0 | 100 | 90.0 | 20.0 | 10.0 | 0.0 | 49.9753094E+00 | DC11V |  | 89.999840752500006 | 9.7828159919928038e-06 | 1517955654.943 | 1517955670.072 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9998 | 89.9998 | 89.9998
| 100 | 100.0 | 10 | 0.0 | 100 | 100.0 | 20.0 | 10.0 | 0.0 | 49.9978963E+00 | DC11V |  | 99.999859879500008 | 1.0072622040267127e-05 | 1517955684.896 | 1517955700.024 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9998 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9999 | 99.9998 | 99.9999 | 99.9999 | 99.9998 | 99.9999
| 100 | 100.0 | 10 | 10.0 | 100 | 90.0 | 20.0 | 10.0 | 0.0 | 50.0102285E+00 | DC11V |  | 89.99984052100001 | 8.2653710748992318e-06 | 1517955714.853 | 1517955729.967 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9999 | 89.9998 | 89.9998 | 89.9998 | 89.9998 | 89.9999
| 100 | 90.0 | 10 | 0.0 | 100 | 90.0 | 20.0 | 10.0 | 0.0 | 49.9835432E+00 | DC11V |  | 89.999867609500001 | 9.3662161381135067e-06 | 1517955744.805 | 1517955759.92 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999 | 89.9999
| 100 | 90.0 | 10 | 10.0 | 100 | 80.0 | 20.0 | 10.0 | 0.0 | 50.0228595E+00 | DC11V |  | 79.999860211999987 | 9.476669562737347e-06 | 1517955774.754 | 1517955789.873 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9998 | 79.9999 | 79.9999 | 79.9999 | 79.9998 | 79.9999
| 100 | 80.0 | 10 | 0.0 | 100 | 80.0 | 20.0 | 10.0 | 0.0 | 49.9690286E+00 | DC11V |  | 79.999884593999994 | 9.3192555513210747e-06 | 1517955804.719 | 1517955819.835 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999 | 79.9999
| 100 | 80.0 | 10 | 10.0 | 100 | 70.0 | 20.0 | 10.0 | 0.0 | 50.0044202E+00 | DC11V |  | 69.99986652150001 | 9.7484872043764982e-06 | 1517955834.683 | 1517955849.804 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999
| 100 | 70.0 | 10 | 0.0 | 100 | 70.0 | 20.0 | 10.0 | 0.0 | 49.9963729E+00 | DC11V |  | 69.999893762999989 | 7.9194261799976762e-06 | 1517955864.66 | 1517955879.774 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999 | 69.9999
| 100 | 70.0 | 10 | 10.0 | 100 | 60.0 | 20.0 | 10.0 | 0.0 | 50.0221041E+00 | DC11V |  | 59.999878632000012 | 7.8525146278828298e-06 | 1517955894.643 | 1517955909.76 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999
| 100 | 60.0 | 10 | 0.0 | 100 | 60.0 | 20.0 | 10.0 | 0.0 | 49.9874843E+00 | DC11V |  | 59.999902612000007 | 6.93769313196694e-06 | 1517955924.592 | 1517955939.688 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999 | 59.9999
| 100 | 60.0 | 10 | 10.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 50.0504754E+00 | DC11V |  | 49.999884909500004 | 7.2567978305264408e-06 | 1517955954.528 | 1517955969.657 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 50.0 | 10 | 0.0 | 100 | 50.0 | 20.0 | 10.0 | 0.0 | 50.0167084E+00 | DC11V |  | 49.9999132405 | 9.8888100774450084e-06 | 1517955984.482 | 1517955999.571 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999 | 49.9999
| 100 | 50.0 | 10 | 10.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 49.9784244E+00 | DC11V |  | 39.999895188499998 | 1.070758809213362e-05 | 1517956014.407 | 1517956029.51 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 40.0 | 10 | 0.0 | 100 | 40.0 | 20.0 | 10.0 | 0.0 | 50.0018208E+00 | DC11V |  | 39.999919096499994 | 7.8487675941318298e-06 | 1517956044.337 | 1517956059.461 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999 | 39.9999
| 100 | 40.0 | 10 | 10.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9976558E+00 | DC11V |  | 29.999898389999998 | 6.1991967224390346e-06 | 1517956074.3 | 1517956089.409 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 30.0 | 10 | 0.0 | 100 | 30.0 | 20.0 | 10.0 | 0.0 | 49.9947553E+00 | DC11V |  | 29.999925532499997 | 9.6305108251429403e-06 | 1517956104.251 | 1517956119.376 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999 | 29.9999
| 100 | 30.0 | 10 | 10.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 50.0585733E+00 | DC11V |  | 19.999905714000004 | 9.5830487838356285e-06 | 1517956134.202 | 1517956149.316 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 20.0 | 10 | 0.0 | 100 | 20.0 | 20.0 | 10.0 | 0.0 | 49.9529500E+00 | DC11V |  | 19.999930913 | 8.508684445888179e-06 | 1517956164.143 | 1517956179.228 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999 | 19.9999
| 100 | 20.0 | 10 | 10.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 50.0123049E+00 | DC11V |  | 9.9999112490499993 | 9.2892777354304473e-06 | 1517956194.059 | 1517956209.203 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 50.0909434E+00 | DC11V |  | 9.9999364574000005 | 9.1249422757425072e-06 | 1517956224.061 | 1517956239.178 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 10 | 9.9999 | 9.9999 | 10 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | 10.0 | 10 | 10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9856551E+00 | DC11V |  | -8.1011577849000014e-05 | 7.7698130158447307e-06 | 1517956253.999 | 1517956269.115 | -7.6227e-05 | -7.9471e-05 | -7.2983e-05 | -7.6669e-05 | -8.3452e-05 | -7.5048e-05 | -7.2983e-05 | -9.908e-05 | -8.9644e-05 | -8.3747e-05 | -8.4336e-05 | -7.5342e-05 | -0.00010085 | -7.9323e-05 | -8.065e-05 | -7.4605e-05 | -7.4163e-05 | -7.6522e-05 | -7.8586e-05 | -8.6548e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0476208E+00 | DC11V |  | -6.0063968030499988e-05 | 7.4609862972688321e-06 | 1517956283.973 | 1517956299.131 | -6.5495e-05 | -6.991e-05 | -5.4015e-05 | -5.3721e-05 | -6.6525e-05 | -6.1963e-05 | -5.8136e-05 | -5.8136e-05 | -5.4162e-05 | -6.0049e-05 | -5.6958e-05 | -6.7114e-05 | -5.9461e-05 | -7.7269e-05 | -5.6517e-05 | -7.3001e-05 | -4.7097e-05 | -5.2249e-05 | -5.5487e-05 | -5.4015e-05
| 100 | 10.0 | 10 | 10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9744329E+00 | DC11V |  | -8.3579543364999988e-05 | 1.0163641613366985e-05 | 1517956313.951 | 1517956329.093 | -6.9515e-05 | -9.4404e-05 | -8.1739e-05 | -8.3359e-05 | -7.8351e-05 | -0.00010118 | -8.3948e-05 | -7.57e-05 | -9.3521e-05 | -8.3506e-05 | -9.1017e-05 | -6.6864e-05 | -8.4979e-05 | -6.5685e-05 | -9.5435e-05 | -7.0546e-05 | -9.1164e-05 | -7.6142e-05 | -9.5288e-05 | -8.925e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0073926E+00 | DC11V |  | -6.3167847326e-05 | 1.0311652435995787e-05 | 1517956343.925 | 1517956359.092 | -8.1213e-05 | -6.4293e-05 | -7.0767e-05 | -5.4583e-05 | -6.2822e-05 | -3.8547e-05 | -5.4877e-05 | -5.5319e-05 | -6.1645e-05 | -7.2827e-05 | -6.2234e-05 | -6.6794e-05 | -5.8261e-05 | -7.4592e-05 | -6.2969e-05 | -5.3553e-05 | -7.4298e-05 | -8.1801e-05 | -5.6496e-05 | -5.5466e-05
| 100 | 10.0 | 10 | 0.0 | 100 | 10.0 | 20.0 | 10.0 | 0.0 | 50.0252203E+00 | DC11V |  | 9.9999333354999997 | 8.197950234685584e-06 | 1517956373.943 | 1517956389.033 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999 | 9.9999
| 100 | -10.0 | 10 | 0.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 49.9666072E+00 | DC11V |  | -10.000052830500001 | 9.6912468107578859e-06 | 1517956403.883 | 1517956418.986 | -10 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10 | -10 | -10.0001 | -10 | -10 | -10 | -10.0001 | -10.0001 | -10 | -10.0001 | -10.0001 | -10 | -10 | -10
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0259474E+00 | DC11V |  | -6.4122495633499988e-05 | 9.381649418940916e-06 | 1517956433.841 | 1517956449.02 | -5.8527e-05 | -6.2351e-05 | -6.8968e-05 | -5.9851e-05 | -4.8675e-05 | -5.7498e-05 | -9.2202e-05 | -6.6615e-05 | -7.1615e-05 | -6.7791e-05 | -7.2644e-05 | -6.4997e-05 | -4.5586e-05 | -6.6762e-05 | -6.3527e-05 | -7.1909e-05 | -6.1468e-05 | -5.8821e-05 | -6.2203e-05 | -6.0439e-05
| 100 | -10.0 | 10 | -10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9457241E+00 | DC11V |  | -4.2889695018499996e-05 | 8.6694195466229334e-06 | 1517956463.842 | 1517956478.991 | -3.1803e-05 | -4.6674e-05 | -4.4171e-05 | -5.6686e-05 | -5.6391e-05 | -4.2846e-05 | -3.0478e-05 | -4.0343e-05 | -6.2428e-05 | -4.7999e-05 | -2.6944e-05 | -3.6809e-05 | -4.4465e-05 | -4.4465e-05 | -3.6956e-05 | -3.6073e-05 | -4.741e-05 | -3.9165e-05 | -4.5201e-05 | -4.049e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0851069E+00 | DC11V |  | -6.1688236676499996e-05 | 9.865177332890996e-06 | 1517956493.825 | 1517956508.999 | -6.3689e-05 | -6.6483e-05 | -4.4861e-05 | -6.4865e-05 | -6.2365e-05 | -6.6189e-05 | -7.9868e-05 | -6.8248e-05 | -8.3251e-05 | -6.1041e-05 | -5.8835e-05 | -4.7215e-05 | -6.8984e-05 | -6.4571e-05 | -6.5159e-05 | -5.501e-05 | -5.7511e-05 | -4.2802e-05 | -5.707e-05 | -5.5746e-05
| 100 | -10.0 | 10 | -10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9541694E+00 | DC11V |  | -0.000118419449331 | 0.0002529292835740188 | 1517956523.821 | 1517956538.937 | -4.1429e-05 | -3.9512e-05 | -4.6884e-05 | -2.7275e-05 | -4.3493e-05 | -3.8628e-05 | -5.0865e-05 | -3.8333e-05 | -4.2314e-05 | -4.9538e-05 | -3.3173e-05 | -3.4795e-05 | -3.3468e-05 | -0.001072 | -0.00061996 | -2.8602e-05 | -2.5359e-05 | -2.7865e-05 | -4.0545e-05 | -3.4352e-05
| 100 | -10.0 | 10 | 0.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 49.9993255E+00 | DC11V |  | -10.000145701499999 | 0.00021519796949031646 | 1517956553.8 | 1517956568.922 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0009 | -10.0007 | -10.0003 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001
| 100 | -20.0 | 10 | -10.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 50.0598499E+00 | DC11V |  | -10.0000281245 | 6.3404956232040973e-06 | 1517956583.768 | 1517956598.867 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10 | -10
| 100 | -20.0 | 10 | 0.0 | 100 | -20.0 | 20.0 | 10.0 | 0.0 | 49.9707059E+00 | DC11V |  | -20.000110054499999 | 0.00014980261711549796 | 1517956613.715 | 1517956629.62 | -20 | -20.0001 | -20.0001 | -20 | -20 | -20.0005 | -20.0005 | -20.0004 | -20.0001 | -20.0001 | -20 | -20.0001 | -20.0001 | -20 | -20 | -20 | -20.0001 | -20.0001 | -20 | -20
| 100 | -30.0 | 10 | -10.0 | 100 | -20.0 | 20.0 | 10.0 | 0.0 | 50.0165574E+00 | DC11V |  | -20.000010648500002 | 6.5377551764681402e-06 | 1517956644.467 | 1517956659.577 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20
| 100 | -30.0 | 10 | 0.0 | 100 | -30.0 | 20.0 | 10.0 | 0.0 | 49.9718366E+00 | DC11V |  | -30.000036258499996 | 1.0242508128157109e-05 | 1517956674.429 | 1517956689.516 | -30 | -30 | -30.0001 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30.0001 | -30 | -30 | -30
| 100 | -40.0 | 10 | -10.0 | 100 | -30.0 | 20.0 | 10.0 | 0.0 | 50.0429803E+00 | DC11V |  | -30.000008196 | 9.7248513615250799e-06 | 1517956704.362 | 1517956719.484 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30
| 100 | -40.0 | 10 | 0.0 | 100 | -40.0 | 20.0 | 10.0 | 0.0 | 49.9814173E+00 | DC11V |  | -40.000023991500001 | 9.7151403872036246e-06 | 1517956734.321 | 1517956749.41 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40
| 100 | -50.0 | 10 | -10.0 | 100 | -40.0 | 20.0 | 10.0 | 0.0 | 50.0337208E+00 | DC11V |  | -39.999997931999999 | 8.4571978820724127e-06 | 1517956764.243 | 1517956779.36 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40
| 100 | -50.0 | 10 | 0.0 | 100 | -50.0 | 20.0 | 10.0 | 0.0 | 49.9239272E+00 | DC11V |  | -50.000020964500003 | 7.722212426075155e-06 | 1517956794.199 | 1517956809.284 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50
| 100 | -60.0 | 10 | -10.0 | 100 | -50.0 | 20.0 | 10.0 | 0.0 | 50.0287713E+00 | DC11V |  | -49.999988481500004 | 9.0242491511894347e-06 | 1517956824.129 | 1517956839.277 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50
| 100 | -60.0 | 10 | 0.0 | 100 | -60.0 | 20.0 | 10.0 | 0.0 | 50.0257963E+00 | DC11V |  | -60.002877514999987 | 0.004378226072111446 | 1517956854.125 | 1517956869.212 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60.0009 | -60.0012 | -60.0014 | -60.0017 | -60.0018 | -60.0006 | -60.0041 | -60.0118 | -60.0115 | -60.0113 | -60.0111 | -60
| 100 | -70.0 | 10 | -10.0 | 100 | -60.0 | 20.0 | 10.0 | 0.0 | 49.9951514E+00 | DC11V |  | -60.000354533500001 | 0.00018544484102521168 | 1517956884.045 | 1517956899.144 | -60 | -60 | -60 | -60.0003 | -60.0005 | -60.0005 | -60.0005 | -60.0005 | -60.0005 | -60.0005 | -60.0004 | -60.0003 | -60.0004 | -60.0004 | -60.0003 | -60.0004 | -60.0004 | -60.0003 | -60.0005 | -60.0007
| 100 | -70.0 | 10 | 0.0 | 100 | -70.0 | 20.0 | 10.0 | 0.0 | 49.9826947E+00 | DC11V |  | -70.000119263499997 | 0.00020833918756017378 | 1517956913.998 | 1517956929.092 | -70.0006 | -70.0004 | -70.0007 | -70.0003 | -70.0002 | -70.0002 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70
| 100 | -80.0 | 10 | -10.0 | 100 | -70.0 | 20.0 | 10.0 | 0.0 | 49.9957362E+00 | DC11V |  | -69.999969684499987 | 1.0476098738237914e-05 | 1517956943.914 | 1517956959.037 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70
| 100 | -80.0 | 10 | 0.0 | 100 | -80.0 | 20.0 | 10.0 | 0.0 | 50.0126541E+00 | DC11V |  | -79.999995608500001 | 9.3502290217252144e-06 | 1517956973.87 | 1517956988.984 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80
| 100 | -90.0 | 10 | -10.0 | 100 | -80.0 | 20.0 | 10.0 | 0.0 | 49.9467831E+00 | DC11V |  | -80.000020911500002 | 0.00023477575991403887 | 1517957003.804 | 1517957018.917 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -79.9999 | -80 | -80 | -80.001
| 100 | -90.0 | 10 | 0.0 | 100 | -90.0 | 20.0 | 10.0 | 0.0 | 49.9875691E+00 | DC11V |  | -89.999991629999982 | 1.0536305329736216e-05 | 1517957033.743 | 1517957048.871 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90
| 100 | -100.0 | 10 | -10.0 | 100 | -90.0 | 20.0 | 10.0 | 0.0 | 49.9975662E+00 | DC11V |  | -89.999974925999993 | 9.2104118768706605e-05 | 1517957063.696 | 1517957078.816 | -89.9999 | -90 | -89.9999 | -90 | -90 | -89.9999 | -89.9999 | -90 | -89.9999 | -90 | -90 | -90 | -89.9999 | -90 | -90 | -90 | -90 | -90 | -89.9999 | -90.0004
| 100 | -100.0 | 10 | 0.0 | 100 | -100.0 | 20.0 | 10.0 | 0.0 | 50.0395549E+00 | DC11V |  | -100.00022391 | 5.1415123261986733e-05 | 1517957093.639 | 1517957108.733 | -100.0002 | -100.0002 | -100.0002 | -100.0002 | -100.0002 | -100.0002 | -100.0002 | -100.0002 | -100.0002 | -100.0003 | -100.0002 | -100.0002 | -100.0001 | -100.0003 | -100.0002 | -100.0002 | -100.0002 | -100.0003 | -100.0003 | -100.0003
| 100 | -100.0 | 10 | -10.0 | 100 | -90.0 | 20.0 | 10.0 | 0.0 | 49.9673421E+00 | DC11V |  | -89.999956880499994 | 1.028002600752527e-05 | 1517957123.575 | 1517957138.678 | -90 | -90 | -90 | -89.9999 | -90 | -90 | -90 | -90 | -89.9999 | -90 | -90 | -89.9999 | -90 | -90 | -89.9999 | -90 | -90 | -89.9999 | -90 | -90
| 100 | -90.0 | 10 | 0.0 | 100 | -90.0 | 20.0 | 10.0 | 0.0 | 50.0180206E+00 | DC11V |  | -89.99998740449999 | 9.8982303862872986e-06 | 1517957153.512 | 1517957168.633 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90 | -90
| 100 | -90.0 | 10 | -10.0 | 100 | -80.0 | 20.0 | 10.0 | 0.0 | 50.0304619E+00 | DC11V |  | -79.9999689725 | 6.3703350568399551e-06 | 1517957183.421 | 1517957198.544 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80
| 100 | -80.0 | 10 | 0.0 | 100 | -80.0 | 20.0 | 10.0 | 0.0 | 49.9731182E+00 | DC11V |  | -80.000010696499999 | 2.0130155308191806e-05 | 1517957213.397 | 1517957228.487 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80 | -80.0001 | -80.0001
| 100 | -80.0 | 10 | -10.0 | 100 | -70.0 | 20.0 | 10.0 | 0.0 | 49.9751303E+00 | DC11V |  | -69.999974630000011 | 8.9430738565931272e-06 | 1517957243.315 | 1517957258.443 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70
| 100 | -70.0 | 10 | 0.0 | 100 | -70.0 | 20.0 | 10.0 | 0.0 | 50.0161986E+00 | DC11V |  | -70.000005835000024 | 7.997671222002523e-06 | 1517957273.276 | 1517957288.386 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70 | -70
| 100 | -70.0 | 10 | -10.0 | 100 | -60.0 | 20.0 | 10.0 | 0.0 | 49.9880265E+00 | DC11V |  | -60.000014475500009 | 1.7839840098298022e-05 | 1517957303.212 | 1517957318.307 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60 | -60
| 100 | -60.0 | 10 | 0.0 | 100 | -60.0 | 20.0 | 10.0 | 0.0 | 50.0155142E+00 | DC11V |  | -60.000071176500001 | 1.4665182669772081e-05 | 1517957333.16 | 1517957348.26 | -60.0001 | -60.0001 | -60.0001 | -60 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001 | -60.0001
| 100 | -60.0 | 10 | -10.0 | 100 | -50.0 | 20.0 | 10.0 | 0.0 | 49.9855183E+00 | DC11V |  | -50.000017628999998 | 1.2587759093178885e-05 | 1517957363.11 | 1517957378.211 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50
| 100 | -50.0 | 10 | 0.0 | 100 | -50.0 | 20.0 | 10.0 | 0.0 | 50.0021558E+00 | DC11V |  | -50.000035924500004 | 1.2030823112355835e-05 | 1517957393.058 | 1517957408.15 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50.0001 | -50.0001 | -50.0001 | -50 | -50 | -50 | -50
| 100 | -50.0 | 10 | -10.0 | 100 | -40.0 | 20.0 | 10.0 | 0.0 | 49.9941139E+00 | DC11V |  | -40.000020885000005 | 8.0056601845817073e-06 | 1517957422.978 | 1517957438.082 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40 | -40
| 100 | -40.0 | 10 | 0.0 | 100 | -40.0 | 20.0 | 10.0 | 0.0 | 50.0180159E+00 | DC11V |  | -40.000051885500007 | 6.7175475254119536e-06 | 1517957452.919 | 1517957468.025 | -40.0001 | -40.0001 | -40 | -40.0001 | -40.0001 | -40.0001 | -40.0001 | -40 | -40 | -40.0001 | -40.0001 | -40.0001 | -40.0001 | -40.0001 | -40.0001 | -40.0001 | -40 | -40 | -40 | -40.0001
| 100 | -40.0 | 10 | -10.0 | 100 | -30.0 | 20.0 | 10.0 | 0.0 | 49.9582236E+00 | DC11V |  | -30.000030545000005 | 1.1035277749178459e-05 | 1517957482.872 | 1517957497.964 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30 | -30
| 100 | -30.0 | 10 | 0.0 | 100 | -30.0 | 20.0 | 10.0 | 0.0 | 50.0021274E+00 | DC11V |  | -30.000061235500006 | 1.3290997131483453e-05 | 1517957512.8 | 1517957527.913 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30 | -30.0001 | -30 | -30 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30.0001 | -30.0001
| 100 | -30.0 | 10 | -10.0 | 100 | -20.0 | 20.0 | 10.0 | 0.0 | 49.9903273E+00 | DC11V |  | -20.0000370815 | 8.7829068507565718e-06 | 1517957542.756 | 1517957557.851 | -20 | -20 | -20.0001 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20 | -20.0001 | -20 | -20 | -20 | -20 | -20 | -20 | -20
| 100 | -20.0 | 10 | 0.0 | 100 | -20.0 | 20.0 | 10.0 | 0.0 | 49.9656227E+00 | DC11V |  | -20.000066514499999 | 8.6050493752930792e-06 | 1517957572.692 | 1517957587.794 | -20.0001 | -20.0001 | -20.0001 | -20 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001 | -20.0001
| 100 | -20.0 | 10 | -10.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 50.0240446E+00 | DC11V |  | -10.000048261 | 8.6516310022513019e-06 | 1517957602.617 | 1517957617.731 | -10.0001 | -10 | -10.0001 | -10.0001 | -10 | -10.0001 | -10 | -10 | -10 | -10 | -10.0001 | -10.0001 | -10 | -10 | -10.0001 | -10 | -10 | -10.0001 | -10 | -10
| 100 | -10.0 | 10 | 0.0 | 100 | -10.0 | 20.0 | 10.0 | 0.0 | 49.9760634E+00 | DC11V |  | -10.0000874405 | 1.5023493094167023e-05 | 1517957632.575 | 1517957647.668 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001 | -10.0001
| 100 | -10.0 | 10 | -10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0391816E+00 | DC11V |  | -5.7123406048000001e-05 | 1.1683740409103601e-05 | 1517957662.492 | 1517957677.67 | -3.6621e-05 | -7.7949e-05 | -4.5299e-05 | -6.3536e-05 | -4.7505e-05 | -7.1919e-05 | -7.3537e-05 | -5.1917e-05 | -5.9271e-05 | -4.9417e-05 | -5.7653e-05 | -5.3241e-05 | -5.7947e-05 | -6.5301e-05 | -5.8535e-05 | -7.0301e-05 | -3.9857e-05 | -4.7946e-05 | -7.1184e-05 | -4.3534e-05
| 100 | 0.0 | 10 | 0.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 50.0084543E+00 | DC11V |  | -8.7037371482000001e-05 | 1.0086192620014197e-05 | 1517957692.524 | 1517957707.65 | -7.678e-05 | -8.8423e-05 | -9.4465e-05 | -9.7707e-05 | -0.00010449 | -8.857e-05 | -6.2338e-05 | -8.1938e-05 | -8.1054e-05 | -8.5623e-05 | -9.4318e-05 | -7.6338e-05 | -8.9012e-05 | -6.7643e-05 | -9.3433e-05 | -8.7981e-05 | -9.6086e-05 | -9.697e-05 | -8.9749e-05 | -8.7833e-05
| 100 | -10.0 | 10 | -10.0 | 100 | 0.0 | 20.0 | 10.0 | 0.0 | 49.9849526E+00 | DC11V |  | -6.7774385895499995e-05 | 1.14457187754296e-05 | 1517957722.48 | 1517957737.642 | -6.3867e-05 | -7.6376e-05 | -6.5339e-05 | -8.2262e-05 | -7.461e-05 | -7.9908e-05 | -5.9158e-05 | -7.5787e-05 | -5.0034e-05 | -6.166e-05 | -7.6229e-05 | -7.5493e-05 | -4.3854e-05 | -7.873e-05 | -7.8142e-05 | -5.3272e-05 | -6.9018e-05 | -5.4596e-05 | -8.1085e-05 | -5.6068e-05
| 100 | 0 | 10 | 0.0 | 100 | 0 | 20.0 | 10.0 | 0.0 | 50.0247150E+00 | DC11V |  | -7.7390868436500002e-05 | 1.0703189605263014e-05 | 1517957752.483 | 1517957767.642 | -7.5948e-05 | -8.9342e-05 | -8.8018e-05 | -6.8736e-05 | -7.742e-05 | -8.4485e-05 | -8.5516e-05 | -8.6399e-05 | -6.9031e-05 | -6.1524e-05 | -6.7264e-05 | -7.8745e-05 | -8.4632e-05 | -7.9775e-05 | -7.9922e-05 | -9.4936e-05 | -8.5368e-05 | -6.9914e-05 | -7.1974e-05 | -4.8866e-05

## Results
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1.0002 | 0.00049367 | 49.0085 | 0.1 | 0.001
x gain | 1.0005 | 0.00061789 | 34.6293 | 0.1 | 0.001
x linearity | 1.0001 | 0.00051024 | 48.8452 | 0.1 | 0.001
x gain | 1 | 0.00057417 | 44.0937 | 0.1 | 0.001
m linearity | 0.99684 | 0.0055349 | 19.3918 | 0.1 | 0.001
m gain | -0.99602 | 0.0055702 | 19.9529 | 0.1 | 0.001
m linearity | 0.9967 | 0.0055291 | 19.3396 | 0.1 | 0.001
m gain | -0.99679 | 0.0055304 | 19.3513 | 0.1 | 0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 0.00045801 | 51.8717 | 0.1 | -0.001
x gain | 1 | 0.00071226 | 44.4284 | 0.1 | -0.001
x linearity | 1.0003 | 0.00056961 | 46.6474 | 0.1 | -0.001
x gain | 0.99975 | 0.00067054 | 43.8438 | 0.1 | -0.001
m linearity | 1.0002 | 0.0004693 | 51.7373 | 0.1 | -0.001
m gain | -1.0002 | 0.00078648 | 41.9116 | 0.1 | -0.001
m linearity | 0.99994 | 0.00068542 | 40.3018 | 0.1 | -0.001
m gain | -0.010002 | 1.9626e-06 | 314.3204 | 0.1 | -0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 7.5519e-05 | 35.2731 | 0.1 | 0.01
x gain | 0.99999 | 4.5262e-05 | 61.4132 | 0.1 | 0.01
x linearity | 1 | 5.8133e-05 | 40.2455 | 0.1 | 0.01
x gain | 1 | 4.8985e-05 | 54.301 | 0.1 | 0.01
m linearity | 1 | 8.2225e-05 | 40.2124 | 0.1 | 0.01
m gain | -0.99999 | 5.0103e-05 | 52.5604 | 0.1 | 0.01
m linearity | 0.99998 | 6.7231e-05 | 45.8658 | 0.1 | 0.01
m gain | -0.1 | 2.2051e-06 | 352.5632 | 0.1 | 0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.7572e-05 | 44.3336 | 0.1 | -0.01
x gain | 1 | 7.9356e-05 | 34.0213 | 0.1 | -0.01
x linearity | 1 | 6.4264e-05 | 37.5443 | 0.1 | -0.01
x gain | 1 | 4.9923e-05 | 49.7943 | 0.1 | -0.01
m linearity | 0.99999 | 6.4429e-05 | 45.2702 | 0.1 | -0.01
m gain | -0.1 | 2.1467e-06 | 357.6442 | 0.1 | -0.01
m linearity | 1 | 4.9622e-05 | 49.6955 | 0.1 | -0.01
m gain | -1 | 5.5008e-05 | 47.9803 | 0.1 | -0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.8226e-06 | 69.3873 | 0.1 | 0.05
x gain | 1 | 7.0629e-06 | 48.9766 | 0.1 | 0.05
x linearity | 1 | 5.7953e-06 | 58.3855 | 0.1 | 0.05
x gain | 1 | 5.7479e-06 | 56.3082 | 0.1 | 0.05
m linearity | 1 | 6.6782e-06 | 42.9984 | 0.1 | 0.05
m gain | -1 | 6.1654e-06 | 62.2562 | 0.1 | 0.05
m linearity | 1 | 5.3392e-06 | 59.0104 | 0.1 | 0.05
m gain | -0.1 | 2.872e-07 | 180.9392 | 0.1 | 0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99999 | 6.5708e-06 | 38.5756 | 0.1 | -0.05
x gain | 1 | 6.7644e-06 | 54.3099 | 0.1 | -0.05
x linearity | 1 | 5.227e-06 | 56.1941 | 0.1 | -0.05
x gain | 1 | 7.4661e-06 | 48.0487 | 0.1 | -0.05
m linearity | 1 | 6.6156e-06 | 39.2373 | 0.1 | -0.05
m gain | -0.1 | 2.9142e-07 | 167.6028 | 0.1 | -0.05
m linearity | 1 | 4.485e-06 | 61.9252 | 0.1 | -0.05
m gain | -1 | 6.6704e-06 | 50.3136 | 0.1 | -0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.4091e-06 | 48.1606 | 0.1 | 0.1
x gain | 1 | 5.0559e-06 | 51.9456 | 0.1 | 0.1
x linearity | 1 | 6.5673e-06 | 43.6876 | 0.1 | 0.1
x gain | 1 | 6.5856e-06 | 41.7134 | 0.1 | 0.1
m linearity | 1 | 5.1138e-06 | 50.044 | 0.1 | 0.1
m gain | -1 | 5.7553e-06 | 44.5399 | 0.1 | 0.1
m linearity | 0.99999 | 6.0502e-06 | 40.886 | 0.1 | 0.1
m gain | -1 | 5.4414e-06 | 48.8524 | 0.1 | 0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.474e-06 | 43.2458 | 0.1 | -0.1
x gain | 1 | 6.196e-06 | 47.171 | 0.1 | -0.1
x linearity | 1 | 5.9711e-06 | 45.5006 | 0.1 | -0.1
x gain | 1 | 6.5384e-06 | 45.2034 | 0.1 | -0.1
m linearity | 1 | 5.6848e-06 | 45.5556 | 0.1 | -0.1
m gain | -1 | 5.6689e-06 | 43.9932 | 0.1 | -0.1
m linearity | 1 | 6.0835e-06 | 45.31 | 0.1 | -0.1
m gain | -1 | 5.9228e-06 | 44.7807 | 0.1 | -0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.6052e-06 | 37.532 | 2 | 0.5
x gain | 1 | 1.4839e-06 | 47.5475 | 2 | 0.5
x linearity | 1 | 1.5216e-06 | 46.5368 | 2 | 0.5
x gain | 1 | 2.3778e-06 | 41.5333 | 2 | 0.5
m linearity | 1 | 1.8629e-06 | 44.0769 | 1 | 0.5
m gain | -1 | 1.6016e-06 | 48.5358 | 1 | 0.5
m linearity | 1 | 1.4364e-06 | 47.6326 | 1 | 0.5
m gain | -1 | 2.098e-06 | 40.3694 | 1 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.8872e-06 | 41.1299 | 2 | -0.5
x gain | 1 | 2.0769e-06 | 36.7942 | 2 | -0.5
x linearity | 1 | 2.1497e-06 | 54.3409 | 2 | -0.5
x gain | 1 | 2.4839e-06 | 60.6698 | 2 | -0.5
m linearity | 1 | 1.7166e-05 | 44.0427 | 1 | -0.5
m gain | -0.1 | 5.3094e-07 | 377.4682 | 1 | -0.5
m linearity | 0.99999 | 2.1242e-05 | 53.2845 | 1 | -0.5
m gain | -0.099999 | 1.6315e-06 | 24.793 | 1 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.6817e-07 | 44.985 | 2 | 1.0
x gain | 1 | 8.6367e-07 | 48.4833 | 2 | 1.0
x linearity | 1 | 9.1925e-07 | 46.2568 | 2 | 1.0
x gain | 1 | 8.7926e-07 | 49.9028 | 2 | 1.0
m linearity | 1 | 8.3717e-07 | 43.4993 | 1 | 1.0
m gain | -1 | 8.7774e-07 | 48.3705 | 1 | 1.0
m linearity | 1 | 8.8806e-07 | 47.8013 | 1 | 1.0
m gain | -1 | 9.8072e-07 | 44.4814 | 1 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 9.8588e-07 | 46.1616 | 2 | -1.0
x gain | 1 | 9.09e-07 | 43.9653 | 2 | -1.0
x linearity | 1 | 6.8908e-06 | 19.3618 | 2 | -1.0
x gain | 0.99999 | 4.5647e-06 | 21.6134 | 2 | -1.0
m linearity | 1 | 9.7287e-07 | 46.5678 | 1 | -1.0
m gain | -1 | 8.7222e-07 | 46.7561 | 1 | -1.0
m linearity | 1 | 8.7255e-07 | 48.0097 | 1 | -1.0
m gain | -1 | 1.0414e-06 | 41.515 | 1 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.0627e-06 | 44.1167 | 20 | 0.5
x gain | 1 | 8.9862e-06 | 42.4872 | 20 | 0.5
x linearity | 0.99999 | 7.8867e-06 | 46.7492 | 20 | 0.5
x gain | 1 | 6.5614e-06 | 52.6338 | 20 | 0.5
m linearity | 1 | 6.7773e-06 | 46.7937 | 10 | 0.5
m gain | -0.99999 | 1.0019e-05 | 43.8029 | 10 | 0.5
m linearity | 1 | 7.2761e-06 | 46.511 | 10 | 0.5
m gain | -0.99999 | 9.1368e-06 | 33.3227 | 10 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.5238e-06 | 44.9309 | 20 | -0.5
x gain | 1 | 8.7093e-06 | 39.9327 | 20 | -0.5
x linearity | 1 | 8.6892e-06 | 45.2815 | 20 | -0.5
x gain | 1 | 9.4891e-06 | 43.9021 | 20 | -0.5
m linearity | 1 | 6.1156e-06 | 57.0155 | 10 | -0.5
m gain | -1 | 7.3687e-06 | 49.9674 | 10 | -0.5
m linearity | 1 | 9.159e-06 | 43.7299 | 10 | -0.5
m gain | -1 | 9.7491e-06 | 44.0569 | 10 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.9346e-06 | 43.6545 | 20 | 1.0
x gain | 1 | 2.4015e-05 | 19.6434 | 20 | 1.0
x linearity | 1 | 4.6445e-06 | 44.6107 | 20 | 1.0
x gain | 1 | 4.6221e-06 | 44.6971 | 20 | 1.0
m linearity | 1 | 4.3665e-06 | 45.1212 | 10 | 1.0
m gain | -1 | 2.4096e-05 | 19.908 | 10 | 1.0
m linearity | 1 | 4.0187e-06 | 40.7988 | 10 | 1.0
m gain | -1 | 4.7847e-06 | 44.9325 | 10 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.845e-06 | 48.3955 | 20 | -1.0
x gain | 1 | 4.2152e-06 | 46.9228 | 20 | -1.0
x linearity | 1 | 3.1942e-06 | 43.9948 | 20 | -1.0
x gain | 1 | 4.3753e-06 | 41.4273 | 20 | -1.0
m linearity | 1 | 4.9055e-06 | 38.252 | 10 | -1.0
m gain | -1 | 4.4204e-06 | 46.5619 | 10 | -1.0
m linearity | 1 | 3.832e-06 | 48.1552 | 10 | -1.0
m gain | -1 | 4.9182e-06 | 44.3428 | 10 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.6444e-07 | 53.0941 | 20 | 5.0
x gain | 1 | 4.5425e-07 | 54.5783 | 20 | 5.0
x linearity | 1 | 3.3189e-07 | 66.8714 | 20 | 5.0
x gain | 1 | 5.0766e-07 | 48.0114 | 20 | 5.0
m linearity | 1 | 4.184e-07 | 52.439 | 10 | 5.0
m gain | -1 | 5.1975e-07 | 52.9764 | 10 | 5.0
m linearity | 1 | 3.6292e-07 | 56.0706 | 10 | 5.0
m gain | -1 | 4.8957e-07 | 50.1564 | 10 | 5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.997e-07 | 54.6342 | 20 | -5.0
x gain | 1 | 4.0378e-07 | 61.4212 | 20 | -5.0
x linearity | 1 | 3.5759e-07 | 65.7501 | 20 | -5.0
x gain | 1 | 4.7001e-07 | 51.613 | 20 | -5.0
m linearity | 1 | 4.0506e-07 | 54.4373 | 10 | -5.0
m gain | -1 | 4.3405e-07 | 55.346 | 10 | -5.0
m linearity | 1 | 4.1565e-07 | 50.9243 | 10 | -5.0
m gain | -1 | 4.3656e-07 | 58.9781 | 10 | -5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.6682e-07 | 52.4447 | 20 | 10.0
x gain | 1 | 4.4462e-07 | 41.8406 | 20 | 10.0
x linearity | 1 | 1.4351e-07 | 79.5498 | 20 | 10.0
x gain | 1 | 1.2234e-07 | 173.4429 | 20 | 10.0
m linearity | 1 | 3.847e-07 | 53.3231 | 10 | 10.0
m gain | -1 | 5.098e-07 | 46.5276 | 10 | 10.0
m linearity | 1 | 2.2089e-07 | 48.1264 | 10 | 10.0
m gain | -1 | 1.2544e-07 | 220.3776 | 10 | 10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.2511e-07 | 171.9043 | 20 | -10.0
x gain | 1 | 1.5084e-07 | 100.241 | 20 | -10.0
x linearity | 1 | 1.1747e-07 | 182.8919 | 20 | -10.0
x gain | 1 | 1.4225e-07 | 90.4906 | 20 | -10.0
m linearity | 1 | 1.3805e-07 | 169.3247 | 10 | -10.0
m gain | -1 | 1.3216e-07 | 193.2511 | 10 | -10.0
m linearity | 1 | 2.2021e-07 | 34.8952 | 10 | -10.0
m gain | -1 | 1.2452e-07 | 208.6669 | 10 | -10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.7887e-06 | 28.5451 | 100 | 50.0
x gain | 1 | 4.3727e-06 | 111.52 | 100 | 50.0
x linearity | 0.99998 | 6.5347e-06 | 38.388 | 100 | 50.0
x gain | 0.99998 | 6.7544e-06 | 42.8707 | 100 | 50.0
m linearity | 1 | 2.3237e-05 | 24.5669 | 100 | 50.0
m gain | -1 | 4.6368e-06 | 112.7143 | 100 | 50.0
m linearity | 1 | 6.4531e-06 | 36.6446 | 100 | 50.0
m gain | -1 | 6.6155e-06 | 40.1001 | 100 | 50.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.1366e-06 | 47.5088 | 100 | 100.0
x gain | 1 | 1.3367e-06 | 43.2855 | 100 | 100.0
x linearity | 1 | 1.0388e-06 | 49.8772 | 100 | 100.0
x gain | 1 | 1.499e-06 | 43.6304 | 100 | 100.0
m linearity | 1 | 1.2356e-06 | 45.7186 | 100 | 100.0
m gain | -1 | 1.3041e-06 | 42.3789 | 100 | 100.0
m linearity | 1 | 1.1421e-06 | 46.507 | 100 | 100.0
m gain | -1 | 1.375e-06 | 42.8925 | 100 | 100.0

## Results1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1.0002 | 0.00049367 | 49.0085 | 0.1 | 0.001
x gain | 1.0005 | 0.00061789 | 34.6293 | 0.1 | 0.001
x linearity | 1.0001 | 0.00051024 | 48.8452 | 0.1 | 0.001
x gain | 1 | 0.00057417 | 44.0937 | 0.1 | 0.001
m linearity | 0.99684 | 0.0055349 | 19.3918 | 0.1 | 0.001
m gain | -0.99602 | 0.0055702 | 19.9529 | 0.1 | 0.001
m linearity | 0.9967 | 0.0055291 | 19.3396 | 0.1 | 0.001
m gain | -0.99679 | 0.0055304 | 19.3513 | 0.1 | 0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 0.00045801 | 51.8717 | 0.1 | -0.001
x gain | 1 | 0.00071226 | 44.4284 | 0.1 | -0.001
x linearity | 1.0003 | 0.00056961 | 46.6474 | 0.1 | -0.001
x gain | 0.99975 | 0.00067054 | 43.8438 | 0.1 | -0.001
m linearity | 1.0002 | 0.0004693 | 51.7373 | 0.1 | -0.001
m gain | -1.0002 | 0.00078648 | 41.9116 | 0.1 | -0.001
m linearity | 0.99994 | 0.00068542 | 40.3018 | 0.1 | -0.001
m gain | -0.010002 | 1.9626e-06 | 314.3204 | 0.1 | -0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 7.5519e-05 | 35.2731 | 0.1 | 0.01
x gain | 0.99999 | 4.5262e-05 | 61.4132 | 0.1 | 0.01
x linearity | 1 | 5.8133e-05 | 40.2455 | 0.1 | 0.01
x gain | 1 | 4.8985e-05 | 54.301 | 0.1 | 0.01
m linearity | 1 | 8.2225e-05 | 40.2124 | 0.1 | 0.01
m gain | -0.99999 | 5.0103e-05 | 52.5604 | 0.1 | 0.01
m linearity | 0.99998 | 6.7231e-05 | 45.8658 | 0.1 | 0.01
m gain | -0.1 | 2.2051e-06 | 352.5632 | 0.1 | 0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.7572e-05 | 44.3336 | 0.1 | -0.01
x gain | 1 | 7.9356e-05 | 34.0213 | 0.1 | -0.01
x linearity | 1 | 6.4264e-05 | 37.5443 | 0.1 | -0.01
x gain | 1 | 4.9923e-05 | 49.7943 | 0.1 | -0.01
m linearity | 0.99999 | 6.4429e-05 | 45.2702 | 0.1 | -0.01
m gain | -0.1 | 2.1467e-06 | 357.6442 | 0.1 | -0.01
m linearity | 1 | 4.9622e-05 | 49.6955 | 0.1 | -0.01
m gain | -1 | 5.5008e-05 | 47.9803 | 0.1 | -0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.8226e-06 | 69.3873 | 0.1 | 0.05
x gain | 1 | 7.0629e-06 | 48.9766 | 0.1 | 0.05
x linearity | 1 | 5.7953e-06 | 58.3855 | 0.1 | 0.05
x gain | 1 | 5.7479e-06 | 56.3082 | 0.1 | 0.05
m linearity | 1 | 6.6782e-06 | 42.9984 | 0.1 | 0.05
m gain | -1 | 6.1654e-06 | 62.2562 | 0.1 | 0.05
m linearity | 1 | 5.3392e-06 | 59.0104 | 0.1 | 0.05
m gain | -0.1 | 2.872e-07 | 180.9392 | 0.1 | 0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99999 | 6.5708e-06 | 38.5756 | 0.1 | -0.05
x gain | 1 | 6.7644e-06 | 54.3099 | 0.1 | -0.05
x linearity | 1 | 5.227e-06 | 56.1941 | 0.1 | -0.05
x gain | 1 | 7.4661e-06 | 48.0487 | 0.1 | -0.05
m linearity | 1 | 6.6156e-06 | 39.2373 | 0.1 | -0.05
m gain | -0.1 | 2.9142e-07 | 167.6028 | 0.1 | -0.05
m linearity | 1 | 4.485e-06 | 61.9252 | 0.1 | -0.05
m gain | -1 | 6.6704e-06 | 50.3136 | 0.1 | -0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.4091e-06 | 48.1606 | 0.1 | 0.1
x gain | 1 | 5.0559e-06 | 51.9456 | 0.1 | 0.1
x linearity | 1 | 6.5673e-06 | 43.6876 | 0.1 | 0.1
x gain | 1 | 6.5856e-06 | 41.7134 | 0.1 | 0.1
m linearity | 1 | 5.1138e-06 | 50.044 | 0.1 | 0.1
m gain | -1 | 5.7553e-06 | 44.5399 | 0.1 | 0.1
m linearity | 0.99999 | 6.0502e-06 | 40.886 | 0.1 | 0.1
m gain | -1 | 5.4414e-06 | 48.8524 | 0.1 | 0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.474e-06 | 43.2458 | 0.1 | -0.1
x gain | 1 | 6.196e-06 | 47.171 | 0.1 | -0.1
x linearity | 1 | 5.9711e-06 | 45.5006 | 0.1 | -0.1
x gain | 1 | 6.5384e-06 | 45.2034 | 0.1 | -0.1
m linearity | 1 | 5.6848e-06 | 45.5556 | 0.1 | -0.1
m gain | -1 | 5.6689e-06 | 43.9932 | 0.1 | -0.1
m linearity | 1 | 6.0835e-06 | 45.31 | 0.1 | -0.1
m gain | -1 | 5.9228e-06 | 44.7807 | 0.1 | -0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.6052e-06 | 37.532 | 2 | 0.5
x gain | 1 | 1.4839e-06 | 47.5475 | 2 | 0.5
x linearity | 1 | 1.5216e-06 | 46.5368 | 2 | 0.5
x gain | 1 | 2.3778e-06 | 41.5333 | 2 | 0.5
m linearity | 1 | 1.8629e-06 | 44.0769 | 1 | 0.5
m gain | -1 | 1.6016e-06 | 48.5358 | 1 | 0.5
m linearity | 1 | 1.4364e-06 | 47.6326 | 1 | 0.5
m gain | -1 | 2.098e-06 | 40.3694 | 1 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.8872e-06 | 41.1299 | 2 | -0.5
x gain | 1 | 2.0769e-06 | 36.7942 | 2 | -0.5
x linearity | 1 | 2.1497e-06 | 54.3409 | 2 | -0.5
x gain | 1 | 2.4839e-06 | 60.6698 | 2 | -0.5
m linearity | 1 | 1.7166e-05 | 44.0427 | 1 | -0.5
m gain | -0.1 | 5.3094e-07 | 377.4682 | 1 | -0.5
m linearity | 0.99999 | 2.1242e-05 | 53.2845 | 1 | -0.5
m gain | -0.099999 | 1.6315e-06 | 24.793 | 1 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.6817e-07 | 44.985 | 2 | 1.0
x gain | 1 | 8.6367e-07 | 48.4833 | 2 | 1.0
x linearity | 1 | 9.1925e-07 | 46.2568 | 2 | 1.0
x gain | 1 | 8.7926e-07 | 49.9028 | 2 | 1.0
m linearity | 1 | 8.3717e-07 | 43.4993 | 1 | 1.0
m gain | -1 | 8.7774e-07 | 48.3705 | 1 | 1.0
m linearity | 1 | 8.8806e-07 | 47.8013 | 1 | 1.0
m gain | -1 | 9.8072e-07 | 44.4814 | 1 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 9.8588e-07 | 46.1616 | 2 | -1.0
x gain | 1 | 9.09e-07 | 43.9653 | 2 | -1.0
x linearity | 1 | 6.8908e-06 | 19.3618 | 2 | -1.0
x gain | 0.99999 | 4.5647e-06 | 21.6134 | 2 | -1.0
m linearity | 1 | 9.7287e-07 | 46.5678 | 1 | -1.0
m gain | -1 | 8.7222e-07 | 46.7561 | 1 | -1.0
m linearity | 1 | 8.7255e-07 | 48.0097 | 1 | -1.0
m gain | -1 | 1.0414e-06 | 41.515 | 1 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.0627e-06 | 44.1167 | 20 | 0.5
x gain | 1 | 8.9862e-06 | 42.4872 | 20 | 0.5
x linearity | 0.99999 | 7.8867e-06 | 46.7492 | 20 | 0.5
x gain | 1 | 6.5614e-06 | 52.6338 | 20 | 0.5
m linearity | 1 | 6.7773e-06 | 46.7937 | 10 | 0.5
m gain | -0.99999 | 1.0019e-05 | 43.8029 | 10 | 0.5
m linearity | 1 | 7.2761e-06 | 46.511 | 10 | 0.5
m gain | -0.99999 | 9.1368e-06 | 33.3227 | 10 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.5238e-06 | 44.9309 | 20 | -0.5
x gain | 1 | 8.7093e-06 | 39.9327 | 20 | -0.5
x linearity | 1 | 8.6892e-06 | 45.2815 | 20 | -0.5
x gain | 1 | 9.4891e-06 | 43.9021 | 20 | -0.5
m linearity | 1 | 6.1156e-06 | 57.0155 | 10 | -0.5
m gain | -1 | 7.3687e-06 | 49.9674 | 10 | -0.5
m linearity | 1 | 9.159e-06 | 43.7299 | 10 | -0.5
m gain | -1 | 9.7491e-06 | 44.0569 | 10 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.9346e-06 | 43.6545 | 20 | 1.0
x gain | 1 | 2.4015e-05 | 19.6434 | 20 | 1.0
x linearity | 1 | 4.6445e-06 | 44.6107 | 20 | 1.0
x gain | 1 | 4.6221e-06 | 44.6971 | 20 | 1.0
m linearity | 1 | 4.3665e-06 | 45.1212 | 10 | 1.0
m gain | -1 | 2.4096e-05 | 19.908 | 10 | 1.0
m linearity | 1 | 4.0187e-06 | 40.7988 | 10 | 1.0
m gain | -1 | 4.7847e-06 | 44.9325 | 10 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.845e-06 | 48.3955 | 20 | -1.0
x gain | 1 | 4.2152e-06 | 46.9228 | 20 | -1.0
x linearity | 1 | 3.1942e-06 | 43.9948 | 20 | -1.0
x gain | 1 | 4.3753e-06 | 41.4273 | 20 | -1.0
m linearity | 1 | 4.9055e-06 | 38.252 | 10 | -1.0
m gain | -1 | 4.4204e-06 | 46.5619 | 10 | -1.0
m linearity | 1 | 3.832e-06 | 48.1552 | 10 | -1.0
m gain | -1 | 4.9182e-06 | 44.3428 | 10 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.6444e-07 | 53.0941 | 20 | 5.0
x gain | 1 | 4.5425e-07 | 54.5783 | 20 | 5.0
x linearity | 1 | 3.3189e-07 | 66.8714 | 20 | 5.0
x gain | 1 | 5.0766e-07 | 48.0114 | 20 | 5.0
m linearity | 1 | 4.184e-07 | 52.439 | 10 | 5.0
m gain | -1 | 5.1975e-07 | 52.9764 | 10 | 5.0
m linearity | 1 | 3.6292e-07 | 56.0706 | 10 | 5.0
m gain | -1 | 4.8957e-07 | 50.1564 | 10 | 5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.997e-07 | 54.6342 | 20 | -5.0
x gain | 1 | 4.0378e-07 | 61.4212 | 20 | -5.0
x linearity | 1 | 3.5759e-07 | 65.7501 | 20 | -5.0
x gain | 1 | 4.7001e-07 | 51.613 | 20 | -5.0
m linearity | 1 | 4.0506e-07 | 54.4373 | 10 | -5.0
m gain | -1 | 4.3405e-07 | 55.346 | 10 | -5.0
m linearity | 1 | 4.1565e-07 | 50.9243 | 10 | -5.0
m gain | -1 | 4.3656e-07 | 58.9781 | 10 | -5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.6682e-07 | 52.4447 | 20 | 10.0
x gain | 1 | 4.4462e-07 | 41.8406 | 20 | 10.0
x linearity | 1 | 1.4351e-07 | 79.5498 | 20 | 10.0
x gain | 1 | 1.2234e-07 | 173.4429 | 20 | 10.0
m linearity | 1 | 3.847e-07 | 53.3231 | 10 | 10.0
m gain | -1 | 5.098e-07 | 46.5276 | 10 | 10.0
m linearity | 1 | 2.2089e-07 | 48.1264 | 10 | 10.0
m gain | -1 | 1.2544e-07 | 220.3776 | 10 | 10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.2511e-07 | 171.9043 | 20 | -10.0
x gain | 1 | 1.5084e-07 | 100.241 | 20 | -10.0
x linearity | 1 | 1.1747e-07 | 182.8919 | 20 | -10.0
x gain | 1 | 1.4225e-07 | 90.4906 | 20 | -10.0
m linearity | 1 | 1.3805e-07 | 169.3247 | 10 | -10.0
m gain | -1 | 1.3216e-07 | 193.2511 | 10 | -10.0
m linearity | 1 | 2.2021e-07 | 34.8952 | 10 | -10.0
m gain | -1 | 1.2452e-07 | 208.6669 | 10 | -10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.7887e-06 | 28.5451 | 100 | 50.0
x gain | 1 | 4.3727e-06 | 111.52 | 100 | 50.0
x linearity | 0.99998 | 6.5347e-06 | 38.388 | 100 | 50.0
x gain | 0.99998 | 6.7544e-06 | 42.8707 | 100 | 50.0
m linearity | 1 | 2.3237e-05 | 24.5669 | 100 | 50.0
m gain | -1 | 4.6368e-06 | 112.7143 | 100 | 50.0
m linearity | 1 | 6.4531e-06 | 36.6446 | 100 | 50.0
m gain | -1 | 6.6155e-06 | 40.1001 | 100 | 50.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.1366e-06 | 47.5088 | 100 | 100.0
x gain | 1 | 1.3367e-06 | 43.2855 | 100 | 100.0
x linearity | 1 | 1.0388e-06 | 49.8772 | 100 | 100.0
x gain | 1 | 1.499e-06 | 43.6304 | 100 | 100.0
m linearity | 1 | 1.2356e-06 | 45.7186 | 100 | 100.0
m gain | -1 | 1.3041e-06 | 42.3789 | 100 | 100.0
m linearity | 1 | 1.1421e-06 | 46.507 | 100 | 100.0
m gain | -1 | 1.375e-06 | 42.8925 | 100 | 100.0

## Results2
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1.0002 | 0.00049367 | 49.0085 | 0.1 | 0.001
x gain | 1.0005 | 0.00061789 | 34.6293 | 0.1 | 0.001
x linearity | 1.0001 | 0.00051024 | 48.8452 | 0.1 | 0.001
x gain | 1 | 0.00057417 | 44.0937 | 0.1 | 0.001
m linearity | 0.99684 | 0.0055349 | 19.3918 | 0.1 | 0.001
m gain | -0.99602 | 0.0055702 | 19.9529 | 0.1 | 0.001
m linearity | 0.9967 | 0.0055291 | 19.3396 | 0.1 | 0.001
m gain | -0.99679 | 0.0055304 | 19.3513 | 0.1 | 0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 0.00045801 | 51.8717 | 0.1 | -0.001
x gain | 1 | 0.00071226 | 44.4284 | 0.1 | -0.001
x linearity | 1.0003 | 0.00056961 | 46.6474 | 0.1 | -0.001
x gain | 0.99975 | 0.00067054 | 43.8438 | 0.1 | -0.001
m linearity | 1.0002 | 0.0004693 | 51.7373 | 0.1 | -0.001
m gain | -1.0002 | 0.00078648 | 41.9116 | 0.1 | -0.001
m linearity | 0.99994 | 0.00068542 | 40.3018 | 0.1 | -0.001
m gain | -0.010002 | 1.9626e-06 | 314.3204 | 0.1 | -0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 7.5519e-05 | 35.2731 | 0.1 | 0.01
x gain | 0.99999 | 4.5262e-05 | 61.4132 | 0.1 | 0.01
x linearity | 1 | 5.8133e-05 | 40.2455 | 0.1 | 0.01
x gain | 1 | 4.8985e-05 | 54.301 | 0.1 | 0.01
m linearity | 1 | 8.2225e-05 | 40.2124 | 0.1 | 0.01
m gain | -0.99999 | 5.0103e-05 | 52.5604 | 0.1 | 0.01
m linearity | 0.99998 | 6.7231e-05 | 45.8658 | 0.1 | 0.01
m gain | -0.1 | 2.2051e-06 | 352.5632 | 0.1 | 0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.7572e-05 | 44.3336 | 0.1 | -0.01
x gain | 1 | 7.9356e-05 | 34.0213 | 0.1 | -0.01
x linearity | 1 | 6.4264e-05 | 37.5443 | 0.1 | -0.01
x gain | 1 | 4.9923e-05 | 49.7943 | 0.1 | -0.01
m linearity | 0.99999 | 6.4429e-05 | 45.2702 | 0.1 | -0.01
m gain | -0.1 | 2.1467e-06 | 357.6442 | 0.1 | -0.01
m linearity | 1 | 4.9622e-05 | 49.6955 | 0.1 | -0.01
m gain | -1 | 5.5008e-05 | 47.9803 | 0.1 | -0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.8226e-06 | 69.3873 | 0.1 | 0.05
x gain | 1 | 7.0629e-06 | 48.9766 | 0.1 | 0.05
x linearity | 1 | 5.7953e-06 | 58.3855 | 0.1 | 0.05
x gain | 1 | 5.7479e-06 | 56.3082 | 0.1 | 0.05
m linearity | 1 | 6.6782e-06 | 42.9984 | 0.1 | 0.05
m gain | -1 | 6.1654e-06 | 62.2562 | 0.1 | 0.05
m linearity | 1 | 5.3392e-06 | 59.0104 | 0.1 | 0.05
m gain | -0.1 | 2.872e-07 | 180.9392 | 0.1 | 0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99999 | 6.5708e-06 | 38.5756 | 0.1 | -0.05
x gain | 1 | 6.7644e-06 | 54.3099 | 0.1 | -0.05
x linearity | 1 | 5.227e-06 | 56.1941 | 0.1 | -0.05
x gain | 1 | 7.4661e-06 | 48.0487 | 0.1 | -0.05
m linearity | 1 | 6.6156e-06 | 39.2373 | 0.1 | -0.05
m gain | -0.1 | 2.9142e-07 | 167.6028 | 0.1 | -0.05
m linearity | 1 | 4.485e-06 | 61.9252 | 0.1 | -0.05
m gain | -1 | 6.6704e-06 | 50.3136 | 0.1 | -0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.4091e-06 | 48.1606 | 0.1 | 0.1
x gain | 1 | 5.0559e-06 | 51.9456 | 0.1 | 0.1
x linearity | 1 | 6.5673e-06 | 43.6876 | 0.1 | 0.1
x gain | 1 | 6.5856e-06 | 41.7134 | 0.1 | 0.1
m linearity | 1 | 5.1138e-06 | 50.044 | 0.1 | 0.1
m gain | -1 | 5.7553e-06 | 44.5399 | 0.1 | 0.1
m linearity | 0.99999 | 6.0502e-06 | 40.886 | 0.1 | 0.1
m gain | -1 | 5.4414e-06 | 48.8524 | 0.1 | 0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.474e-06 | 43.2458 | 0.1 | -0.1
x gain | 1 | 6.196e-06 | 47.171 | 0.1 | -0.1
x linearity | 1 | 5.9711e-06 | 45.5006 | 0.1 | -0.1
x gain | 1 | 6.5384e-06 | 45.2034 | 0.1 | -0.1
m linearity | 1 | 5.6848e-06 | 45.5556 | 0.1 | -0.1
m gain | -1 | 5.6689e-06 | 43.9932 | 0.1 | -0.1
m linearity | 1 | 6.0835e-06 | 45.31 | 0.1 | -0.1
m gain | -1 | 5.9228e-06 | 44.7807 | 0.1 | -0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.6052e-06 | 37.532 | 2 | 0.5
x gain | 1 | 1.4839e-06 | 47.5475 | 2 | 0.5
x linearity | 1 | 1.5216e-06 | 46.5368 | 2 | 0.5
x gain | 1 | 2.3778e-06 | 41.5333 | 2 | 0.5
m linearity | 1 | 1.8629e-06 | 44.0769 | 1 | 0.5
m gain | -1 | 1.6016e-06 | 48.5358 | 1 | 0.5
m linearity | 1 | 1.4364e-06 | 47.6326 | 1 | 0.5
m gain | -1 | 2.098e-06 | 40.3694 | 1 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.8872e-06 | 41.1299 | 2 | -0.5
x gain | 1 | 2.0769e-06 | 36.7942 | 2 | -0.5
x linearity | 1 | 2.1497e-06 | 54.3409 | 2 | -0.5
x gain | 1 | 2.4839e-06 | 60.6698 | 2 | -0.5
m linearity | 1 | 1.7166e-05 | 44.0427 | 1 | -0.5
m gain | -0.1 | 5.3094e-07 | 377.4682 | 1 | -0.5
m linearity | 0.99999 | 2.1242e-05 | 53.2845 | 1 | -0.5
m gain | -0.099999 | 1.6315e-06 | 24.793 | 1 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.6817e-07 | 44.985 | 2 | 1.0
x gain | 1 | 8.6367e-07 | 48.4833 | 2 | 1.0
x linearity | 1 | 9.1925e-07 | 46.2568 | 2 | 1.0
x gain | 1 | 8.7926e-07 | 49.9028 | 2 | 1.0
m linearity | 1 | 8.3717e-07 | 43.4993 | 1 | 1.0
m gain | -1 | 8.7774e-07 | 48.3705 | 1 | 1.0
m linearity | 1 | 8.8806e-07 | 47.8013 | 1 | 1.0
m gain | -1 | 9.8072e-07 | 44.4814 | 1 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 9.8588e-07 | 46.1616 | 2 | -1.0
x gain | 1 | 9.09e-07 | 43.9653 | 2 | -1.0
x linearity | 1 | 6.8908e-06 | 19.3618 | 2 | -1.0
x gain | 0.99999 | 4.5647e-06 | 21.6134 | 2 | -1.0
m linearity | 1 | 9.7287e-07 | 46.5678 | 1 | -1.0
m gain | -1 | 8.7222e-07 | 46.7561 | 1 | -1.0
m linearity | 1 | 8.7255e-07 | 48.0097 | 1 | -1.0
m gain | -1 | 1.0414e-06 | 41.515 | 1 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.0627e-06 | 44.1167 | 20 | 0.5
x gain | 1 | 8.9862e-06 | 42.4872 | 20 | 0.5
x linearity | 0.99999 | 7.8867e-06 | 46.7492 | 20 | 0.5
x gain | 1 | 6.5614e-06 | 52.6338 | 20 | 0.5
m linearity | 1 | 6.7773e-06 | 46.7937 | 10 | 0.5
m gain | -0.99999 | 1.0019e-05 | 43.8029 | 10 | 0.5
m linearity | 1 | 7.2761e-06 | 46.511 | 10 | 0.5
m gain | -0.99999 | 9.1368e-06 | 33.3227 | 10 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.5238e-06 | 44.9309 | 20 | -0.5
x gain | 1 | 8.7093e-06 | 39.9327 | 20 | -0.5
x linearity | 1 | 8.6892e-06 | 45.2815 | 20 | -0.5
x gain | 1 | 9.4891e-06 | 43.9021 | 20 | -0.5
m linearity | 1 | 6.1156e-06 | 57.0155 | 10 | -0.5
m gain | -1 | 7.3687e-06 | 49.9674 | 10 | -0.5
m linearity | 1 | 9.159e-06 | 43.7299 | 10 | -0.5
m gain | -1 | 9.7491e-06 | 44.0569 | 10 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.9346e-06 | 43.6545 | 20 | 1.0
x gain | 1 | 2.4015e-05 | 19.6434 | 20 | 1.0
x linearity | 1 | 4.6445e-06 | 44.6107 | 20 | 1.0
x gain | 1 | 4.6221e-06 | 44.6971 | 20 | 1.0
m linearity | 1 | 4.3665e-06 | 45.1212 | 10 | 1.0
m gain | -1 | 2.4096e-05 | 19.908 | 10 | 1.0
m linearity | 1 | 4.0187e-06 | 40.7988 | 10 | 1.0
m gain | -1 | 4.7847e-06 | 44.9325 | 10 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.845e-06 | 48.3955 | 20 | -1.0
x gain | 1 | 4.2152e-06 | 46.9228 | 20 | -1.0
x linearity | 1 | 3.1942e-06 | 43.9948 | 20 | -1.0
x gain | 1 | 4.3753e-06 | 41.4273 | 20 | -1.0
m linearity | 1 | 4.9055e-06 | 38.252 | 10 | -1.0
m gain | -1 | 4.4204e-06 | 46.5619 | 10 | -1.0
m linearity | 1 | 3.832e-06 | 48.1552 | 10 | -1.0
m gain | -1 | 4.9182e-06 | 44.3428 | 10 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.6444e-07 | 53.0941 | 20 | 5.0
x gain | 1 | 4.5425e-07 | 54.5783 | 20 | 5.0
x linearity | 1 | 3.3189e-07 | 66.8714 | 20 | 5.0
x gain | 1 | 5.0766e-07 | 48.0114 | 20 | 5.0
m linearity | 1 | 4.184e-07 | 52.439 | 10 | 5.0
m gain | -1 | 5.1975e-07 | 52.9764 | 10 | 5.0
m linearity | 1 | 3.6292e-07 | 56.0706 | 10 | 5.0
m gain | -1 | 4.8957e-07 | 50.1564 | 10 | 5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.997e-07 | 54.6342 | 20 | -5.0
x gain | 1 | 4.0378e-07 | 61.4212 | 20 | -5.0
x linearity | 1 | 3.5759e-07 | 65.7501 | 20 | -5.0
x gain | 1 | 4.7001e-07 | 51.613 | 20 | -5.0
m linearity | 1 | 4.0506e-07 | 54.4373 | 10 | -5.0
m gain | -1 | 4.3405e-07 | 55.346 | 10 | -5.0
m linearity | 1 | 4.1565e-07 | 50.9243 | 10 | -5.0
m gain | -1 | 4.3656e-07 | 58.9781 | 10 | -5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.6682e-07 | 52.4447 | 20 | 10.0
x gain | 1 | 4.4462e-07 | 41.8406 | 20 | 10.0
x linearity | 1 | 1.4351e-07 | 79.5498 | 20 | 10.0
x gain | 1 | 1.2234e-07 | 173.4429 | 20 | 10.0
m linearity | 1 | 3.847e-07 | 53.3231 | 10 | 10.0
m gain | -1 | 5.098e-07 | 46.5276 | 10 | 10.0
m linearity | 1 | 2.2089e-07 | 48.1264 | 10 | 10.0
m gain | -1 | 1.2544e-07 | 220.3776 | 10 | 10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.2511e-07 | 171.9043 | 20 | -10.0
x gain | 1 | 1.5084e-07 | 100.241 | 20 | -10.0
x linearity | 1 | 1.1747e-07 | 182.8919 | 20 | -10.0
x gain | 1 | 1.4225e-07 | 90.4906 | 20 | -10.0
m linearity | 1 | 1.3805e-07 | 169.3247 | 10 | -10.0
m gain | -1 | 1.3216e-07 | 193.2511 | 10 | -10.0
m linearity | 1 | 2.2021e-07 | 34.8952 | 10 | -10.0
m gain | -1 | 1.2452e-07 | 208.6669 | 10 | -10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.7887e-06 | 28.5451 | 100 | 50.0
x gain | 1 | 4.3727e-06 | 111.52 | 100 | 50.0
x linearity | 0.99998 | 6.5347e-06 | 38.388 | 100 | 50.0
x gain | 0.99998 | 6.7544e-06 | 42.8707 | 100 | 50.0
m linearity | 1 | 2.3237e-05 | 24.5669 | 100 | 50.0
m gain | -1 | 4.6368e-06 | 112.7143 | 100 | 50.0
m linearity | 1 | 6.4531e-06 | 36.6446 | 100 | 50.0
m gain | -1 | 6.6155e-06 | 40.1001 | 100 | 50.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.1366e-06 | 47.5088 | 100 | 100.0
x gain | 1 | 1.3367e-06 | 43.2855 | 100 | 100.0
x linearity | 1 | 1.0388e-06 | 49.8772 | 100 | 100.0
x gain | 1 | 1.499e-06 | 43.6304 | 100 | 100.0
m linearity | 1 | 1.2356e-06 | 45.7186 | 100 | 100.0
m gain | -1 | 1.3041e-06 | 42.3789 | 100 | 100.0
m linearity | 1 | 1.1421e-06 | 46.507 | 100 | 100.0
m gain | -1 | 1.375e-06 | 42.8925 | 100 | 100.0

## Results3
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1.0002 | 0.00049367 | 49.0085 | 0.1 | 0.001
x gain | 1.0005 | 0.00061789 | 34.6293 | 0.1 | 0.001
x linearity | 1.0001 | 0.00051024 | 48.8452 | 0.1 | 0.001
x gain | 1 | 0.00057417 | 44.0937 | 0.1 | 0.001
m linearity | 0.99684 | 0.0055349 | 19.3918 | 0.1 | 0.001
m gain | -0.99602 | 0.0055702 | 19.9529 | 0.1 | 0.001
m linearity | 0.9967 | 0.0055291 | 19.3396 | 0.1 | 0.001
m gain | -0.99679 | 0.0055304 | 19.3513 | 0.1 | 0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 0.00045801 | 51.8717 | 0.1 | -0.001
x gain | 1 | 0.00071226 | 44.4284 | 0.1 | -0.001
x linearity | 1.0003 | 0.00056961 | 46.6474 | 0.1 | -0.001
x gain | 0.99975 | 0.00067054 | 43.8438 | 0.1 | -0.001
m linearity | 1.0002 | 0.0004693 | 51.7373 | 0.1 | -0.001
m gain | -1.0002 | 0.00078648 | 41.9116 | 0.1 | -0.001
m linearity | 0.99994 | 0.00068542 | 40.3018 | 0.1 | -0.001
m gain | -0.010002 | 1.9626e-06 | 314.3204 | 0.1 | -0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 7.5519e-05 | 35.2731 | 0.1 | 0.01
x gain | 0.99999 | 4.5262e-05 | 61.4132 | 0.1 | 0.01
x linearity | 1 | 5.8133e-05 | 40.2455 | 0.1 | 0.01
x gain | 1 | 4.8985e-05 | 54.301 | 0.1 | 0.01
m linearity | 1 | 8.2225e-05 | 40.2124 | 0.1 | 0.01
m gain | -0.99999 | 5.0103e-05 | 52.5604 | 0.1 | 0.01
m linearity | 0.99998 | 6.7231e-05 | 45.8658 | 0.1 | 0.01
m gain | -0.1 | 2.2051e-06 | 352.5632 | 0.1 | 0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.7572e-05 | 44.3336 | 0.1 | -0.01
x gain | 1 | 7.9356e-05 | 34.0213 | 0.1 | -0.01
x linearity | 1 | 6.4264e-05 | 37.5443 | 0.1 | -0.01
x gain | 1 | 4.9923e-05 | 49.7943 | 0.1 | -0.01
m linearity | 0.99999 | 6.4429e-05 | 45.2702 | 0.1 | -0.01
m gain | -0.1 | 2.1467e-06 | 357.6442 | 0.1 | -0.01
m linearity | 1 | 4.9622e-05 | 49.6955 | 0.1 | -0.01
m gain | -1 | 5.5008e-05 | 47.9803 | 0.1 | -0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.8226e-06 | 69.3873 | 0.1 | 0.05
x gain | 1 | 7.0629e-06 | 48.9766 | 0.1 | 0.05
x linearity | 1 | 5.7953e-06 | 58.3855 | 0.1 | 0.05
x gain | 1 | 5.7479e-06 | 56.3082 | 0.1 | 0.05
m linearity | 1 | 6.6782e-06 | 42.9984 | 0.1 | 0.05
m gain | -1 | 6.1654e-06 | 62.2562 | 0.1 | 0.05
m linearity | 1 | 5.3392e-06 | 59.0104 | 0.1 | 0.05
m gain | -0.1 | 2.872e-07 | 180.9392 | 0.1 | 0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99999 | 6.5708e-06 | 38.5756 | 0.1 | -0.05
x gain | 1 | 6.7644e-06 | 54.3099 | 0.1 | -0.05
x linearity | 1 | 5.227e-06 | 56.1941 | 0.1 | -0.05
x gain | 1 | 7.4661e-06 | 48.0487 | 0.1 | -0.05
m linearity | 1 | 6.6156e-06 | 39.2373 | 0.1 | -0.05
m gain | -0.1 | 2.9142e-07 | 167.6028 | 0.1 | -0.05
m linearity | 1 | 4.485e-06 | 61.9252 | 0.1 | -0.05
m gain | -1 | 6.6704e-06 | 50.3136 | 0.1 | -0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.4091e-06 | 48.1606 | 0.1 | 0.1
x gain | 1 | 5.0559e-06 | 51.9456 | 0.1 | 0.1
x linearity | 1 | 6.5673e-06 | 43.6876 | 0.1 | 0.1
x gain | 1 | 6.5856e-06 | 41.7134 | 0.1 | 0.1
m linearity | 1 | 5.1138e-06 | 50.044 | 0.1 | 0.1
m gain | -1 | 5.7553e-06 | 44.5399 | 0.1 | 0.1
m linearity | 0.99999 | 6.0502e-06 | 40.886 | 0.1 | 0.1
m gain | -1 | 5.4414e-06 | 48.8524 | 0.1 | 0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.474e-06 | 43.2458 | 0.1 | -0.1
x gain | 1 | 6.196e-06 | 47.171 | 0.1 | -0.1
x linearity | 1 | 5.9711e-06 | 45.5006 | 0.1 | -0.1
x gain | 1 | 6.5384e-06 | 45.2034 | 0.1 | -0.1
m linearity | 1 | 5.6848e-06 | 45.5556 | 0.1 | -0.1
m gain | -1 | 5.6689e-06 | 43.9932 | 0.1 | -0.1
m linearity | 1 | 6.0835e-06 | 45.31 | 0.1 | -0.1
m gain | -1 | 5.9228e-06 | 44.7807 | 0.1 | -0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.6052e-06 | 37.532 | 2 | 0.5
x gain | 1 | 1.4839e-06 | 47.5475 | 2 | 0.5
x linearity | 1 | 1.5216e-06 | 46.5368 | 2 | 0.5
x gain | 1 | 2.3778e-06 | 41.5333 | 2 | 0.5
m linearity | 1 | 1.8629e-06 | 44.0769 | 1 | 0.5
m gain | -1 | 1.6016e-06 | 48.5358 | 1 | 0.5
m linearity | 1 | 1.4364e-06 | 47.6326 | 1 | 0.5
m gain | -1 | 2.098e-06 | 40.3694 | 1 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.8872e-06 | 41.1299 | 2 | -0.5
x gain | 1 | 2.0769e-06 | 36.7942 | 2 | -0.5
x linearity | 1 | 2.1497e-06 | 54.3409 | 2 | -0.5
x gain | 1 | 2.4839e-06 | 60.6698 | 2 | -0.5
m linearity | 1 | 1.7166e-05 | 44.0427 | 1 | -0.5
m gain | -0.1 | 5.3094e-07 | 377.4682 | 1 | -0.5
m linearity | 0.99999 | 2.1242e-05 | 53.2845 | 1 | -0.5
m gain | -0.099999 | 1.6315e-06 | 24.793 | 1 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.6817e-07 | 44.985 | 2 | 1.0
x gain | 1 | 8.6367e-07 | 48.4833 | 2 | 1.0
x linearity | 1 | 9.1925e-07 | 46.2568 | 2 | 1.0
x gain | 1 | 8.7926e-07 | 49.9028 | 2 | 1.0
m linearity | 1 | 8.3717e-07 | 43.4993 | 1 | 1.0
m gain | -1 | 8.7774e-07 | 48.3705 | 1 | 1.0
m linearity | 1 | 8.8806e-07 | 47.8013 | 1 | 1.0
m gain | -1 | 9.8072e-07 | 44.4814 | 1 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 9.8588e-07 | 46.1616 | 2 | -1.0
x gain | 1 | 9.09e-07 | 43.9653 | 2 | -1.0
x linearity | 1 | 6.8908e-06 | 19.3618 | 2 | -1.0
x gain | 0.99999 | 4.5647e-06 | 21.6134 | 2 | -1.0
m linearity | 1 | 9.7287e-07 | 46.5678 | 1 | -1.0
m gain | -1 | 8.7222e-07 | 46.7561 | 1 | -1.0
m linearity | 1 | 8.7255e-07 | 48.0097 | 1 | -1.0
m gain | -1 | 1.0414e-06 | 41.515 | 1 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.0627e-06 | 44.1167 | 20 | 0.5
x gain | 1 | 8.9862e-06 | 42.4872 | 20 | 0.5
x linearity | 0.99999 | 7.8867e-06 | 46.7492 | 20 | 0.5
x gain | 1 | 6.5614e-06 | 52.6338 | 20 | 0.5
m linearity | 1 | 6.7773e-06 | 46.7937 | 10 | 0.5
m gain | -0.99999 | 1.0019e-05 | 43.8029 | 10 | 0.5
m linearity | 1 | 7.2761e-06 | 46.511 | 10 | 0.5
m gain | -0.99999 | 9.1368e-06 | 33.3227 | 10 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.5238e-06 | 44.9309 | 20 | -0.5
x gain | 1 | 8.7093e-06 | 39.9327 | 20 | -0.5
x linearity | 1 | 8.6892e-06 | 45.2815 | 20 | -0.5
x gain | 1 | 9.4891e-06 | 43.9021 | 20 | -0.5
m linearity | 1 | 6.1156e-06 | 57.0155 | 10 | -0.5
m gain | -1 | 7.3687e-06 | 49.9674 | 10 | -0.5
m linearity | 1 | 9.159e-06 | 43.7299 | 10 | -0.5
m gain | -1 | 9.7491e-06 | 44.0569 | 10 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.9346e-06 | 43.6545 | 20 | 1.0
x gain | 1 | 2.4015e-05 | 19.6434 | 20 | 1.0
x linearity | 1 | 4.6445e-06 | 44.6107 | 20 | 1.0
x gain | 1 | 4.6221e-06 | 44.6971 | 20 | 1.0
m linearity | 1 | 4.3665e-06 | 45.1212 | 10 | 1.0
m gain | -1 | 2.4096e-05 | 19.908 | 10 | 1.0
m linearity | 1 | 4.0187e-06 | 40.7988 | 10 | 1.0
m gain | -1 | 4.7847e-06 | 44.9325 | 10 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.845e-06 | 48.3955 | 20 | -1.0
x gain | 1 | 4.2152e-06 | 46.9228 | 20 | -1.0
x linearity | 1 | 3.1942e-06 | 43.9948 | 20 | -1.0
x gain | 1 | 4.3753e-06 | 41.4273 | 20 | -1.0
m linearity | 1 | 4.9055e-06 | 38.252 | 10 | -1.0
m gain | -1 | 4.4204e-06 | 46.5619 | 10 | -1.0
m linearity | 1 | 3.832e-06 | 48.1552 | 10 | -1.0
m gain | -1 | 4.9182e-06 | 44.3428 | 10 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.6444e-07 | 53.0941 | 20 | 5.0
x gain | 1 | 4.5425e-07 | 54.5783 | 20 | 5.0
x linearity | 1 | 3.3189e-07 | 66.8714 | 20 | 5.0
x gain | 1 | 5.0766e-07 | 48.0114 | 20 | 5.0
m linearity | 1 | 4.184e-07 | 52.439 | 10 | 5.0
m gain | -1 | 5.1975e-07 | 52.9764 | 10 | 5.0
m linearity | 1 | 3.6292e-07 | 56.0706 | 10 | 5.0
m gain | -1 | 4.8957e-07 | 50.1564 | 10 | 5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.997e-07 | 54.6342 | 20 | -5.0
x gain | 1 | 4.0378e-07 | 61.4212 | 20 | -5.0
x linearity | 1 | 3.5759e-07 | 65.7501 | 20 | -5.0
x gain | 1 | 4.7001e-07 | 51.613 | 20 | -5.0
m linearity | 1 | 4.0506e-07 | 54.4373 | 10 | -5.0
m gain | -1 | 4.3405e-07 | 55.346 | 10 | -5.0
m linearity | 1 | 4.1565e-07 | 50.9243 | 10 | -5.0
m gain | -1 | 4.3656e-07 | 58.9781 | 10 | -5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.6682e-07 | 52.4447 | 20 | 10.0
x gain | 1 | 4.4462e-07 | 41.8406 | 20 | 10.0
x linearity | 1 | 1.4351e-07 | 79.5498 | 20 | 10.0
x gain | 1 | 1.2234e-07 | 173.4429 | 20 | 10.0
m linearity | 1 | 3.847e-07 | 53.3231 | 10 | 10.0
m gain | -1 | 5.098e-07 | 46.5276 | 10 | 10.0
m linearity | 1 | 2.2089e-07 | 48.1264 | 10 | 10.0
m gain | -1 | 1.2544e-07 | 220.3776 | 10 | 10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.2511e-07 | 171.9043 | 20 | -10.0
x gain | 1 | 1.5084e-07 | 100.241 | 20 | -10.0
x linearity | 1 | 1.1747e-07 | 182.8919 | 20 | -10.0
x gain | 1 | 1.4225e-07 | 90.4906 | 20 | -10.0
m linearity | 1 | 1.3805e-07 | 169.3247 | 10 | -10.0
m gain | -1 | 1.3216e-07 | 193.2511 | 10 | -10.0
m linearity | 1 | 2.2021e-07 | 34.8952 | 10 | -10.0
m gain | -1 | 1.2452e-07 | 208.6669 | 10 | -10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.7887e-06 | 28.5451 | 100 | 50.0
x gain | 1 | 4.3727e-06 | 111.52 | 100 | 50.0
x linearity | 0.99998 | 6.5347e-06 | 38.388 | 100 | 50.0
x gain | 0.99998 | 6.7544e-06 | 42.8707 | 100 | 50.0
m linearity | 1 | 2.3237e-05 | 24.5669 | 100 | 50.0
m gain | -1 | 4.6368e-06 | 112.7143 | 100 | 50.0
m linearity | 1 | 6.4531e-06 | 36.6446 | 100 | 50.0
m gain | -1 | 6.6155e-06 | 40.1001 | 100 | 50.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.1366e-06 | 47.5088 | 100 | 100.0
x gain | 1 | 1.3367e-06 | 43.2855 | 100 | 100.0
x linearity | 1 | 1.0388e-06 | 49.8772 | 100 | 100.0
x gain | 1 | 1.499e-06 | 43.6304 | 100 | 100.0
m linearity | 1 | 1.2356e-06 | 45.7186 | 100 | 100.0
m gain | -1 | 1.3041e-06 | 42.3789 | 100 | 100.0
m linearity | 1 | 1.1421e-06 | 46.507 | 100 | 100.0
m gain | -1 | 1.375e-06 | 42.8925 | 100 | 100.0

## Results4
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1.0002 | 0.00049367 | 49.0085 | 0.1 | 0.001
x gain | 1.0005 | 0.00061789 | 34.6293 | 0.1 | 0.001
x linearity | 1.0001 | 0.00051024 | 48.8452 | 0.1 | 0.001
x gain | 1 | 0.00057417 | 44.0937 | 0.1 | 0.001
m linearity | 0.99684 | 0.0055349 | 19.3918 | 0.1 | 0.001
m gain | -0.99602 | 0.0055702 | 19.9529 | 0.1 | 0.001
m linearity | 0.9967 | 0.0055291 | 19.3396 | 0.1 | 0.001
m gain | -0.99679 | 0.0055304 | 19.3513 | 0.1 | 0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 0.00045801 | 51.8717 | 0.1 | -0.001
x gain | 1 | 0.00071226 | 44.4284 | 0.1 | -0.001
x linearity | 1.0003 | 0.00056961 | 46.6474 | 0.1 | -0.001
x gain | 0.99975 | 0.00067054 | 43.8438 | 0.1 | -0.001
m linearity | 1.0002 | 0.0004693 | 51.7373 | 0.1 | -0.001
m gain | -1.0002 | 0.00078648 | 41.9116 | 0.1 | -0.001
m linearity | 0.99994 | 0.00068542 | 40.3018 | 0.1 | -0.001
m gain | -0.010002 | 1.9626e-06 | 314.3204 | 0.1 | -0.001
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99998 | 7.5519e-05 | 35.2731 | 0.1 | 0.01
x gain | 0.99999 | 4.5262e-05 | 61.4132 | 0.1 | 0.01
x linearity | 1 | 5.8133e-05 | 40.2455 | 0.1 | 0.01
x gain | 1 | 4.8985e-05 | 54.301 | 0.1 | 0.01
m linearity | 1 | 8.2225e-05 | 40.2124 | 0.1 | 0.01
m gain | -0.99999 | 5.0103e-05 | 52.5604 | 0.1 | 0.01
m linearity | 0.99998 | 6.7231e-05 | 45.8658 | 0.1 | 0.01
m gain | -0.1 | 2.2051e-06 | 352.5632 | 0.1 | 0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.7572e-05 | 44.3336 | 0.1 | -0.01
x gain | 1 | 7.9356e-05 | 34.0213 | 0.1 | -0.01
x linearity | 1 | 6.4264e-05 | 37.5443 | 0.1 | -0.01
x gain | 1 | 4.9923e-05 | 49.7943 | 0.1 | -0.01
m linearity | 0.99999 | 6.4429e-05 | 45.2702 | 0.1 | -0.01
m gain | -0.1 | 2.1467e-06 | 357.6442 | 0.1 | -0.01
m linearity | 1 | 4.9622e-05 | 49.6955 | 0.1 | -0.01
m gain | -1 | 5.5008e-05 | 47.9803 | 0.1 | -0.01
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.8226e-06 | 69.3873 | 0.1 | 0.05
x gain | 1 | 7.0629e-06 | 48.9766 | 0.1 | 0.05
x linearity | 1 | 5.7953e-06 | 58.3855 | 0.1 | 0.05
x gain | 1 | 5.7479e-06 | 56.3082 | 0.1 | 0.05
m linearity | 1 | 6.6782e-06 | 42.9984 | 0.1 | 0.05
m gain | -1 | 6.1654e-06 | 62.2562 | 0.1 | 0.05
m linearity | 1 | 5.3392e-06 | 59.0104 | 0.1 | 0.05
m gain | -0.1 | 2.872e-07 | 180.9392 | 0.1 | 0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 0.99999 | 6.5708e-06 | 38.5756 | 0.1 | -0.05
x gain | 1 | 6.7644e-06 | 54.3099 | 0.1 | -0.05
x linearity | 1 | 5.227e-06 | 56.1941 | 0.1 | -0.05
x gain | 1 | 7.4661e-06 | 48.0487 | 0.1 | -0.05
m linearity | 1 | 6.6156e-06 | 39.2373 | 0.1 | -0.05
m gain | -0.1 | 2.9142e-07 | 167.6028 | 0.1 | -0.05
m linearity | 1 | 4.485e-06 | 61.9252 | 0.1 | -0.05
m gain | -1 | 6.6704e-06 | 50.3136 | 0.1 | -0.05
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.4091e-06 | 48.1606 | 0.1 | 0.1
x gain | 1 | 5.0559e-06 | 51.9456 | 0.1 | 0.1
x linearity | 1 | 6.5673e-06 | 43.6876 | 0.1 | 0.1
x gain | 1 | 6.5856e-06 | 41.7134 | 0.1 | 0.1
m linearity | 1 | 5.1138e-06 | 50.044 | 0.1 | 0.1
m gain | -1 | 5.7553e-06 | 44.5399 | 0.1 | 0.1
m linearity | 0.99999 | 6.0502e-06 | 40.886 | 0.1 | 0.1
m gain | -1 | 5.4414e-06 | 48.8524 | 0.1 | 0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 5.474e-06 | 43.2458 | 0.1 | -0.1
x gain | 1 | 6.196e-06 | 47.171 | 0.1 | -0.1
x linearity | 1 | 5.9711e-06 | 45.5006 | 0.1 | -0.1
x gain | 1 | 6.5384e-06 | 45.2034 | 0.1 | -0.1
m linearity | 1 | 5.6848e-06 | 45.5556 | 0.1 | -0.1
m gain | -1 | 5.6689e-06 | 43.9932 | 0.1 | -0.1
m linearity | 1 | 6.0835e-06 | 45.31 | 0.1 | -0.1
m gain | -1 | 5.9228e-06 | 44.7807 | 0.1 | -0.1
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.6052e-06 | 37.532 | 2 | 0.5
x gain | 1 | 1.4839e-06 | 47.5475 | 2 | 0.5
x linearity | 1 | 1.5216e-06 | 46.5368 | 2 | 0.5
x gain | 1 | 2.3778e-06 | 41.5333 | 2 | 0.5
m linearity | 1 | 1.8629e-06 | 44.0769 | 1 | 0.5
m gain | -1 | 1.6016e-06 | 48.5358 | 1 | 0.5
m linearity | 1 | 1.4364e-06 | 47.6326 | 1 | 0.5
m gain | -1 | 2.098e-06 | 40.3694 | 1 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.8872e-06 | 41.1299 | 2 | -0.5
x gain | 1 | 2.0769e-06 | 36.7942 | 2 | -0.5
x linearity | 1 | 2.1497e-06 | 54.3409 | 2 | -0.5
x gain | 1 | 2.4839e-06 | 60.6698 | 2 | -0.5
m linearity | 1 | 1.7166e-05 | 44.0427 | 1 | -0.5
m gain | -0.1 | 5.3094e-07 | 377.4682 | 1 | -0.5
m linearity | 0.99999 | 2.1242e-05 | 53.2845 | 1 | -0.5
m gain | -0.099999 | 1.6315e-06 | 24.793 | 1 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.6817e-07 | 44.985 | 2 | 1.0
x gain | 1 | 8.6367e-07 | 48.4833 | 2 | 1.0
x linearity | 1 | 9.1925e-07 | 46.2568 | 2 | 1.0
x gain | 1 | 8.7926e-07 | 49.9028 | 2 | 1.0
m linearity | 1 | 8.3717e-07 | 43.4993 | 1 | 1.0
m gain | -1 | 8.7774e-07 | 48.3705 | 1 | 1.0
m linearity | 1 | 8.8806e-07 | 47.8013 | 1 | 1.0
m gain | -1 | 9.8072e-07 | 44.4814 | 1 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 9.8588e-07 | 46.1616 | 2 | -1.0
x gain | 1 | 9.09e-07 | 43.9653 | 2 | -1.0
x linearity | 1 | 6.8908e-06 | 19.3618 | 2 | -1.0
x gain | 0.99999 | 4.5647e-06 | 21.6134 | 2 | -1.0
m linearity | 1 | 9.7287e-07 | 46.5678 | 1 | -1.0
m gain | -1 | 8.7222e-07 | 46.7561 | 1 | -1.0
m linearity | 1 | 8.7255e-07 | 48.0097 | 1 | -1.0
m gain | -1 | 1.0414e-06 | 41.515 | 1 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.0627e-06 | 44.1167 | 20 | 0.5
x gain | 1 | 8.9862e-06 | 42.4872 | 20 | 0.5
x linearity | 0.99999 | 7.8867e-06 | 46.7492 | 20 | 0.5
x gain | 1 | 6.5614e-06 | 52.6338 | 20 | 0.5
m linearity | 1 | 6.7773e-06 | 46.7937 | 10 | 0.5
m gain | -0.99999 | 1.0019e-05 | 43.8029 | 10 | 0.5
m linearity | 1 | 7.2761e-06 | 46.511 | 10 | 0.5
m gain | -0.99999 | 9.1368e-06 | 33.3227 | 10 | 0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 7.5238e-06 | 44.9309 | 20 | -0.5
x gain | 1 | 8.7093e-06 | 39.9327 | 20 | -0.5
x linearity | 1 | 8.6892e-06 | 45.2815 | 20 | -0.5
x gain | 1 | 9.4891e-06 | 43.9021 | 20 | -0.5
m linearity | 1 | 6.1156e-06 | 57.0155 | 10 | -0.5
m gain | -1 | 7.3687e-06 | 49.9674 | 10 | -0.5
m linearity | 1 | 9.159e-06 | 43.7299 | 10 | -0.5
m gain | -1 | 9.7491e-06 | 44.0569 | 10 | -0.5
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.9346e-06 | 43.6545 | 20 | 1.0
x gain | 1 | 2.4015e-05 | 19.6434 | 20 | 1.0
x linearity | 1 | 4.6445e-06 | 44.6107 | 20 | 1.0
x gain | 1 | 4.6221e-06 | 44.6971 | 20 | 1.0
m linearity | 1 | 4.3665e-06 | 45.1212 | 10 | 1.0
m gain | -1 | 2.4096e-05 | 19.908 | 10 | 1.0
m linearity | 1 | 4.0187e-06 | 40.7988 | 10 | 1.0
m gain | -1 | 4.7847e-06 | 44.9325 | 10 | 1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.845e-06 | 48.3955 | 20 | -1.0
x gain | 1 | 4.2152e-06 | 46.9228 | 20 | -1.0
x linearity | 1 | 3.1942e-06 | 43.9948 | 20 | -1.0
x gain | 1 | 4.3753e-06 | 41.4273 | 20 | -1.0
m linearity | 1 | 4.9055e-06 | 38.252 | 10 | -1.0
m gain | -1 | 4.4204e-06 | 46.5619 | 10 | -1.0
m linearity | 1 | 3.832e-06 | 48.1552 | 10 | -1.0
m gain | -1 | 4.9182e-06 | 44.3428 | 10 | -1.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 4.6444e-07 | 53.0941 | 20 | 5.0
x gain | 1 | 4.5425e-07 | 54.5783 | 20 | 5.0
x linearity | 1 | 3.3189e-07 | 66.8714 | 20 | 5.0
x gain | 1 | 5.0766e-07 | 48.0114 | 20 | 5.0
m linearity | 1 | 4.184e-07 | 52.439 | 10 | 5.0
m gain | -1 | 5.1975e-07 | 52.9764 | 10 | 5.0
m linearity | 1 | 3.6292e-07 | 56.0706 | 10 | 5.0
m gain | -1 | 4.8957e-07 | 50.1564 | 10 | 5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.997e-07 | 54.6342 | 20 | -5.0
x gain | 1 | 4.0378e-07 | 61.4212 | 20 | -5.0
x linearity | 1 | 3.5759e-07 | 65.7501 | 20 | -5.0
x gain | 1 | 4.7001e-07 | 51.613 | 20 | -5.0
m linearity | 1 | 4.0506e-07 | 54.4373 | 10 | -5.0
m gain | -1 | 4.3405e-07 | 55.346 | 10 | -5.0
m linearity | 1 | 4.1565e-07 | 50.9243 | 10 | -5.0
m gain | -1 | 4.3656e-07 | 58.9781 | 10 | -5.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 3.6682e-07 | 52.4447 | 20 | 10.0
x gain | 1 | 4.4462e-07 | 41.8406 | 20 | 10.0
x linearity | 1 | 1.4351e-07 | 79.5498 | 20 | 10.0
x gain | 1 | 1.2234e-07 | 173.4429 | 20 | 10.0
m linearity | 1 | 3.847e-07 | 53.3231 | 10 | 10.0
m gain | -1 | 5.098e-07 | 46.5276 | 10 | 10.0
m linearity | 1 | 2.2089e-07 | 48.1264 | 10 | 10.0
m gain | -1 | 1.2544e-07 | 220.3776 | 10 | 10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.2511e-07 | 171.9043 | 20 | -10.0
x gain | 1 | 1.5084e-07 | 100.241 | 20 | -10.0
x linearity | 1 | 1.1747e-07 | 182.8919 | 20 | -10.0
x gain | 1 | 1.4225e-07 | 90.4906 | 20 | -10.0
m linearity | 1 | 1.3805e-07 | 169.3247 | 10 | -10.0
m gain | -1 | 1.3216e-07 | 193.2511 | 10 | -10.0
m linearity | 1 | 2.2021e-07 | 34.8952 | 10 | -10.0
m gain | -1 | 1.2452e-07 | 208.6669 | 10 | -10.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 8.7887e-06 | 28.5451 | 100 | 50.0
x gain | 1 | 4.3727e-06 | 111.52 | 100 | 50.0
x linearity | 0.99998 | 6.5347e-06 | 38.388 | 100 | 50.0
x gain | 0.99998 | 6.7544e-06 | 42.8707 | 100 | 50.0
m linearity | 1 | 2.3237e-05 | 24.5669 | 100 | 50.0
m gain | -1 | 4.6368e-06 | 112.7143 | 100 | 50.0
m linearity | 1 | 6.4531e-06 | 36.6446 | 100 | 50.0
m gain | -1 | 6.6155e-06 | 40.1001 | 100 | 50.0
Label | Ratio | STDEV | Effct. DoF | Range | Max/Min V Setting
x linearity | 1 | 1.1366e-06 | 47.5088 | 100 | 100.0
x gain | 1 | 1.3367e-06 | 43.2855 | 100 | 100.0
x linearity | 1 | 1.0388e-06 | 49.8772 | 100 | 100.0
x gain | 1 | 1.499e-06 | 43.6304 | 100 | 100.0
m linearity | 1 | 1.2356e-06 | 45.7186 | 100 | 100.0
m gain | -1 | 1.3041e-06 | 42.3789 | 100 | 100.0
m linearity | 1 | 1.1421e-06 | 46.507 | 100 | 100.0
m gain | -1 | 1.375e-06 | 42.8925 | 100 | 100.0